## log
Exp800.10
What is the optimal concentration of IAC in Legionella pneumophila assay?
20210415
Legionella species assay uses 1e3 copies/ul as IAC concentration. This results in a Cq of 38-40.
Greg explains that this reaction out-competes Leg spp. Reaction so want a late crossover.
This reaction doesn't need to be successful if sample is successful. Thus, okay if it fails.
However, if sample fails, then this reaction needs to succeed for troubleshooting and control purposes.
I think 38-40 is too late and increases risk of not-observing control.
Will aim for 35 cycles.
Neither Greg nor I have observed IAC fluor in Leg pneumophila assay using various sources and concentrations of IAC.
In one study I used std_3 (high conc) and did not observe IAC fluor.
I used 38 cycles because I was interested in determining LOD.
However, didn't observe IAC at 45 cycles either.
yesterday, I used DeNovix spec to calc concentration of undiluted IAC.
Raquel says that tube arrives at concentration of 10uM, not 100uM. | [9 empty] | 20210414
[10 empty] | 1943400 | IAC_DEER Extinction Co
measured absorbance of 10uM from Raquel on spectrophotometer using DeNovix.
[10 empty] | 202.115 | ng/ul
|  | IDT DNA Data (IAC) |  |  |  | Spec Data |  |  |  | 6.1247 | abs@260
| TubeName | Conc (uM) | pM | Copy Number | Copy Number/rxn | Conc (uM) | pM | Copy Number | Copy Number/ul | Conc | 3.1515 | uM | should be 10uM.
| IDT Source | 10 | 10000000 |  |  | 3.1515 | 3151538.5407
| std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 0.031515 | 31515.3854 | 1.8979e+13 | 37957130184.2132
| std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 0.0031515 | 3151.5385 | 1.8979e+12 | 3795713018.4213
| std_3 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 0.00031515 | 315.1539 | 189785650921.0662 | 379571301.8421
| std_4 | 0.0001 | 100 | 60220000000 | 120440000 | 3.1515e-05 | 31.5154 | 18978565092.1066 | 37957130.1842
| std_5 | 1e-05 | 10 | 6022000000 | 12044000 | 3.1515e-06 | 3.1515 | 1897856509.2107 | 3795713.0184
| std_6 | 1e-06 | 1 | 602200000 | 1204400 | 3.1515e-07 | 0.31515 | 189785650.9211 | 379571.3018
| std_7 | 1e-07 | 0.1 | 60220000 | 120440 | 3.1515e-08 | 0.031515 | 18978565.0921 | 37957.1302
| std_8 | 1e-08 | 0.01 | 6022000 | 12044 | 3.1515e-09 | 0.0031515 | 1897856.5092 | 3795.713
| std_9 | 1e-09 | 0.001 | 602200 | 1204.4 | 3.1515e-10 | 0.00031515 | 189785.6509 | 379.5713
| std_10 | 1e-10 | 0.0001 | 60220 | 120.44 | 3.1515e-11 | 3.1515e-05 | 18978.5651 | 37.9571
| std_11 | 1e-11 | 1e-05 | 6022 | 12.044 | 3.1515e-12 | 3.1515e-06 | 1897.8565 | 3.7957
| std_12 | 1e-12 | 1e-06 | 602.2 | 1.2044 | 3.1515e-13 | 3.1515e-07 | 189.7857 | 0.37957
| std_13 | 1e-13 | 1e-07 | 60.22 | 0.12044 | 3.1515e-14 | 3.1515e-08 | 18.9786 | 0.037957
| std_14 | 1e-14 | 1e-08 | 6.022 | 0.012044 | 3.1515e-15 | 3.1515e-09 | 1.8979 | 0.0037957
| std_15 | 1e-15 | 1e-09 | 0.6022 | 0.0012044 | 3.1515e-16 | 3.1515e-10 | 0.18979 | 0.00037957
| std_16 | 1e-16 | 1e-10 | 0.06022 | 0.00012044 | 3.1515e-17 | 3.1515e-11 | 0.018979 | 3.7957e-05
| std_17 | 1e-17 | 1e-11 | 0.006022 | 1.2044e-05 | 3.1515e-18 | 3.1515e-12 | 0.0018979 | 3.7957e-06
| std_18 | 1e-18 | 1e-12 | 0.0006022 | 1.2044e-06 | 3.1515e-19 | 3.1515e-13 | 0.00018979 | 3.7957e-07
|  | Supplier | Cat # | Component | Size (Average) | Units | Stock Concentration | Units
|  | IDT | 10006833 | N2 Forward | 1000 | uL | 100 | uM
|  | IDT | 10006834 | N2 Reverse | 1000 | uL | 100 | uM | Raquel found this in a COVID-19 doc
|  | IDT | 10006835 | N2 Probe | 500 | uL | 100 | uM | https://luminultratechnologiesltd.sharepoint.com/:x:/r/LuminUltra%20Corporate/qms/_layouts/15/Doc.aspx?sourcedoc=%7B0C151CC3-2386-4BB6-A817-F6B1779387FB%7D&file=SARS-CoV-2%20Direct%20-%20Updated%20BOM%20Calculations.xlsx&action=default&mobileredirect=true&DefaultItemOpen=1
|  | IDT | Custom | IAC Forward Primer (100uM) | 1300 | uL | 100 | uM
|  | IDT | Custom | IAC Reverse Primer (100uM) | 1300 | uL | 100 | uM
|  | IDT | Custom | IAC Probe (100uM) | 550 | uL | 100 | uM
|  | IDT | Custom | IAC Template (Ultramer) | 40 | uL | 100 | uM | <-- Stock is ~ 6x10^13 copies/uL. Dilute 1.5uL into 898.5uL TE to make 1x10^11 solution and then dilute by serial 1:100 dilutions (2uL in 198uL TE) to 1x10^5 for final use. Use H20 for final dilution to 1x10^5.
This agrees with my value in 10ul (think it's 4x higher in 100uM), but I disagree with per ul. Another way to say this is that there are 6e13 copies in 10ul. If this 10ul is added to 990ul water, then it can be 6e13/1000 or 6e10 copies/ul (Red is from Neil)
20210415
Ran 8 tubes just as a quick check to see if IAC would even amplify at std_1 concentration.
20210427
Running primer matrix and probe matrix shows that optimal concentration of primer and probe is 50, 50nM and 400nM.
Running another plate using Exp800.04 std curve using optimized conditions in presence of Leg pneumo shows failed wells.
Was this incorrect setup? Peformed Thursday before trip to TX. Did I miss something?
Want to make a whole plate and compare it to optimized concentration. Will use std_6 Leg pneumo. Ask Raquel to make two mmixes:
mmix base =
9.36
Component | Desired | Units | PROBE RXN | 115.2
mmix |  |  | 9.36 | 1078.272
Leg pneumo Probe | 500 |  | 1 | 115.2
Leg pneumo "std_5" |  |  | 0.2 | 23.04
water |  |  | 7.44 | 857.088
TOTAL |  |  | 18 | 2073.6
Will use "std_5" instead of "std_6" because vol is too high.
Good data. See sheet "std_vs_opt"
Talked with Greg, he reminds me that with lyophilization, sample fluorescence will take a hit.
Greg also wondered what IAC would do in presence of lower Leg pneumo pos control.
So, I'm repeating the assay with std_12 in place of std_5.
20210430
I repeated assay with std_12, "LO_std_vs_opt"
posted details in Monday.com: https://luminultra.monday.com/boards/1049377695/pulses/1242653111
Neil and Greg still concerned that IAC concentration is so high.
Neil and I discussed assays.
See also email to Pat Whalen for a good synposis and updating thinking, wordage.
Why is IAC positive control high?
The high (3e9 copies/ul) concentration is disconcerting for three reasons:
1) It's a lot of DNA. Inefficient.
2) It's not high in Legionella spp. assay.
3) It might indicate other systemic problems with this assay e.g. Leg pneumo primer, probe interactions, incompatibility with or miscalculation with reagent like dNTPs or polymerases etc.
Three factors might be causing it:
1) Leg pneumo effects. Maybe there's something about this F, R primer and probe that interact with IAC F, R primers and probe. The IAC positive control is a synthetic construct that does not match any sequence in GenBank source.
2) My calculations could be off. Maybe, but not too difficult to calculate the concentration. I've been through these calculations several times to confirm and check.
3) The original concentration of IAC is incorrect. @Raquel Jaramillo says the IDT DNA tube is 10uM. I've confirmed this with the spectrophotometer also confirming my calculations.
My hypothesis is that IAC is high because of Leg pneumo effects. This explains the null fluorescence in previous reactions prepared by @Greg Howard and myself.
Spoke with @Neil Sharma today. We're doing the following to test the hypothesis that IAC concentration is a function of the presence of Leg pneumo primers and probe:
1) Raquel will make two master mixes with everything similar except the Leg pneumo primers and probe will be swapped out with Leg spp primers and probe. I'll use the same standards I previously prepared. I expect to find a lower Cq for IAC concentration in Leg spp. assay.
2) Raquel will make two master mixes. The master mix will contain a different polymerase. One master mix will contain Leg pneumo primers and probe in addition to IAC primers and probe. The other master mix will only contain IAC primers and probe. I expect to find a lower Cq for IAC concentration in the absence of Leg pneumo primers and probe.
20210503
None of the master mixes provided by Raquel have any fluorescence. Wrong primer?
found this graphic when searching for LOQ. Want to incorporate a version into my reports.
https://www.eflm.eu/files/efcc/Zagreb-Theodorsson_2.pdf
20210505
Raquel repeated Leg pneumo master mix from "Why_so_high_mmix". Confirmed IAC primers. Had her check sequences. However, no Leg pneumo probe was added because this is on backorder.
Wrote Laara for status update and she wrote company, ThermoFisher, but no response yet.
Made 96 well plate yesterday using master mix. No fluorescence in any channel. To be expected in FAM channel because not Leg pneumo probe available, but expected to see it in IAC.
Testing with my own IAC primers probe to see if:
1) Concentration of IAC primers probe is a factor (currently at my optimized concentrations: 50, 50, 400nM)
2) Incorrect IAC primers probe.
mmix base =
9.36
| Component | Desired | Units | PROBE RXN | 115.2
| mmix |  |  | 9.36 | 1078.272
| Leg pneumo "std_5" |  |  | 0.2 | 23.04
| IAC_F | 0.25uM |  | 0.5 | 57.6
| IAC_R | 0.25uM |  | 0.5 | 57.6
| IAC_P | 0.5uM |  | 1 | 115.2
| water |  |  | 6.44 | 741.888
| TOTAL |  |  | 18 | 2073.6
again, no fluorescence.
Making new IAC std control.
From 10uM undiluted template that Raquel provided me.
Add 10ul to 990ul H2O.
use a little bit of the remaining master mix to test the reaction mix to see if it works.
| Component | Desired | Units | PROBE RXN | 16.8
| mmix |  |  | 9.36 | 157.248
| Leg pneumo "std_5" |  |  | 0.2 | 3.36
| Leg pneumo probe |  |  | 1 | 16.8
| IAC_F | 0.25uM |  | 0.5 | 8.4
| IAC_R | 0.25uM |  | 0.5 | 8.4
| IAC_P | 0.5uM |  | 1 | 16.8
| water |  |  | 5.44 | 91.392
| TOTAL |  |  | 18 | 302.4
20210706
20210709
Lyo samples provided by Raquel work well.
Testing LOD on 12 replicates dispensed by new program.
Used Tube 11, 12, 13, 14.
Duplicated 11, 12, 13.
Calculations in 3rd LOD.
Also included 1st LOD and 2nd LOD in sheet called "LOD_sort".
Note on C24: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is expected concentration from the tube.
Note on E24: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on F24: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul. Reply: assumes copy number is in 1000ul
Note on I24: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on J24: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul.
Note on B25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Note, that this is different than micH std or other standards as I usually have 100x dilution from std_0 to std_1. However, this standard is simply 10x dilution.
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Note on I25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Note on A99: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Leg pneumo probe has been ordered but not yet arrived. I asked Raquel to make mastermix without probe. I would add that down here because I have leftover probe from reactions.
Note on B171: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Leg pneumo probe has been ordered but not yet arrived. I asked Raquel to make mastermix without probe. I would add that down here because I have leftover probe from reactions.
Note on B187: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Leg pneumo probe has been ordered but not yet arrived. I asked Raquel to make mastermix without probe. I would add that down here because I have leftover probe from reactions.
Note on B188: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I have enough probe to complete these reactions.

## ref
Legionella - Master Mix for Lyophilization |  |  | Number of Reactions = |  | 400
Stock Concentrations |  |  | Overage = |  | 5 | %
LegGF | 100 | uM
LegGR | 100 | uM
Leg spp MGB probe | 100 | uM
IAC Forward | 100 | uM
IAC Reverse | 100 | uM
IAC Probe RED (CY5) | 100 | uM
Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 1000
dNTP Mix | 25 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
KAPA3G | 30 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Dextran | 25 | Percent
Component | Desired |  | Per 20 uL |  | Multiple Reactions |  |  | Lot
DotAF | 0.4 | uM | 0.08 | uL | 33.6 | uL | Forward 1
DotAR | 0.4 | uM | 0.08 | uL | 33.6 | uL | Reverse 1
Leg pne DotA MGB probe | 0.5 | uM | 0.1 | uL | 42 | uL | Probe 1
IAC Forward | 0.25 | uM | 0 | uL | 0 | uL | Forward 1
IAC Reverse | 0.25 | uM | 0 | uL | 0 | uL | Reverse 1
IAC Probe RED (CY5) | 0.5 | uM | 0 | uL | 0 | uL
Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL | 0 | uL
dNTP Mix | 0.25 | mM | 0.32 | uL | 134.4 | uL | dNTP Mix
AptaTaq | 2.5 | Units | 0.05 | uL | 21 | uL | AptaTaq
KAPA3G | 1 | Units | 0.033333 | uL | 14 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 29.4 | uL | Mg2+
Trehalose | 10 | Percent | 4 | uL | 1680 | uL | Trehalose
BSA | 0.5 | mg/mL | 0.2 | uL | 84 | uL | BSA
PCR Buffer 10X | 1 | x | 2 | uL | 840 | uL | PCR Buffer 10X
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 933.3333 | uL | Ficoll
Dextran | 0 | Percent | 0 | uL | 0 | uL
Water |  |  | 0 | uL | 0 | uL | Water
|  | Total: | 9.1556 | uL | 3845.3333 | uL
[3/10 1:44 PM] Neil Sharma
GeneCount Legionella pneumophlia (In Development)
LegGF
LegGR
Leg pne DotA MGB | 6-FAM-ATACAGCAAATGTATGTGACTT
DotAF | ATTGTCTCGCGCGATTGC
DotAR | CCGGATCATTATTAACCATSAC
IAC Probe RED | Cy5-agctagtcgatgcactccagtcctcct-TAO/ 3' IBRQ
IAC DNA (Deer_IAC_Ultramer) | tacagcaccctagcttggtagaatcgatcagctacgtgaggtcctacgacgatcgccaagcatgccctagctaagatgcatcgattgctcatcacgtacgttaggtcgactaggaggactggagtgcatcgactagctaagatggttcgattgctcatcacgaaggttaggtcgactacgaacgagtcgtacgttaggtt
Positive Control (Leg_pne_UltraPOS) | CCGGATCATTATTAACCATCACCTGTGCCACCGTAGATAAAGTCACATACATTTGCTGTATAGCAATCGCGCGAGACAAT
[3/10 1:50 PM] Neil Sharma
GeneCount Legionella Species
LegGF | CTAATTggCTgATTgTCTTgAC
LegGR | ggCgATgACCTACTTTCg
Leg Spp MGB Probe (Order From Invitrogen) | 6FAM- gTTTCACTTCTgAgTTCg
IAC Forward | ctaaccttcgtgatgagcaatcg
IAC Reverse | gatcagctacgtgaggtcctac
IAC Probe RED | Cy5-agctagtcgatgcactccagtcctcct-TAO/ 3' IBRQ
IAC DNA (Deer_IAC_Ultramer) | tacagcaccctagcttggtagaatcgatcagctacgtgaggtcctacgacgatcgccaagcatgccctagctaagatgcatcgattgctcatcacgtacgttaggtcgactaggaggactggagtgcatcgactagctaagatggttcgattgctcatcacgaaggttaggtcgactacgaacgagtcgtacgttaggtt
Positive Control (Leg_spp_UltraPOS) | TCTAATTGGCTGATTGTCTTGACCATATAATCTGAGTGACTTCAGAATGTGATATTGATTTGTATACGTGAAACGTATCGTGTAAACTCTGACTCTTTACCAAACCTGTGGCTTACTGATACCATCTCGAACTCAGAAGTGAAACATTTCCGCGCCAATGATAGTGTGAGGCTTCCTCATGCGAAAGTAGGTCATCGCCT

## plate1_setup
Process
Used the same process for making std curve.
The std_curve mastermix had Legionella pneumophila positive control doped in.
Used stds 6, 9 , 12 from Leg pneumo.
mastermix |  |  |  | per 3 | per 8 sets of 3 | per 4 sets of 24 | Component | Desired |  | Per 20 uL |  | Multiple Reactions |  |  | Lot
Component | Desired | Units | per 20 ul | 3.42 |  |  | DotAF | 0.4 | uM | 0.08 | uL | 29.4 | uL | Forward 1
mastermix |  |  | 9.3556 | 31.996 |  | 1077.76 | DotAR | 0.4 | uM | 0.08 | uL | 29.4 | uL | Reverse 1
water |  |  | 5.6444 | 19.304 |  | 650.24 | Leg pne DotA MGB probe | 0.5 | uM | 0.1 | uL | 36.75 | uL | Probe 1
TOTAL |  |  | 15 | 51.3 | 410.4 | 1728 | IAC Forward | 0.25 | uM | 0.05 | uL | 18.375 | uL | Forward 1
[7 empty] | IAC Reverse | 0.25 | uM | 0.05 | uL | 18.375 | uL | Reverse 1
each mastermix has their own conc of Legionella. 4 sets of 24 so 4 diff conc of Leg stds can be used. | [6 empty] | IAC Probe RED (CY5) | 0.5 | uM | 0.1 | uL | 36.75 | uL
[5 empty] | per 8 sets of 3 | overage = 1.14 | Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL | 0 | uL
Component | Desired | Units | per 20 ul | *3 | 24 | 27.36 | dNTP Mix | 0.25 | mM | 0.32 | uL | 117.6 | uL | dNTP Mix
mastermix |  |  | 9.3556 | 28.0667 | 224.5333 | 255.968 | AptaTaq | 2.5 | Units | 0.05 | uL | 18.375 | uL | AptaTaq
water |  |  | 6.6444 | 19.9333 | 159.4667 | 181.792 | KAPA3G | 1 | Units | 0.033333 | uL | 12.25 | uL
Leg std DNA |  |  | 2 | 6 | 48 | 54.72 | Mg2+ | 3.5 | mM | 0.07 | uL | 25.725 | uL | Mg2+
TOTAL |  |  | 18 | 54 | 432 | 492.48 | Trehalose | 10 | Percent | 4 | uL | 1470 | uL | Trehalose
[7 empty] | BSA | 0.5 | mg/mL | 0.2 | uL | 73.5 | uL | BSA
[7 empty] | PCR Buffer 10X | 1 | x | 2 | uL | 735 | uL | PCR Buffer 10X
[7 empty] | 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 816.6667 | uL | Ficoll
[7 empty] | Dextran | 0 | Percent | 0 | uL | 0 | uL
[7 empty] | Water |  |  | 5.6444 | uL | 2074.3333 | uL | Water
[9 empty] | Total: | 15 | uL | 5512.5 | uL
Note on D21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel did not add water to this mastermix.

## 210416_plate1
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | Ct_mean |  | std_6 |  |  | std_9 |  |  | std_12 |  |  | water
A01 | IAC | Unknown | Cy5 | 32.69 | A01 | IAC | Unknown | Cy5 | 32.69 | 32.7867 | 0.087369 | IAC Tube | Copy Number/ul | Ct_mean | IAC Tube | Copy Number/ul | Ct_mean | IAC Tube | Copy Number/ul | Ct_mean | IAC Tube | Copy Number/ul | Ct_mean
A01 | Leg pneumo | Unknown | FAM | 28.51 | A02 | IAC | Unknown | Cy5 | 32.86 |  |  | std_1 | 60220000000 | 32.7867 | std_1 | 60220000000 | 31.9333 | std_1 | 60220000000 | 32.2567 | std_1 | 60220000000 | 32.8167
A02 | IAC | Unknown | Cy5 | 32.86 | A03 | IAC | Unknown | Cy5 | 32.81 |  |  | std_2 | 12044000000 | 37.2233 | std_2 | 12044000000 | 36.3967 | std_2 | 12044000000 | 36.3933 | std_2 | 12044000000 | 36.86
A02 | Leg pneumo | Unknown | FAM | 28.57 | A04 | IAC | Unknown | Cy5 | 36.95 | 37.2233 | 0.41429 | std_3 | 1204400000 | 42.4367 | std_3 | 1204400000 | 41.2833 | std_3 | 1204400000 | 41.2233 | std_3 | 1204400000 | 41.3833
A03 | IAC | Unknown | Cy5 | 32.81 | A05 | IAC | Unknown | Cy5 | 37.7 |  |  | std_4 | 120440000 |  | std_4 | 120440000 |  | std_4 | 120440000 |  | std_4 | 120440000
A03 | Leg pneumo | Unknown | FAM | 28.77 | A06 | IAC | Unknown | Cy5 | 37.02 |  |  | std_5 | 12044000 |  | std_5 | 12044000 |  | std_5 | 12044000 |  | std_5 | 12044000
A04 | IAC | Unknown | Cy5 | 36.95 | A07 | IAC | Unknown | Cy5 | 42.02 | 42.4367 | 0.48563 | std_6 | 1204400 |  | std_6 | 1204400 |  | std_6 | 1204400 |  | std_6 | 1204400
A04 | Leg pneumo | Unknown | FAM | 29.86 | A08 | IAC | Unknown | Cy5 | 42.32 |  |  | std_7 | 120440 |  | std_7 | 120440 |  | std_7 | 120440 |  | std_7 | 120440
A05 | IAC | Unknown | Cy5 | 37.7 | A09 | IAC | Unknown | Cy5 | 42.97 |  |  | std_8 | 12044 |  | std_8 | 12044 |  | std_8 | 12044 |  | std_8 | 12044
A05 | Leg pneumo | Unknown | FAM | 29.44 | A10 | IAC | Unknown | Cy5
A06 | IAC | Unknown | Cy5 | 37.02 | A11 | IAC | Unknown | Cy5
A06 | Leg pneumo | Unknown | FAM | 29.31 | A12 | IAC | Unknown | Cy5
A07 | IAC | Unknown | Cy5 | 42.02 | B01 | IAC | Unknown | Cy5 |  |  |  | IAC Tube | Copy Number/ul | Log(Copy#) | Lpneumo "std_6" | Ct_6_STDEV | 2*STDEV | Lpneumo "std_9" | Ct_9_STDEV | 2*STDEV | Lpneumo "std_12" | Ct_12_STDEV | 2*STDEV | Lpneumo "H2O" | Ct_H2O_STDEV | 2*STDEV
A07 | Leg pneumo | Unknown | FAM | 29.86 | B02 | IAC | Unknown | Cy5 |  |  |  | std_1 | 60220000000 | 10.7797 | 32.7867 | 0.087369 | 0.17474 | 31.9333 | 0.50521 | 1.0104 | 32.2567 | 0.13503 | 0.27006 | 32.8167 | 0.31786 | 0.63571
A08 | IAC | Unknown | Cy5 | 42.32 | B03 | IAC | Unknown | Cy5 |  |  |  | std_2 | 12044000000 | 10.0808 | 37.2233 | 0.41429 | 0.82857 | 36.3967 | 0.20793 | 0.41585 | 36.3933 | 0.11846 | 0.23692 | 36.86 | 0.5246 | 1.0492
A08 | Leg pneumo | Unknown | FAM | 29.96 | B04 | IAC | Unknown | Cy5 |  |  |  | std_3 | 1204400000 | 9.0808 | 42.4367 | 0.48563 | 0.97125 | 41.2833 | 0.32716 | 0.65432 | 41.2233 | 0.306 | 0.61199 | 41.3833 | 0.16503 | 0.33005
A09 | IAC | Unknown | Cy5 | 42.97 | B05 | IAC | Unknown | Cy5
A09 | Leg pneumo | Unknown | FAM | 30.2 | B06 | IAC | Unknown | Cy5
A10 | IAC | Unknown | Cy5 |  | B07 | IAC | Unknown | Cy5 | [18 empty] | x | b | PCR Efficiency
A10 | Leg pneumo | Unknown | FAM | 29.96 | B08 | IAC | Unknown | Cy5 | [17 empty] | std_6 | -5.6478 | 93.8498 | 50.33%
A11 | IAC | Unknown | Cy5 |  | B09 | IAC | Unknown | Cy5 | [17 empty] | std_9 | -5.4609 | 91.0399 | 52.45%
A11 | Leg pneumo | Unknown | FAM | 29.95 | B10 | IAC | Unknown | Cy5 | [17 empty] | std_12 | -5.2469 | 88.9908 | 55.09%
A12 | IAC | Unknown | Cy5 |  | B11 | IAC | Unknown | Cy5 | [17 empty] | H2O | -5.0066 | 86.9877 | 58.39%
A12 | Leg pneumo | Unknown | FAM | 30.14 | B12 | IAC | Unknown | Cy5
B01 | IAC | Unknown | Cy5 |  | C01 | IAC | Unknown | Cy5 | 32.23 | 31.9333 | 0.50521
B01 | Leg pneumo | Unknown | FAM | 29.81 | C02 | IAC | Unknown | Cy5 | 32.22
B02 | IAC | Unknown | Cy5 |  | C03 | IAC | Unknown | Cy5 | 31.35
B02 | Leg pneumo | Unknown | FAM | 29.96 | C04 | IAC | Unknown | Cy5 | 36.16 | 36.3967 | 0.20793
B03 | IAC | Unknown | Cy5 |  | C05 | IAC | Unknown | Cy5 | 36.55
B03 | Leg pneumo | Unknown | FAM | 29.69 | C06 | IAC | Unknown | Cy5 | 36.48
B04 | IAC | Unknown | Cy5 |  | C07 | IAC | Unknown | Cy5 | 41.42 | 41.2833 | 0.32716
B04 | Leg pneumo | Unknown | FAM | 29.86 | C08 | IAC | Unknown | Cy5 | 41.52
B05 | IAC | Unknown | Cy5 |  | C09 | IAC | Unknown | Cy5 | 40.91
B05 | Leg pneumo | Unknown | FAM | 30.14 | C10 | IAC | Unknown | Cy5
B06 | IAC | Unknown | Cy5 |  | C11 | IAC | Unknown | Cy5
B06 | Leg pneumo | Unknown | FAM | 30.06 | C12 | IAC | Unknown | Cy5
B07 | IAC | Unknown | Cy5 |  | D01 | IAC | Unknown | Cy5
B07 | Leg pneumo | Unknown | FAM | 30.19 | D02 | IAC | Unknown | Cy5
B08 | IAC | Unknown | Cy5 |  | D03 | IAC | Unknown | Cy5
B08 | Leg pneumo | Unknown | FAM | 30.04 | D04 | IAC | Unknown | Cy5
B09 | IAC | Unknown | Cy5 |  | D05 | IAC | Unknown | Cy5
B09 | Leg pneumo | Unknown | FAM | 30.26 | D06 | IAC | Unknown | Cy5
B10 | IAC | Unknown | Cy5 |  | D07 | IAC | Unknown | Cy5
B10 | Leg pneumo | Unknown | FAM | 29.99 | D08 | IAC | Unknown | Cy5
B11 | IAC | Unknown | Cy5 |  | D09 | IAC | Unknown | Cy5 |  |  |  |  | CONCLUSIONS
B11 | Leg pneumo | Unknown | FAM | 29.96 | D10 | IAC | Unknown | Cy5 |  |  |  |  | IAC only amplify above concentration 1.2e9 copies/ul. Four, 10 fold serial dilutions beyond this fail. And they fail at every background concentration of Leg pneumo suggesting that Leg pneumo primers, probe, DNA does not interfere with IAC reaction. Assuming normal distribution of Cq, error bars represent two sigma (95%) confidence interval (CI). Overlapping error bars among different dilutions of Leg pneumo suggest Leg pneumo does not interfere. No fluorescence in IAC in std tubes 4-7 suggests nonoptimal primer probe concentrations.
B12 | IAC | Unknown | Cy5 |  | D11 | IAC | Unknown | Cy5
B12 | Leg pneumo | Unknown | FAM | 30.37 | D12 | IAC | Unknown | Cy5
C01 | IAC | Unknown | Cy5 | 32.23 | E01 | IAC | Unknown | Cy5 | 32.26 | 32.2567 | 0.13503
C01 | Leg pneumo | Unknown | FAM | 37.79 | E02 | IAC | Unknown | Cy5 | 32.39
C02 | IAC | Unknown | Cy5 | 32.22 | E03 | IAC | Unknown | Cy5 | 32.12
C02 | Leg pneumo | Unknown | FAM | 37.93 | E04 | IAC | Unknown | Cy5 | 36.33 | 36.3933 | 0.11846
C03 | IAC | Unknown | Cy5 | 31.35 | E05 | IAC | Unknown | Cy5 | 36.53
C03 | Leg pneumo | Unknown | FAM | 38.52 | E06 | IAC | Unknown | Cy5 | 36.32
C04 | IAC | Unknown | Cy5 | 36.16 | E07 | IAC | Unknown | Cy5 | 40.87 | 41.2233 | 0.306
C04 | Leg pneumo | Unknown | FAM | 37.74 | E08 | IAC | Unknown | Cy5 | 41.4
C05 | IAC | Unknown | Cy5 | 36.55 | E09 | IAC | Unknown | Cy5 | 41.4
C05 | Leg pneumo | Unknown | FAM | 38.27 | E10 | IAC | Unknown | Cy5
C06 | IAC | Unknown | Cy5 | 36.48 | E11 | IAC | Unknown | Cy5
C06 | Leg pneumo | Unknown | FAM | 38.17 | E12 | IAC | Unknown | Cy5
C07 | IAC | Unknown | Cy5 | 41.42 | F01 | IAC | Unknown | Cy5
C07 | Leg pneumo | Unknown | FAM | 39.64 | F02 | IAC | Unknown | Cy5
C08 | IAC | Unknown | Cy5 | 41.52 | F03 | IAC | Unknown | Cy5
C08 | Leg pneumo | Unknown | FAM | 38.96 | F04 | IAC | Unknown | Cy5
C09 | IAC | Unknown | Cy5 | 40.91 | F05 | IAC | Unknown | Cy5
C09 | Leg pneumo | Unknown | FAM | 39.41 | F06 | IAC | Unknown | Cy5
C10 | IAC | Unknown | Cy5 |  | F07 | IAC | Unknown | Cy5
C10 | Leg pneumo | Unknown | FAM | 38.99 | F08 | IAC | Unknown | Cy5
C11 | IAC | Unknown | Cy5 |  | F09 | IAC | Unknown | Cy5
C11 | Leg pneumo | Unknown | FAM | 39.22 | F10 | IAC | Unknown | Cy5
C12 | IAC | Unknown | Cy5 |  | F11 | IAC | Unknown | Cy5
C12 | Leg pneumo | Unknown | FAM | 39.25 | F12 | IAC | Unknown | Cy5
D01 | IAC | Unknown | Cy5 |  | G01 | IAC | Unknown | Cy5 | 32.92 | 32.8167 | 0.31786
D01 | Leg pneumo | Unknown | FAM | 38.98 | G02 | IAC | Unknown | Cy5 | 33.07
D02 | IAC | Unknown | Cy5 |  | G03 | IAC | Unknown | Cy5 | 32.46
D02 | Leg pneumo | Unknown | FAM | 39.4 | G04 | IAC | Unknown | Cy5 | 36.78 | 36.86 | 0.5246
D03 | IAC | Unknown | Cy5 |  | G05 | IAC | Unknown | Cy5 | 37.42
D03 | Leg pneumo | Unknown | FAM | 38.78 | G06 | IAC | Unknown | Cy5 | 36.38
D04 | IAC | Unknown | Cy5 |  | G07 | IAC | Unknown | Cy5 | 41.43 | 41.3833 | 0.16503
D04 | Leg pneumo | Unknown | FAM | 39.4 | G08 | IAC | Unknown | Cy5 | 41.2
D05 | IAC | Unknown | Cy5 |  | G09 | IAC | Unknown | Cy5 | 41.52
D05 | Leg pneumo | Unknown | FAM | 39.57 | G10 | IAC | Unknown | Cy5
D06 | IAC | Unknown | Cy5 |  | G11 | IAC | Unknown | Cy5
D06 | Leg pneumo | Unknown | FAM | 39.47 | G12 | IAC | Unknown | Cy5
D07 | IAC | Unknown | Cy5 |  | H01 | IAC | Unknown | Cy5
D07 | Leg pneumo | Unknown | FAM | 38.38 | H02 | IAC | Unknown | Cy5
D08 | IAC | Unknown | Cy5 |  | H03 | IAC | Unknown | Cy5
D08 | Leg pneumo | Unknown | FAM | 38.85 | H04 | IAC | Unknown | Cy5
D09 | IAC | Unknown | Cy5 |  | H05 | IAC | Unknown | Cy5
D09 | Leg pneumo | Unknown | FAM | 40.32 | H06 | IAC | Unknown | Cy5
D10 | IAC | Unknown | Cy5 |  | H07 | IAC | Unknown | Cy5
D10 | Leg pneumo | Unknown | FAM | 39.15 | H08 | IAC | Unknown | Cy5
D11 | IAC | Unknown | Cy5 |  | H09 | IAC | Unknown | Cy5
D11 | Leg pneumo | Unknown | FAM | 39.5 | H10 | IAC | Negative | Cy5
D12 | IAC | Unknown | Cy5 |  | H11 | IAC | Negative | Cy5
D12 | Leg pneumo | Unknown | FAM | 39.34 | H12 | IAC | Negative | Cy5
E01 | IAC | Unknown | Cy5 | 32.26 | A01 | Leg pneumo | Unknown | FAM | 28.51 | 28.6167
E01 | Leg pneumo | Unknown | FAM | 40.72 | A02 | Leg pneumo | Unknown | FAM | 28.57
E02 | IAC | Unknown | Cy5 | 32.39 | A03 | Leg pneumo | Unknown | FAM | 28.77
E02 | Leg pneumo | Unknown | FAM | 41.62 | A04 | Leg pneumo | Unknown | FAM | 29.86 | 29.5367
E03 | IAC | Unknown | Cy5 | 32.12 | A05 | Leg pneumo | Unknown | FAM | 29.44
E03 | Leg pneumo | Unknown | FAM | 43.08 | A06 | Leg pneumo | Unknown | FAM | 29.31
E04 | IAC | Unknown | Cy5 | 36.33 | A07 | Leg pneumo | Unknown | FAM | 29.86 | 30.0067
E04 | Leg pneumo | Unknown | FAM | 41.41 | A08 | Leg pneumo | Unknown | FAM | 29.96
E05 | IAC | Unknown | Cy5 | 36.53 | A09 | Leg pneumo | Unknown | FAM | 30.2
E05 | Leg pneumo | Unknown | FAM | 41.93 | A10 | Leg pneumo | Unknown | FAM | 29.96 | 30.0167
E06 | IAC | Unknown | Cy5 | 36.32 | A11 | Leg pneumo | Unknown | FAM | 29.95
E06 | Leg pneumo | Unknown | FAM | 41.53 | A12 | Leg pneumo | Unknown | FAM | 30.14
E07 | IAC | Unknown | Cy5 | 40.87 | B01 | Leg pneumo | Unknown | FAM | 29.81 | 29.82
E07 | Leg pneumo | Unknown | FAM | 41.06 | B02 | Leg pneumo | Unknown | FAM | 29.96
E08 | IAC | Unknown | Cy5 | 41.4 | B03 | Leg pneumo | Unknown | FAM | 29.69
E08 | Leg pneumo | Unknown | FAM | 41.04 | B04 | Leg pneumo | Unknown | FAM | 29.86 | 30.02
E09 | IAC | Unknown | Cy5 | 41.4 | B05 | Leg pneumo | Unknown | FAM | 30.14
E09 | Leg pneumo | Unknown | FAM | 42.76 | B06 | Leg pneumo | Unknown | FAM | 30.06
E10 | IAC | Unknown | Cy5 |  | B07 | Leg pneumo | Unknown | FAM | 30.19 | 30.1633
E10 | Leg pneumo | Unknown | FAM | 42.08 | B08 | Leg pneumo | Unknown | FAM | 30.04
E11 | IAC | Unknown | Cy5 |  | B09 | Leg pneumo | Unknown | FAM | 30.26
E11 | Leg pneumo | Unknown | FAM | 41.71 | B10 | Leg pneumo | Unknown | FAM | 29.99 | 30.1067
E12 | IAC | Unknown | Cy5 |  | B11 | Leg pneumo | Unknown | FAM | 29.96
E12 | Leg pneumo | Unknown | FAM | 42.13 | B12 | Leg pneumo | Unknown | FAM | 30.37
F01 | IAC | Unknown | Cy5 |  | C01 | Leg pneumo | Unknown | FAM | 37.79 | 38.08
F01 | Leg pneumo | Unknown | FAM | 41.74 | C02 | Leg pneumo | Unknown | FAM | 37.93
F02 | IAC | Unknown | Cy5 |  | C03 | Leg pneumo | Unknown | FAM | 38.52
F02 | Leg pneumo | Unknown | FAM | 41.5 | C04 | Leg pneumo | Unknown | FAM | 37.74 | 38.06
F03 | IAC | Unknown | Cy5 |  | C05 | Leg pneumo | Unknown | FAM | 38.27
F03 | Leg pneumo | Unknown | FAM | 41.98 | C06 | Leg pneumo | Unknown | FAM | 38.17
F04 | IAC | Unknown | Cy5 |  | C07 | Leg pneumo | Unknown | FAM | 39.64 | 39.3367
F04 | Leg pneumo | Unknown | FAM | 41.17 | C08 | Leg pneumo | Unknown | FAM | 38.96
F05 | IAC | Unknown | Cy5 |  | C09 | Leg pneumo | Unknown | FAM | 39.41
F05 | Leg pneumo | Unknown | FAM | 41.01 | C10 | Leg pneumo | Unknown | FAM | 38.99 | 39.1533
F06 | IAC | Unknown | Cy5 |  | C11 | Leg pneumo | Unknown | FAM | 39.22
F06 | Leg pneumo | Unknown | FAM | 41.98 | C12 | Leg pneumo | Unknown | FAM | 39.25
F07 | IAC | Unknown | Cy5 |  | D01 | Leg pneumo | Unknown | FAM | 38.98 | 39.0533
F07 | Leg pneumo | Unknown | FAM | 41.08 | D02 | Leg pneumo | Unknown | FAM | 39.4
F08 | IAC | Unknown | Cy5 |  | D03 | Leg pneumo | Unknown | FAM | 38.78
F08 | Leg pneumo | Unknown | FAM | 42.05 | D04 | Leg pneumo | Unknown | FAM | 39.4 | 39.48
F09 | IAC | Unknown | Cy5 |  | D05 | Leg pneumo | Unknown | FAM | 39.57
F09 | Leg pneumo | Unknown | FAM | 42.08 | D06 | Leg pneumo | Unknown | FAM | 39.47
F10 | IAC | Unknown | Cy5 |  | D07 | Leg pneumo | Unknown | FAM | 38.38 | 39.1833
F10 | Leg pneumo | Unknown | FAM | 41.04 | D08 | Leg pneumo | Unknown | FAM | 38.85
F11 | IAC | Unknown | Cy5 |  | D09 | Leg pneumo | Unknown | FAM | 40.32
F11 | Leg pneumo | Unknown | FAM | 41.83 | D10 | Leg pneumo | Unknown | FAM | 39.15 | 39.33
F12 | IAC | Unknown | Cy5 |  | D11 | Leg pneumo | Unknown | FAM | 39.5
F12 | Leg pneumo | Unknown | FAM | 42.39 | D12 | Leg pneumo | Unknown | FAM | 39.34
G01 | IAC | Unknown | Cy5 | 32.92 | E01 | Leg pneumo | Unknown | FAM | 40.72 | 41.8067
G01 | Leg pneumo | Unknown | FAM |  | E02 | Leg pneumo | Unknown | FAM | 41.62
G02 | IAC | Unknown | Cy5 | 33.07 | E03 | Leg pneumo | Unknown | FAM | 43.08
G02 | Leg pneumo | Unknown | FAM |  | E04 | Leg pneumo | Unknown | FAM | 41.41 | 41.6233
G03 | IAC | Unknown | Cy5 | 32.46 | E05 | Leg pneumo | Unknown | FAM | 41.93
G03 | Leg pneumo | Unknown | FAM |  | E06 | Leg pneumo | Unknown | FAM | 41.53
G04 | IAC | Unknown | Cy5 | 36.78 | E07 | Leg pneumo | Unknown | FAM | 41.06 | 41.62
G04 | Leg pneumo | Unknown | FAM |  | E08 | Leg pneumo | Unknown | FAM | 41.04
G05 | IAC | Unknown | Cy5 | 37.42 | E09 | Leg pneumo | Unknown | FAM | 42.76
G05 | Leg pneumo | Unknown | FAM |  | E10 | Leg pneumo | Unknown | FAM | 42.08 | 41.9733
G06 | IAC | Unknown | Cy5 | 36.38 | E11 | Leg pneumo | Unknown | FAM | 41.71
G06 | Leg pneumo | Unknown | FAM |  | E12 | Leg pneumo | Unknown | FAM | 42.13
G07 | IAC | Unknown | Cy5 | 41.43 | F01 | Leg pneumo | Unknown | FAM | 41.74 | 41.74
G07 | Leg pneumo | Unknown | FAM |  | F02 | Leg pneumo | Unknown | FAM | 41.5
G08 | IAC | Unknown | Cy5 | 41.2 | F03 | Leg pneumo | Unknown | FAM | 41.98
G08 | Leg pneumo | Unknown | FAM |  | F04 | Leg pneumo | Unknown | FAM | 41.17 | 41.3867
G09 | IAC | Unknown | Cy5 | 41.52 | F05 | Leg pneumo | Unknown | FAM | 41.01
G09 | Leg pneumo | Unknown | FAM |  | F06 | Leg pneumo | Unknown | FAM | 41.98
G10 | IAC | Unknown | Cy5 |  | F07 | Leg pneumo | Unknown | FAM | 41.08 | 41.7367
G10 | Leg pneumo | Unknown | FAM |  | F08 | Leg pneumo | Unknown | FAM | 42.05
G11 | IAC | Unknown | Cy5 |  | F09 | Leg pneumo | Unknown | FAM | 42.08
G11 | Leg pneumo | Unknown | FAM |  | F10 | Leg pneumo | Unknown | FAM | 41.04 | 41.7533
G12 | IAC | Unknown | Cy5 |  | F11 | Leg pneumo | Unknown | FAM | 41.83
G12 | Leg pneumo | Unknown | FAM |  | F12 | Leg pneumo | Unknown | FAM | 42.39
H01 | IAC | Unknown | Cy5 |  | G01 | Leg pneumo | Unknown | FAM
H01 | Leg pneumo | Unknown | FAM |  | G02 | Leg pneumo | Unknown | FAM
H02 | IAC | Unknown | Cy5 |  | G03 | Leg pneumo | Unknown | FAM
H02 | Leg pneumo | Unknown | FAM |  | G04 | Leg pneumo | Unknown | FAM
H03 | IAC | Unknown | Cy5 |  | G05 | Leg pneumo | Unknown | FAM
H03 | Leg pneumo | Unknown | FAM |  | G06 | Leg pneumo | Unknown | FAM
H04 | IAC | Unknown | Cy5 |  | G07 | Leg pneumo | Unknown | FAM
H04 | Leg pneumo | Unknown | FAM |  | G08 | Leg pneumo | Unknown | FAM
H05 | IAC | Unknown | Cy5 |  | G09 | Leg pneumo | Unknown | FAM
H05 | Leg pneumo | Unknown | FAM |  | G10 | Leg pneumo | Unknown | FAM
H06 | IAC | Unknown | Cy5 |  | G11 | Leg pneumo | Unknown | FAM
H06 | Leg pneumo | Unknown | FAM |  | G12 | Leg pneumo | Unknown | FAM
H07 | IAC | Unknown | Cy5 |  | H01 | Leg pneumo | Unknown | FAM
H07 | Leg pneumo | Unknown | FAM |  | H02 | Leg pneumo | Unknown | FAM
H08 | IAC | Unknown | Cy5 |  | H03 | Leg pneumo | Unknown | FAM
H08 | Leg pneumo | Unknown | FAM |  | H04 | Leg pneumo | Unknown | FAM
H09 | IAC | Unknown | Cy5 |  | H05 | Leg pneumo | Unknown | FAM
H09 | Leg pneumo | Unknown | FAM |  | H06 | Leg pneumo | Unknown | FAM
H10 | IAC | Negative | Cy5 |  | H07 | Leg pneumo | Unknown | FAM
H10 | Leg pneumo | Negative | FAM |  | H08 | Leg pneumo | Unknown | FAM
H11 | IAC | Negative | Cy5 |  | H09 | Leg pneumo | Unknown | FAM
H11 | Leg pneumo | Negative | FAM |  | H10 | Leg pneumo | Negative | FAM
H12 | IAC | Negative | Cy5 |  | H11 | Leg pneumo | Negative | FAM
H12 | Leg pneumo | Negative | FAM |  | H12 | Leg pneumo | Negative | FAM
Note on O1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Legionella pneumophila positive control.
20210422 A few more thoughts
I need to repeat the above data with updated concentrations. The above was created with IAC 250nM F and R and 500nM probe. Only three of the IAC pos control successfully amplified.
I've collected evidence showing that ideal concentration should be about 50nM F and R. Probe should be about 400nM. The result of this is a leftward (downward) shift of Cq using same concentration of IAC pos control.
I'd like to run the above using new F, R, P concentrations and observe what happens to the curve above. Also want to know if Leg pneumo at the same concentrations unaffects IAC amplification.

## plate2_setup
Process
Used the same process for making std curve.
The std_curve mastermix had Legionella pneumophila positive control doped in.
Used stds 6, 9 , 12 from Leg pneumo.
20210422
selected new F, R and probe concentration: 50, 50, 400 respectively.
Using Mmix with no IAC, no water. R. made 4/16.
testing same standards as in plate1. Expect to see decrease Cq for all detectable standards.
mastermix |  |  |  | per 3 | per 8 sets of 3 | per 4 sets of 24 | Component | Desired |  | Per 20 uL |  | Multiple Reactions |  |  | Lot
Component | Desired | Units | per 20 ul | 3.42 |  |  | DotAF | 0.4 | uM | 0.08 | uL | 29.4 | uL | Forward 1
mastermix |  |  | 9.1556 | 31.312 |  | 1054.72 | DotAR | 0.4 | uM | 0.08 | uL | 29.4 | uL | Reverse 1
water |  |  | 5.8444 | 19.988 |  | 673.28 | Leg pne DotA MGB probe | 0.5 | uM | 0.1 | uL | 36.75 | uL | Probe 1
TOTAL |  |  | 15 | 51.3 | 410.4 | 1728 | IAC Forward | 0.25 | uM | 0 | uL | 18.375 | uL | Forward 1
[7 empty] | IAC Reverse | 0.25 | uM | 0 | uL | 18.375 | uL | Reverse 1
each mastermix has their own conc of Legionella. 4 sets of 24 so 4 diff conc of Leg stds can be used. | [6 empty] | IAC Probe RED (CY5) | 0.5 | uM | 0 | uL | 36.75 | uL
[5 empty] | per 8 sets of 3 | overage = 1.17 | Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL | 0 | uL
Component | Desired | Units | per 20 ul | *3 | 24 | 28.08 | dNTP Mix | 0.25 | mM | 0.32 | uL | 117.6 | uL | dNTP Mix
mastermix |  |  | 9.1556 | 27.4667 | 219.7333 | 257.088 | AptaTaq | 2.5 | Units | 0.05 | uL | 18.375 | uL | AptaTaq
IAC_F | 50 |  | 0.1 | 0.3 | 2.4 | 2.808
IAC_R | 50 |  | 0.1 | 0.3 | 2.4 | 2.808
IAC_P | 400 |  | 0.8 | 2.4 | 19.2 | 22.464
Leg std DNA |  |  | 2 | 6 | 48 | 56.16 | KAPA3G | 1 | Units | 0.033333 | uL | 12.25 | uL
water |  |  | 5.8444 | 17.5333 | 140.2667 | 164.112 | Mg2+ | 3.5 | mM | 0.07 | uL | 25.725 | uL | Mg2+
TOTAL |  |  | 18 | 54 | 432 | 505.44 | Trehalose | 10 | Percent | 4 | uL | 1470 | uL | Trehalose
[7 empty] | BSA | 0.5 | mg/mL | 0.2 | uL | 73.5 | uL | BSA
[7 empty] | PCR Buffer 10X | 1 | x | 2 | uL | 735 | uL | PCR Buffer 10X
[7 empty] | 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 816.6667 | uL | Ficoll
[7 empty] | Dextran | 0 | Percent | 0 | uL | 0 | uL
[7 empty] | Water |  |  |  | uL | 2074.3333 | uL | Water
[9 empty] | Total: | 9.1556 | uL | 5512.5 | uL
Note on D21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel did not add water to this mastermix.

## plate2_data
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | CT_Mean | std_6 |  |  | std_9 |  |  | std_12 |  |  | water
A01 | IAC | Unknown | Cy5 | 26.73 | A01 | IAC | Unknown | Cy5 | 26.73 | 26.49 | IAC Tube | Copy Number/ul | Ct_mean | IAC Tube | Copy Number/ul | Ct_mean | IAC Tube | Copy Number/ul | Ct_mean | IAC Tube | Copy Number/ul | Ct_mean
A01 | Legpneumo | Unknown | FAM | 30.59 | A02 | IAC | Unknown | Cy5 | 26.58 |  | std_1 | 60220000000 | 26.49 | std_1 | 60220000000 |  | std_1 | 60220000000 | 27.68 | std_1 | 60220000000
A02 | IAC | Unknown | Cy5 | 26.58 | A03 | IAC | Unknown | Cy5 | 26.16 |  | std_2 | 12044000000 | 31.5867 | std_2 | 12044000000 | 31.2067 | std_2 | 12044000000 |  | std_2 | 12044000000 | 31.8533
A02 | Legpneumo | Unknown | FAM | 30.08 | A04 | IAC | Unknown | Cy5 | 31.64 | 31.5867 | std_3 | 1204400000 | 43.27 | std_3 | 1204400000 | 41.0667 | std_3 | 1204400000 | 39.2833 | std_3 | 1204400000 | 38.09
A03 | IAC | Unknown | Cy5 | 26.16 | A05 | IAC | Unknown | Cy5 | 31.72 |  | std_4 | 120440000 |  | std_4 | 120440000 |  | std_4 | 120440000 |  | std_4 | 120440000 | 43.355
A03 | Legpneumo | Unknown | FAM | 30.23 | A06 | IAC | Unknown | Cy5 | 31.4 |  | std_5 | 12044000 |  | std_5 | 12044000 |  | std_5 | 12044000 |  | std_5 | 12044000
A04 | IAC | Unknown | Cy5 | 31.64 | A07 | IAC | Unknown | Cy5 | 43.18 | 43.27 | std_6 | 1204400 |  | std_6 | 1204400 |  | std_6 | 1204400 |  | std_6 | 1204400
A04 | Legpneumo | Unknown | FAM | 29.95 | A08 | IAC | Unknown | Cy5 | 42.69 |  | std_7 | 120440 |  | std_7 | 120440 |  | std_7 | 120440 |  | std_7 | 120440
A05 | IAC | Unknown | Cy5 | 31.72 | A09 | IAC | Unknown | Cy5 | 43.94 |  | std_8 | 12044 |  | std_8 | 12044 |  | std_8 | 12044 |  | std_8 | 12044
A05 | Legpneumo | Unknown | FAM | 29.97 | A10 | IAC | Unknown | Cy5
A06 | IAC | Unknown | Cy5 | 31.4 | A11 | IAC | Unknown | Cy5
A06 | Legpneumo | Unknown | FAM | 29.96 | A12 | IAC | Unknown | Cy5
A07 | IAC | Unknown | Cy5 | 43.18 | B01 | IAC | Unknown | Cy5
A07 | Legpneumo | Unknown | FAM | 29.85 | B02 | IAC | Unknown | Cy5
A08 | IAC | Unknown | Cy5 | 42.69 | B03 | IAC | Unknown | Cy5
A08 | Legpneumo | Unknown | FAM | 30.13 | B04 | IAC | Unknown | Cy5
A09 | IAC | Unknown | Cy5 | 43.94 | B05 | IAC | Unknown | Cy5
A09 | Legpneumo | Unknown | FAM | 30.46 | B06 | IAC | Unknown | Cy5
A10 | IAC | Unknown | Cy5 |  | B07 | IAC | Unknown | Cy5
A10 | Legpneumo | Unknown | FAM | 30.42 | B08 | IAC | Unknown | Cy5
A11 | IAC | Unknown | Cy5 |  | B09 | IAC | Unknown | Cy5
A11 | Legpneumo | Unknown | FAM | 29.97 | B10 | IAC | Unknown | Cy5
A12 | IAC | Unknown | Cy5 |  | B11 | IAC | Unknown | Cy5
A12 | Legpneumo | Unknown | FAM | 30.25 | B12 | IAC | Unknown | Cy5
B01 | IAC | Unknown | Cy5 |  | C01 | IAC | Unknown | Cy5
B01 | Legpneumo | Unknown | FAM | 29.88 | C02 | IAC | Unknown | Cy5
B02 | IAC | Unknown | Cy5 |  | C03 | IAC | Unknown | Cy5
B02 | Legpneumo | Unknown | FAM | 30.4 | C04 | IAC | Unknown | Cy5 | 31.12 | 31.2067
B03 | IAC | Unknown | Cy5 |  | C05 | IAC | Unknown | Cy5 | 31.43
B03 | Legpneumo | Unknown | FAM | 29.98 | C06 | IAC | Unknown | Cy5 | 31.07
B04 | IAC | Unknown | Cy5 |  | C07 | IAC | Unknown | Cy5 | 41.29 | 41.0667
B04 | Legpneumo | Unknown | FAM | 29.76 | C08 | IAC | Unknown | Cy5 | 41.6
B05 | IAC | Unknown | Cy5 |  | C09 | IAC | Unknown | Cy5 | 40.31
B05 | Legpneumo | Unknown | FAM | 30.11 | C10 | IAC | Unknown | Cy5
B06 | IAC | Unknown | Cy5 |  | C11 | IAC | Unknown | Cy5
B06 | Legpneumo | Unknown | FAM | 30.36 | C12 | IAC | Unknown | Cy5
B07 | IAC | Unknown | Cy5 |  | D01 | IAC | Unknown | Cy5
B07 | Legpneumo | Unknown | FAM | 30.24 | D02 | IAC | Unknown | Cy5
B08 | IAC | Unknown | Cy5 |  | D03 | IAC | Unknown | Cy5
B08 | Legpneumo | Unknown | FAM | 29.81 | D04 | IAC | Unknown | Cy5
B09 | IAC | Unknown | Cy5 |  | D05 | IAC | Unknown | Cy5
B09 | Legpneumo | Unknown | FAM | 30.05 | D06 | IAC | Unknown | Cy5
B10 | IAC | Unknown | Cy5 |  | D07 | IAC | Unknown | Cy5
B10 | Legpneumo | Unknown | FAM | 29.75 | D08 | IAC | Unknown | Cy5
B11 | IAC | Unknown | Cy5 |  | D09 | IAC | Unknown | Cy5
B11 | Legpneumo | Unknown | FAM | 29.95 | D10 | IAC | Unknown | Cy5
B12 | IAC | Unknown | Cy5 |  | D11 | IAC | Unknown | Cy5
B12 | Legpneumo | Unknown | FAM | 30.07 | D12 | IAC | Unknown | Cy5
C01 | IAC | Unknown | Cy5 |  | E01 | IAC | Unknown | Cy5 | 27.9 | 27.68
C01 | Legpneumo | Unknown | FAM | 32 | E02 | IAC | Unknown | Cy5 | 27.66
C02 | IAC | Unknown | Cy5 |  | E03 | IAC | Unknown | Cy5 | 27.48
C02 | Legpneumo | Unknown | FAM | 32.2 | E04 | IAC | Unknown | Cy5
C03 | IAC | Unknown | Cy5 |  | E05 | IAC | Unknown | Cy5
C03 | Legpneumo | Unknown | FAM | 33.54 | E06 | IAC | Unknown | Cy5
C04 | IAC | Unknown | Cy5 | 31.12 | E07 | IAC | Unknown | Cy5 | 39.57 | 39.2833
C04 | Legpneumo | Unknown | FAM | 32.5 | E08 | IAC | Unknown | Cy5 | 39.61
C05 | IAC | Unknown | Cy5 | 31.43 | E09 | IAC | Unknown | Cy5 | 38.67
C05 | Legpneumo | Unknown | FAM | 32.55 | E10 | IAC | Unknown | Cy5
C06 | IAC | Unknown | Cy5 | 31.07 | E11 | IAC | Unknown | Cy5
C06 | Legpneumo | Unknown | FAM | 32.48 | E12 | IAC | Unknown | Cy5
C07 | IAC | Unknown | Cy5 | 41.29 | F01 | IAC | Unknown | Cy5
C07 | Legpneumo | Unknown | FAM | 32.55 | F02 | IAC | Unknown | Cy5
C08 | IAC | Unknown | Cy5 | 41.6 | F03 | IAC | Unknown | Cy5
C08 | Legpneumo | Unknown | FAM | 32.47 | F04 | IAC | Unknown | Cy5
C09 | IAC | Unknown | Cy5 | 40.31 | F05 | IAC | Unknown | Cy5
C09 | Legpneumo | Unknown | FAM | 32.74 | F06 | IAC | Unknown | Cy5
C10 | IAC | Unknown | Cy5 |  | F07 | IAC | Unknown | Cy5
C10 | Legpneumo | Unknown | FAM | 32.7 | F08 | IAC | Unknown | Cy5
C11 | IAC | Unknown | Cy5 |  | F09 | IAC | Unknown | Cy5
C11 | Legpneumo | Unknown | FAM | 32.28 | F10 | IAC | Unknown | Cy5
C12 | IAC | Unknown | Cy5 |  | F11 | IAC | Unknown | Cy5
C12 | Legpneumo | Unknown | FAM | 32.48 | F12 | IAC | Unknown | Cy5
D01 | IAC | Unknown | Cy5 |  | G01 | IAC | Unknown | Cy5 |  | #DIV/0!
D01 | Legpneumo | Unknown | FAM | 32.2 | G02 | IAC | Unknown | Cy5
D02 | IAC | Unknown | Cy5 |  | G03 | IAC | Unknown | Cy5
D02 | Legpneumo | Unknown | FAM | 32.48 | G04 | IAC | Unknown | Cy5 | 31.92 | 31.8533
D03 | IAC | Unknown | Cy5 |  | G05 | IAC | Unknown | Cy5 | 32.05
D03 | Legpneumo | Unknown | FAM | 32.45 | G06 | IAC | Unknown | Cy5 | 31.59
D04 | IAC | Unknown | Cy5 |  | G07 | IAC | Unknown | Cy5 | 38.67 | 38.09
D04 | Legpneumo | Unknown | FAM | 32.8 | G08 | IAC | Unknown | Cy5 | 37.83
D05 | IAC | Unknown | Cy5 |  | G09 | IAC | Unknown | Cy5 | 37.77
D05 | Legpneumo | Unknown | FAM | 32.59 | G10 | IAC | Unknown | Cy5 | 44.78 | 43.355
D06 | IAC | Unknown | Cy5 |  | G11 | IAC | Unknown | Cy5
D06 | Legpneumo | Unknown | FAM | 32.82 | G12 | IAC | Unknown | Cy5 | 41.93
D07 | IAC | Unknown | Cy5 |  | H01 | IAC | Unknown | Cy5
D07 | Legpneumo | Unknown | FAM | 31.97 | H02 | IAC | Unknown | Cy5
D08 | IAC | Unknown | Cy5 |  | H03 | IAC | Unknown | Cy5
D08 | Legpneumo | Unknown | FAM | 32.43 | H04 | IAC | Unknown | Cy5
D09 | IAC | Unknown | Cy5 |  | H05 | IAC | Unknown | Cy5
D09 | Legpneumo | Unknown | FAM | 32.3 | H06 | IAC | Unknown | Cy5
D10 | IAC | Unknown | Cy5 |  | H07 | IAC | Unknown | Cy5
D10 | Legpneumo | Unknown | FAM | 32.03 | H08 | IAC | Unknown | Cy5
D11 | IAC | Unknown | Cy5 |  | H09 | IAC | Unknown | Cy5
D11 | Legpneumo | Unknown | FAM | 32.62 | H10 | IAC | Unknown | Cy5
D12 | IAC | Unknown | Cy5 |  | H11 | IAC | Unknown | Cy5
D12 | Legpneumo | Unknown | FAM | 31.5 | H12 | IAC | Unknown | Cy5
E01 | IAC | Unknown | Cy5 | 27.9 | A01 | Legpneumo | Unknown | FAM | 30.59
E01 | Legpneumo | Unknown | FAM | 37.08 | A02 | Legpneumo | Unknown | FAM | 30.08
E02 | IAC | Unknown | Cy5 | 27.66 | A03 | Legpneumo | Unknown | FAM | 30.23
E02 | Legpneumo | Unknown | FAM | 39.09 | A04 | Legpneumo | Unknown | FAM | 29.95
E03 | IAC | Unknown | Cy5 | 27.48 | A05 | Legpneumo | Unknown | FAM | 29.97
E03 | Legpneumo | Unknown | FAM | 38.5 | A06 | Legpneumo | Unknown | FAM | 29.96
E04 | IAC | Unknown | Cy5 |  | A07 | Legpneumo | Unknown | FAM | 29.85
E04 | Legpneumo | Unknown | FAM | 35.1 | A08 | Legpneumo | Unknown | FAM | 30.13
E05 | IAC | Unknown | Cy5 |  | A09 | Legpneumo | Unknown | FAM | 30.46
E05 | Legpneumo | Unknown | FAM | 36.2 | A10 | Legpneumo | Unknown | FAM | 30.42
E06 | IAC | Unknown | Cy5 |  | A11 | Legpneumo | Unknown | FAM | 29.97
E06 | Legpneumo | Unknown | FAM | 35.89 | A12 | Legpneumo | Unknown | FAM | 30.25
E07 | IAC | Unknown | Cy5 | 39.57 | B01 | Legpneumo | Unknown | FAM | 29.88
E07 | Legpneumo | Unknown | FAM | 36.67 | B02 | Legpneumo | Unknown | FAM | 30.4
E08 | IAC | Unknown | Cy5 | 39.61 | B03 | Legpneumo | Unknown | FAM | 29.98
E08 | Legpneumo | Unknown | FAM | 35.65 | B04 | Legpneumo | Unknown | FAM | 29.76
E09 | IAC | Unknown | Cy5 | 38.67 | B05 | Legpneumo | Unknown | FAM | 30.11
E09 | Legpneumo | Unknown | FAM | 36.05 | B06 | Legpneumo | Unknown | FAM | 30.36
E10 | IAC | Unknown | Cy5 |  | B07 | Legpneumo | Unknown | FAM | 30.24
E10 | Legpneumo | Unknown | FAM | 36.07 | B08 | Legpneumo | Unknown | FAM | 29.81
E11 | IAC | Unknown | Cy5 |  | B09 | Legpneumo | Unknown | FAM | 30.05
E11 | Legpneumo | Unknown | FAM | 36.18 | B10 | Legpneumo | Unknown | FAM | 29.75
E12 | IAC | Unknown | Cy5 |  | B11 | Legpneumo | Unknown | FAM | 29.95
E12 | Legpneumo | Unknown | FAM | 36.31 | B12 | Legpneumo | Unknown | FAM | 30.07
F01 | IAC | Unknown | Cy5 |  | C01 | Legpneumo | Unknown | FAM | 32
F01 | Legpneumo | Unknown | FAM | 36.1 | C02 | Legpneumo | Unknown | FAM | 32.2
F02 | IAC | Unknown | Cy5 |  | C03 | Legpneumo | Unknown | FAM | 33.54
F02 | Legpneumo | Unknown | FAM | 36.25 | C04 | Legpneumo | Unknown | FAM | 32.5
F03 | IAC | Unknown | Cy5 |  | C05 | Legpneumo | Unknown | FAM | 32.55
F03 | Legpneumo | Unknown | FAM | 36.64 | C06 | Legpneumo | Unknown | FAM | 32.48
F04 | IAC | Unknown | Cy5 |  | C07 | Legpneumo | Unknown | FAM | 32.55
F04 | Legpneumo | Unknown | FAM | 36.47 | C08 | Legpneumo | Unknown | FAM | 32.47
F05 | IAC | Unknown | Cy5 |  | C09 | Legpneumo | Unknown | FAM | 32.74
F05 | Legpneumo | Unknown | FAM | 36.34 | C10 | Legpneumo | Unknown | FAM | 32.7
F06 | IAC | Unknown | Cy5 |  | C11 | Legpneumo | Unknown | FAM | 32.28
F06 | Legpneumo | Unknown | FAM | 35.45 | C12 | Legpneumo | Unknown | FAM | 32.48
F07 | IAC | Unknown | Cy5 |  | D01 | Legpneumo | Unknown | FAM | 32.2
F07 | Legpneumo | Unknown | FAM | 36.59 | D02 | Legpneumo | Unknown | FAM | 32.48
F08 | IAC | Unknown | Cy5 |  | D03 | Legpneumo | Unknown | FAM | 32.45
F08 | Legpneumo | Unknown | FAM | 35.91 | D04 | Legpneumo | Unknown | FAM | 32.8
F09 | IAC | Unknown | Cy5 |  | D05 | Legpneumo | Unknown | FAM | 32.59
F09 | Legpneumo | Unknown | FAM | 37.47 | D06 | Legpneumo | Unknown | FAM | 32.82
F10 | IAC | Unknown | Cy5 |  | D07 | Legpneumo | Unknown | FAM | 31.97
F10 | Legpneumo | Unknown | FAM | 36.26 | D08 | Legpneumo | Unknown | FAM | 32.43
F11 | IAC | Unknown | Cy5 |  | D09 | Legpneumo | Unknown | FAM | 32.3
F11 | Legpneumo | Unknown | FAM | 36.89 | D10 | Legpneumo | Unknown | FAM | 32.03
F12 | IAC | Unknown | Cy5 |  | D11 | Legpneumo | Unknown | FAM | 32.62
F12 | Legpneumo | Unknown | FAM | 35.33 | D12 | Legpneumo | Unknown | FAM | 31.5
G01 | IAC | Unknown | Cy5 |  | E01 | Legpneumo | Unknown | FAM | 37.08
G01 | Legpneumo | Unknown | FAM | 39.95 | E02 | Legpneumo | Unknown | FAM | 39.09
G02 | IAC | Unknown | Cy5 |  | E03 | Legpneumo | Unknown | FAM | 38.5
G02 | Legpneumo | Unknown | FAM | 39.64 | E04 | Legpneumo | Unknown | FAM | 35.1
G03 | IAC | Unknown | Cy5 |  | E05 | Legpneumo | Unknown | FAM | 36.2
G03 | Legpneumo | Unknown | FAM | 40.83 | E06 | Legpneumo | Unknown | FAM | 35.89
G04 | IAC | Unknown | Cy5 | 31.92 | E07 | Legpneumo | Unknown | FAM | 36.67
G04 | Legpneumo | Unknown | FAM | 39.77 | E08 | Legpneumo | Unknown | FAM | 35.65
G05 | IAC | Unknown | Cy5 | 32.05 | E09 | Legpneumo | Unknown | FAM | 36.05
G05 | Legpneumo | Unknown | FAM | 43.05 | E10 | Legpneumo | Unknown | FAM | 36.07
G06 | IAC | Unknown | Cy5 | 31.59 | E11 | Legpneumo | Unknown | FAM | 36.18
G06 | Legpneumo | Unknown | FAM | 41.58 | E12 | Legpneumo | Unknown | FAM | 36.31
G07 | IAC | Unknown | Cy5 | 38.67 | F01 | Legpneumo | Unknown | FAM | 36.1
G07 | Legpneumo | Unknown | FAM | 40.23 | F02 | Legpneumo | Unknown | FAM | 36.25
G08 | IAC | Unknown | Cy5 | 37.83 | F03 | Legpneumo | Unknown | FAM | 36.64
G08 | Legpneumo | Unknown | FAM | 43.02 | F04 | Legpneumo | Unknown | FAM | 36.47
G09 | IAC | Unknown | Cy5 | 37.77 | F05 | Legpneumo | Unknown | FAM | 36.34
G09 | Legpneumo | Unknown | FAM | 42.54 | F06 | Legpneumo | Unknown | FAM | 35.45
G10 | IAC | Unknown | Cy5 | 44.78 | F07 | Legpneumo | Unknown | FAM | 36.59
G10 | Legpneumo | Unknown | FAM |  | F08 | Legpneumo | Unknown | FAM | 35.91
G11 | IAC | Unknown | Cy5 |  | F09 | Legpneumo | Unknown | FAM | 37.47
G11 | Legpneumo | Unknown | FAM | 42.23 | F10 | Legpneumo | Unknown | FAM | 36.26
G12 | IAC | Unknown | Cy5 | 41.93 | F11 | Legpneumo | Unknown | FAM | 36.89
G12 | Legpneumo | Unknown | FAM | 41.12 | F12 | Legpneumo | Unknown | FAM | 35.33
H01 | IAC | Unknown | Cy5 |  | G01 | Legpneumo | Unknown | FAM | 39.95
H01 | Legpneumo | Unknown | FAM | 41.16 | G02 | Legpneumo | Unknown | FAM | 39.64
H02 | IAC | Unknown | Cy5 |  | G03 | Legpneumo | Unknown | FAM | 40.83
H02 | Legpneumo | Unknown | FAM | 39.6 | G04 | Legpneumo | Unknown | FAM | 39.77
H03 | IAC | Unknown | Cy5 |  | G05 | Legpneumo | Unknown | FAM | 43.05
H03 | Legpneumo | Unknown | FAM |  | G06 | Legpneumo | Unknown | FAM | 41.58
H04 | IAC | Unknown | Cy5 |  | G07 | Legpneumo | Unknown | FAM | 40.23
H04 | Legpneumo | Unknown | FAM |  | G08 | Legpneumo | Unknown | FAM | 43.02
H05 | IAC | Unknown | Cy5 |  | G09 | Legpneumo | Unknown | FAM | 42.54
H05 | Legpneumo | Unknown | FAM |  | G10 | Legpneumo | Unknown | FAM
H06 | IAC | Unknown | Cy5 |  | G11 | Legpneumo | Unknown | FAM | 42.23
H06 | Legpneumo | Unknown | FAM | 41.5 | G12 | Legpneumo | Unknown | FAM | 41.12
H07 | IAC | Unknown | Cy5 |  | H01 | Legpneumo | Unknown | FAM | 41.16
H07 | Legpneumo | Unknown | FAM | 41.38 | H02 | Legpneumo | Unknown | FAM | 39.6
H08 | IAC | Unknown | Cy5 |  | H03 | Legpneumo | Unknown | FAM
H08 | Legpneumo | Unknown | FAM | 41.02 | H04 | Legpneumo | Unknown | FAM
H09 | IAC | Unknown | Cy5 |  | H05 | Legpneumo | Unknown | FAM
H09 | Legpneumo | Unknown | FAM | 41.84 | H06 | Legpneumo | Unknown | FAM | 41.5
H10 | IAC | Unknown | Cy5 |  | H07 | Legpneumo | Unknown | FAM | 41.38
H10 | Legpneumo | Unknown | FAM | 41.43 | H08 | Legpneumo | Unknown | FAM | 41.02
H11 | IAC | Unknown | Cy5 |  | H09 | Legpneumo | Unknown | FAM | 41.84
H11 | Legpneumo | Unknown | FAM | 40.8 | H10 | Legpneumo | Unknown | FAM | 41.43
H12 | IAC | Unknown | Cy5 |  | H11 | Legpneumo | Unknown | FAM | 40.8
H12 | Legpneumo | Unknown | FAM | 41.83 | H12 | Legpneumo | Unknown | FAM | 41.83
Conclusions
This plate's data is much worse than plate 1's data. Are the non-fluorescent wells better explained by chemistry or by an additional freeze-thaw cycle? Was there a systemic error? Is moving to 50, 50nM for F, R introduce stochasticity that wasn't observed in F,R and P matrix data?
IAC Cq with std_1 appears reduced, but the cost may be greater stochastic non-fluorescence.
I don't observe any patterns. Even IAC at high concentration fails.
Next Steps:
Compare fluorescence, non-fluorescence of two mastermixes
IAC(HK): 50, 50, 400
IAC(Lspp): 250, 250, 500
20210430
Right to exclude this data. Something happened with the wells or reactions or control that I didn't notice. Glad I performed more reaction with "LO..." and "HI..." those experiments have been informative.

## matrix_setup
Creating a primer matrix from concentrations
| USER_INPUTS | Value | Note
| orig_F_conc | 10 | What is the starting F primer concentration? (in uM) |  |  |  | orig_F_conc = 10, # What is the starting F primer concentration? (in uM)
| orig_R_conc | 10 | What is the starting R primer concentration? (in uM) |  |  |  | orig_R_conc = 10, # What is the starting R primer concentration? (in uM)
| orig_P_conc | 10 | What is probe starting concentration? (in uM) |  |  |  | orig_P_conc = 10, # What is probe starting concentration? (in uM)
| std_F_conc | 250 | What is the constant F primer concentration for standards? This should be guess or from literature. (in nM) |  |  |  | std_F_conc = 250, # What is the constant F primer concentration for standards? This should be guess or from literature. (in nM)
| std_R_cond | 250 | What is the constant R primer concentration for standards? This should be guess or from literature. (in nM) |  |  |  | std_R_cond = 250, # What is the constant R primer concentration for standards? This should be guess or from literature. (in nM)
| P_conc | 500 | What is the probe concentration? This should be constant throughout experiment. |  |  |  | P_conc = 500, # What is the probe concentration? This should be constant throughout experiment.
| tot_stds | 21 | How many wells of standards will be run? (in # wells) |  |  |  | tot_stds = 21, # How many wells of standards will be run? (in # wells)
| tot_NTCs | 3 | How many wells of NTCs will be run? (in # wells) |  |  |  | tot_NTCs = 3, # How many wells of NTCs will be run? (in # wells)
| tot_samp | 72 | How many samples with varying conditions will be run? (in # wells) |  |  |  | tot_samp = 72, # How many samples with varying conditions will be run? (in # wells)
| rxn_base | 8.86 | Everything in PCR buffer (Mg2+, dNTPs, polymerase, enhancers, stabilizers, sugars, etc.). (in uL) |  |  |  | rxn_base = 8.86, # Everything in PCR buffer (Mg2+, dNTPs, polymerase, enhancers, stabilizers, sugars, etc.). (in uL)
| tot_rxn_vol | 20 | What is total volume of PCR reaction in plate/tube? (in ul) |  |  |  | tot_rxn_vol = 20, # What is total volume of PCR reaction in plate/tube? (in ul)
| F_concs | [50, 100, 200, 400, 600, 800] | What are the F primer concentrations? (a list in nM) |  |  |  | F_concs = [50, 100, 200, 400, 600, 800], # What are the F primer concentrations? (a list in nM)
| R_concs | [50, 100, 200, 400, 600, 800] | What are the R primer concentrations? (a list in nM) |  |  |  | R_concs = [50, 100, 200, 400, 600, 800], # What are the R primer concentrations? (a list in nM)
| F_50 | 50 | F primer at lowest concentration. (in nM) |  |  |  | F_50 = 50, # F primer at lowest concentration. (in nM)
| F_100 | 100 | F primer at concentration. (in nM) |  |  |  | F_100 = 100, # F primer at concentration. (in nM)
| F_200 | 200 | F primer at concentration. (in nM) |  |  |  | F_200 = 200, # F primer at concentration. (in nM)
| F_400 | 400 | F primer at concentration. (in nM) |  |  |  | F_400 = 400, # F primer at concentration. (in nM)
| F_600 | 600 | F primer at concentration. (in nM) |  |  |  | F_600 = 600, # F primer at concentration. (in nM)
| F_800 | 800 | F primer at highest concentration. (in nM) |  |  |  | F_800 = 800, # F primer at highest concentration. (in nM)
| R_50 | 50 | R primer at lowest concentration. (in nM) |  |  |  | R_50 = 50, # R primer at lowest concentration. (in nM)
| R_100 | 100 | R primer at concentration. (in nM) |  |  |  | R_100 = 100, # R primer at concentration. (in nM)
| R_200 | 200 | R primer at concentration. (in nM) |  |  |  | R_200 = 200, # R primer at concentration. (in nM)
| R_400 | 400 | R primer at concentration. (in nM) |  |  |  | R_400 = 400, # R primer at concentration. (in nM)
| R_600 | 600 | R primer at concentration. (in nM) |  |  |  | R_600 = 600, # R primer at concentration. (in nM)
| R_800 | 800 | R primer at highest concentration. (in nM) |  |  |  | R_800 = 800, # R primer at highest concentration. (in nM)
| dna_per_rxn | 2 | How much standard, positive control or NTC to add per well. |  |  |  | dna_per_rxn = 2, # How much standard, positive control or NTC to add per well.
| std_NTC_reps | 3 | How many standard and NTC replicates? (int) |  |  |  | std_NTC_reps = 3, # How many standard and NTC replicates? (int)
| R_reps | 12 | How many wells will use R primer at particular concentration? (int) |  |  |  | R_reps = 12, # How many wells will use R primer at particular concentration? (int)
| F_reps | 12 | How many wells will use F primer at particular concentration? (int) |  |  |  | F_reps = 12, # How many wells will use F primer at particular concentration? (int)
| F_int_vol | 100 | What is the volume of F intermediate primer in new tube? (in ul) |  |  |  | F_int_vol = 100, # What is the volume of F intermediate primer in new tube? (in ul)
| percent_waste | 0.2 | What percentage waste? (decimal) |  |  |  | percent_waste = # What percentage waste? (decimal)
| sN_mix_waste_offset | 0.025 | How much percent_waste offset should sN_mix use? This calculated as percent_waste-sN_mix_overage = percent_waste for sN_mix_overage e.g. (20-7=13%) Should not be 0 otherwise offset = percent_waste. (decimal) |  |  |  | sN_mix_waste_offset = # How much percent_waste offset should sN_mix use? This calculated as percent_waste-sN_mix_overage = percent_waste for sN_mix_overage e.g. (20-7=13%) Should not be 0 otherwise offset = percent_waste. (decimal)
| R_mix_waste_offset | 0.1 | How much percent_waste offset should R_mix use? This calculated as percent_waste-R_mix_overage = percent_waste for R_mix_overage e.g. (20-=13%) If 0, then offset = percent_waste. (decimal) |  |  |  | R_mix_waste_offset = # How much percent_waste offset should R_mix use? This calculated as percent_waste-R_mix_overage = percent_waste for R_mix_overage e.g. (20-=13%) If 0, then offset = percent_waste. (decimal)
| std_NTC_waste_offset | 0.028 | How much percent_waste offset should std_NTC use? (decimal) |  |  |  | std_NTC_waste_offset = # How much percent_waste offset should std_NTC use? (decimal)
| bpw_waste_offset | 0.03 | How much percent_waste offset should bpw_waste use? (decimal) |  |  |  | bpw_waste_offset = # How much percent_waste offset should bpw_waste use? (decimal)
CALCS
| tot_rxns | 96 | Calc what is total # rxns. (int) |  |  |  | tot_rxns = # Calc what is total # rxns. (int)
| tot_sds_NTC | 24 | Calc number of standards and nontemplate controls. (int) |  |  |  | tot_sds_NTC = # Calc number of standards and nontemplate controls. (int)
| rxn_vol_no_dna | 18 | Calc what is volume of PCR rxn with no DNA added. (in ul) |  |  |  | rxn_vol_no_dna = # Calc what is volume of PCR rxn with no DNA added. (in ul)
| P_per_rxn | 1 | Calc this constant probe is needed per pcr reaction. (in ul) |  |  |  | P_per_rxn = # Calc this constant probe is needed per pcr reaction. (in ul)
| max_vol_F_per_rxn | 1.6 | Calc greatest vol added from starting F primer conc to rxn to satisfy max test F conc? (in ul) |  |  |  | max_vol_F_per_rxn = # Calc greatest vol added from starting F primer conc to rxn to satisfy max test F conc? (in ul)
| max_vol_R_per_rxn | 1.6 | Calc greatest vol added from starting R primer conc to rxn to satisfy max test R conc? (in ul) |  |  |  | max_vol_R_per_rxn = # Calc greatest vol added from starting R primer conc to rxn to satisfy max test R conc? (in ul)
| water_per_rxn | 4.94 | Calc this is water added per rxn such that greatest F, R primer conc can be accommodated in total rxn vol. (in ul) |  |  |  | water_per_rxn = # Calc this is water added per rxn such that greatest F, R primer conc can be accommodated in total rxn vol. (in ul)
| std_vol_F_per_rxn | 0.5 | Calc adding this much F primer to each std rxn |  |  |  | std_vol_F_per_rxn = # Calc adding this much F primer to each std rxn
| std_vol_R_per_rxn | 0.5 | Calc adding this much F primer to each std rxn |  |  |  | std_vol_R_per_rxn = # Calc adding this much F primer to each std rxn
| BPW_rxn | 14.8 | Mix base + probe + water (no DNA, F, R primer). (in ul) |  |  |  | BPW_rxn = # Mix base + probe + water (no DNA, F, R primer). (in ul)
| std_woff_per_sN_rxn | 2.2 | Calc how much water to add to offset diff in max vol F,R primer per reaction. (Max-Std)*2 (in ul)
| BPW_sN_rxn | 18 | Mix base + probe + water+F,R primer at std conc+water_offset(no DNA). (in ul) |  |  |  | BPW_sN_rxn = # Mix base + probe + water+F,R primer at std conc+water_offset(no DNA). (in ul)
| BWP_mix_xfer_sN_mix | 417.36 | Calc how much tot_BPW_mix to move to make new tube for stds and NTC. (in ul) |  |  |  | BWP_mix_xfer_sN_mix = # Calc how much tot_BPW_mix to move to make new tube for stds and NTC. (in ul)
| std_woff_add_to_sN_mix | 62.04 | Cal how much water to add to sdd_NTC mmix with waste |  |  |  | std_woff_add_to_sN_mix = # Cal how much water to add to sdd_NTC mmix with waste
| F_add_to_sN_mix | 14.1 | Calc how much F primer to add to std_NTC mmix with waste |  |  |  | F_add_to_sN_mix = # Calc how much F primer to add to std_NTC mmix with waste
| R_add_to_sN_mix | 14.1 | Calc how much R primer to add to std_NTC mmix with waste |  |  |  | R_add_to_sN_mix = # Calc how much R primer to add to std_NTC mmix with waste
| sN_mix_xfer_to_stds_mix | 63.288 | Calc how much sn_mix to aliquot to tube in preparation for std DNA mixing. (in ul) |  |  |  | sN_mix_xfer_to_stds_mix = # Calc how much sn_mix to aliquot to tube in preparation for std DNA mixing. (in ul)
| std_DNA_xfer_to_stds_mix | 7.032 | Calc how much DNA from each standard adding to stds_mix with waste. (in ul) |  |  |  | std_DNA_xfer_to_stds_mix = # Calc how much DNA from each standard adding to stds_mix with waste. (in ul)
| bpw_mix_xfer_bpwd_mix | 1246.752 | Calc how much bpw_mix needed in new tube for # total samples |  |  |  | bpw_mix_xfer_bpwd_mix = # Calc how much bpw_mix needed in new tube for # total samples
| std_DNA_xfer_to_bpwd_mix | 168.48 | Calc how much DNA to add to bpw_mix for template in PCR rxns. |  |  |  | std_DNA_xfer_to_bpwd_mix = # Calc how much DNA to add to bpw_mix for template in PCR rxns.
| bpwd_rxn | 16.8 | Calc vol of bpw reaction + DNA vol |  |  |  | bpwd_rxn = # Calc vol of bpw reaction + DNA vol
| bpwd_mix_xfer_R_mix | 221.76 | Calc how muc bpwd_mix to transfer to each tube for specific R conc. Less than 12% to avoid missing vol. |  |  |  | bpwd_mix_xfer_R_mix = # Calc how muc bpwd_mix to transfer to each tube for specific R conc. Less than 12% to avoid missing vol.
| R_50_rxn | 0.1 | Calc how much R primer at lowest concentration to add per reaction? (e.g. 50nM) (in ul) |  |  |  | R_50_rxn = # Calc how much R primer at lowest concentration to add per reaction? (e.g. 50nM) (in ul)
| R_50_mix | 1.32 | Calc how much R primer to add to R_50 mix. (in ul) |  |  |  | R_50_mix = # Calc how much R primer to add to R_50 mix. (in ul)
| R_50_woff_rxn | 1.5 | Calc how much water offset to add to each rxn. (in ul) |  |  |  | R_50_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_50_woff_mix | 19.8 | Calc how much water offset to add to R_50 mix. (in ul) |  |  |  | R_50_woff_mix = # Calc how much water offset to add to R_50 mix. (in ul)
| R_100_rxn | 0.2 | Calc how much R primer at concentration to add. (in ul) |  |  |  | R_100_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_100_mix | 2.64 | Calc how much R primer to add to add. (in ul) |  |  |  | R_100_mix = # Calc how much R primer to add to add. (in ul)
| R_100_woff_rxn | 1.4 | Calc how much water offset to add to each rxn. (in ul) |  |  |  | R_100_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_100_woff_mix | 18.48 | Calc how much water offset to add to mix. (in ul) |  |  |  | R_100_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_200_rxn | 0.4 | Calc how much R primer at concentration to add. (in ul) |  |  |  | R_200_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_200_mix | 5.28 | Calc how much R primer to add to add. (in ul) |  |  |  | R_200_mix = # Calc how much R primer to add to add. (in ul)
| R_200_woff_rxn | 1.2 | Calc how much water offset to add to each rxn. (in ul) |  |  |  | R_200_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_200_woff_mix | 15.84 | Calc how much water offset to add to mix. (in ul) |  |  |  | R_200_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_400_rxn | 0.8 | Calc how much R primer at concentration to add. (in ul) |  |  |  | R_400_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_400_mix | 10.56 | Calc how much R primer to add to add. (in ul) |  |  |  | R_400_mix = # Calc how much R primer to add to add. (in ul)
| R_400_woff_rxn | 0.8 | Calc how much water offset to add to each rxn. (in ul) |  |  |  | R_400_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_400_woff_mix | 10.56 | Calc how much water offset to add to mix. (in ul) |  |  |  | R_400_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_600_rxn | 1.2 | Calc how much R primer at concentration to add. (in ul) |  |  |  | R_600_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_600_mix | 15.84 | Calc how much R primer to add to add. (in ul) |  |  |  | R_600_mix = # Calc how much R primer to add to add. (in ul)
| R_600_woff_rxn | 0.4 | Calc how much water offset to add to each rxn. (in ul) |  |  |  | R_600_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_600_woff_mix | 5.28 | Calc how much water offset to add to mix. (in ul) |  |  |  | R_600_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_800_rxn | 1.6 | Calc how much R primer at concentration to add. (in ul) |  |  |  | R_800_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_800_mix | 21.12 | Calc how much R primer to add to add. (in ul) |  |  |  | R_800_mix = # Calc how much R primer to add to add. (in ul)
| R_800_woff_rxn | 0 | Calc how much water offset to add to each rxn. (in ul) |  |  |  | R_800_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_800_woff_mix | 0 | Calc how much water offset to add to mix. (in ul) |  |  |  | R_800_woff_mix = # Calc how much water offset to add to mix. (in ul)
| F_50_int_conc | 0.625 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) |  |  |  | F_50_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_100_int_conc | 1.25 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) |  |  |  | F_100_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_200_int_conc | 2.5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) |  |  |  | F_200_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_400_int_conc | 5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) |  |  |  | F_400_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_600_int_conc | 7.5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) |  |  |  | F_600_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_800_int_conc | 10 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) |  |  |  | F_800_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_50_int_primer | 6.25 | What amount of primer should be added to generate int F primer conc? (in ul) |  |  |  | F_50_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_50_int_water | 93.75 | What amount of water should be added to generate int F primer conc? (in ul) |  |  |  | F_50_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_100_int_primer | 12.5 | What amount of primer should be added to generate int F primer conc? (in ul) |  |  |  | F_100_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_100_int_water | 87.5 | What amount of water should be added to generate int F primer conc? (in ul) |  |  |  | F_100_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_200_int_primer | 25 | What amount of primer should be added to generate int F primer conc? (in ul) |  |  |  | F_200_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_200_int_water | 75 | What amount of water should be added to generate int F primer conc? (in ul) |  |  |  | F_200_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_400_int_primer | 50 | What amount of primer should be added to generate int F primer conc? (in ul) |  |  |  | F_400_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_400_int_water | 50 | What amount of water should be added to generate int F primer conc? (in ul) |  |  |  | F_400_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_600_int_primer | 75 | What amount of primer should be added to generate int F primer conc? (in ul) |  |  |  | F_600_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_600_int_water | 25 | What amount of water should be added to generate int F primer conc? (in ul) |  |  |  | F_600_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_800_int_primer | 100 | What amount of primer should be added to generate int F primer conc? (in ul) |  |  |  | F_800_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_800_int_water | 0 | What amount of water should be added to generate int F primer conc? (in ul) |  |  |  | F_800_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| R_mix_rxn | 18.4 | What is the rxn vol that should be aliquoted to plate wells after R primer addition? (in ul) |  |  |  | R_mix_rxn = # What is the rxn vol that should be aliquoted to plate wells after R primer addition? (in ul)
MIXES
| BPW_mix_tot | 1704.96 | Mix = base + probe + water (no DNA, F, R primer)*96*waste. (in ul) |  |  |  | BPW_mix_tot = # Mix = base + probe + water (no DNA, F, R primer)*96*waste. (in ul)
| sN_mix_tot | 507.6 | Mix = base + probe + water + F,R primers at std conc + water offset (no DNA) * number stds_NTC * waste |  |  |  | sN_mix_tot = # Mix = base + probe + water + F,R primers at std conc + water offset (no DNA) * number stds_NTC * waste
| bpwd_mix_tot | 1415.232 | Mix = base + probe + DNA (no F, R primer) |  |  |  | bpwd_mix_tot = # Mix = base + probe + DNA (no F, R primer)
| 50_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) |  |  |  | 50_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 100_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) |  |  |  | 100_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 200_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) |  |  |  | 200_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 400_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) |  |  |  | 400_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 600_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) |  |  |  | 600_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 800_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) |  |  |  | 800_mix = # Mix = base + probe + DNA+ R primer (No F primer)
Mastermix ordered from Raquel
| Component | Desired |  | Per 20 uL |  | Multiple Reactions |  | Confirming Volumes and offsets
| DotAF | 0.4 | uM | 0.08 | uL | 29.4 |  | BPW_mix | 1704.96 |  | bpwd_mix_tot | 1415.232 |  | F_10uM use
| DotAR | 0.4 | uM | 0.08 | uL | 29.4 |  | aliquoted (transferred (xfer)) to two mixes |  |  | bpwd_mix_xfer_R_mix | 221.76
| Leg pne DotA MGB probe | 0.5 | uM | 0.1 | uL | 36.75 |  | BWP_mix_xfer_sN_mix | 417.36 |  | # transfers | 6
| IAC Forward | 0.25 | uM | 0 | uL | 0 |  | bpwd_mix | 1246.752 |  |  | 1330.56
| IAC Reverse | 0.25 | uM | 0 | uL | 0 |  |  | 1664.112
| IAC Probe RED (CY5) | 0.5 | uM | 0 | uL | 0
| Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL | 0 |  | sN_mix has other volumes added and then gets aliquoted to 8 tubes containing standards
| dNTP Mix | 0.25 | mM | 0.32 | uL | 117.6 |  | sN_mix_tot | 507.6
| AptaTaq | 2.5 | Units | 0.05 | uL | 18.375 |  | sN_mix_xfer_to_stds_mix | 63.288
| KAPA3G | 1 | Units | 0.033333 | uL | 12.25 |  | * 8 tubes | 506.304
| Mg2+ | 3.5 | mM | 0.07 | uL | 25.725
| Trehalose | 10 | Percent | 4 | uL | 1470
| BSA | 0.5 | mg/mL | 0.2 | uL | 73.5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc
| PCR Buffer 10X | 1 | x | 2 | uL | 735 | uM | ul | nM | ul
| 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 816.6667 | C1 | V1 | C2 | V2
| Dextran | 0 | Percent | 0 | uL | 0 | 0.625 | 1.6 | 50 | 20
| Water |  |  | 0 | uL | 0 | 1.25 | 1.6 | 100 | 20
[7 empty] | 2.5 | 1.6 | 200 | 20
|  |  | Total: | 9.1556 | uL | 3364.6667 | 5 | 1.6 | 400 | 20
[7 empty] | 7.5 | 1.6 | 600 | 20
[7 empty] | 10 | 1.6 | 800 | 20
| tube:"bpw_mix" | 24 wells for std curve and NTCs
| Component | Desired | Units | PROBE RXN | 115.2 |  | Creating F primer intermediate concentrations from 10uM in 1.5mL tube
| mmix |  |  | 9.1556 | 1054.72 |  | uM | ul | uM | ul
| Probe from 10uM | 500nM |  | 1 | 115.2 |  | C1 | V1 | C2 | V2 |  | 10uM N2_Fwd | Water
| Leg pneumo std_6 |  |  | 2 | 230.4 |  | 10 | 6.25 | 0.625 | 100 |  | 6.25 | 93.75
| water |  |  | 2.6444 | 304.64 |  | 10 | 12.5 | 1.25 | 100 |  | 12.5 | 87.5
| TOTAL |  |  | 14.8 | 1704.96 |  | 10 | 25 | 2.5 | 100 |  | 25 | 75
[7 empty] | 10 | 50 | 5 | 100 |  | 50 | 50
[7 empty] | 10 | 75 | 7.5 | 100 |  | 75 | 25
[7 empty] | 10 | 100 | 10 | 100 |  | 100 | 0
| tube:"sN_mix" | 24 wells for std curve and NTCs
| Component | Desired | Units | PROBE RXN | 27.6
| mmix |  |  | 14.8 | 408.48
| F_primer@10uM |  |  | 0.5 | 13.8
| R_primer@10uM |  |  | 0.5 | 13.8
| water |  |  | 2.2 | 60.72
| TOTAL |  |  | 18 | 496.8
| tube:"bpwd_mix" | 72 wells for samples
| Component | Desired | Units | PROBE RXN | 84.24
| mmix |  |  | 14.8 | 1246.752
| F_primer@10uM |  |  | 0 | 0
| R_primer@10uM |  |  | 0 | 0
| water |  |  | 0 | 0
| DNA (from std_1) for |  |  | 2 | 168.48
| TOTAL |  |  | 16.8 | 1415.232
tube:"50_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 50 | nm | 0.1 | 1.32
| water |  |  | 1.5 | 19.8
| TOTAL |  |  | 18.4 | 242.88
tube:"100_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 100 | nm | 0.2 | 2.64
| water |  |  | 1.4 | 18.48
| TOTAL |  |  | 18.4 | 242.88
tube:"200_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 200 | nm | 0.4 | 5.28
| water |  |  | 1.2 | 15.84
| TOTAL |  |  | 18.4 | 242.88
tube:"400_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 400 | nm | 0.8 | 10.56
| water |  |  | 0.4 | 5.28
| TOTAL |  |  | 18 | 237.6
tube:"600_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 600 | nm | 1.2 | 15.84
| water |  |  | 0.4 | 5.28
| TOTAL |  |  | 18.4 | 242.88
tube:"800_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 800 | nm | 1.6 | 21.12
| water |  |  | 0 | 0
| TOTAL |  |  | 18.4 | 242.88

## matrix_data
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | Ct_AVG |  | [F] | [R] | [F]_[R] | Ct_mean | TrendAvg | TrendStDev |  | [F] | [R] | [F]_[R] | Ct_mean | TrendAvg | TrendStDev
A01 | IAC | Unknown | Cy5 | 27.41 | A01 | IAC | Unknown | Cy5 | 27.41 | 27.575 |  | 50 | 50 | 50_50 | 27.575 | 33.0025 | 5.029 |  | 50 | 50 | 50_50 | 27.575 | 27.2987 | 0.18472
A01 | Lpneumo | Unknown | FAM | 27.96 | A02 | IAC | Unknown | Cy5 | 27.74 |  |  | 100 | 50 | 100_50 | 30.66 |  |  |  | 50 | 100 | 50_100 | 27.205
A02 | IAC | Unknown | Cy5 | 27.74 | A03 | IAC | Unknown | Cy5 | 29.99 | 30.66 |  | 200 | 50 | 200_50 | 34.57 |  |  |  | 50 | 200 | 50_200 | 27.225
A02 | Lpneumo | Unknown | FAM | 28.54 | A04 | IAC | Unknown | Cy5 | 31.33 |  |  | 400 | 50 | 400_50 | 39.205 |  |  |  | 50 | 400 | 50_400 | 27.19
A03 | IAC | Unknown | Cy5 | 29.99 | A05 | IAC | Unknown | Cy5 | 34.17 | 34.57 |  | 600 | 50 | 600_50 | 0 |  |  |  | 50 | 600 | 50_600 | 27.115
A03 | Lpneumo | Unknown | FAM | 29.13 | A06 | IAC | Unknown | Cy5 | 34.97 |  |  | 800 | 50 | 800_50 | 0 |  |  |  | 50 | 800 | 50_800 | 26.345
A04 | IAC | Unknown | Cy5 | 31.33 | A07 | IAC | Unknown | Cy5 | 38.92 | 39.205 |  | 50 | 100 | 50_100 | 27.205 | 35.5033 | 6.4209 |  | 100 | 50 | 100_50 | 30.66 | 29.3417 | 0.86554
A04 | Lpneumo | Unknown | FAM | 29.44 | A08 | IAC | Unknown | Cy5 | 39.49 |  |  | 100 | 100 | 100_100 | 29.08 |  |  |  | 100 | 100 | 100_100 | 29.08
A05 | IAC | Unknown | Cy5 | 34.17 | A09 | IAC | Unknown | Cy5 |  |  |  | 200 | 100 | 200_100 | 34.315 |  |  |  | 100 | 200 | 100_200 | 29.41
A05 | Lpneumo | Unknown | FAM | 33.58 | A10 | IAC | Unknown | Cy5 |  |  |  | 400 | 100 | 400_100 | 38.26 |  |  |  | 100 | 400 | 100_400 | 29.53
A06 | IAC | Unknown | Cy5 | 34.97 | A11 | IAC | Unknown | Cy5 |  |  |  | 600 | 100 | 600_100 | 41.58 |  |  |  | 100 | 600 | 100_600 | 29.405
A06 | Lpneumo | Unknown | FAM | 34.77 | A12 | IAC | Unknown | Cy5 |  |  |  | 800 | 100 | 800_100 | 42.58 |  |  |  | 100 | 800 | 100_800 | 27.965
A07 | IAC | Unknown | Cy5 | 38.92 | B01 | IAC | Unknown | Cy5 | 27.31 | 27.205 |  | 50 | 200 | 50_200 | 27.225 | 35.1192 | 5.9003 |  | 200 | 50 | 200_50 | 34.57 | 34.0317 | 0.38168
A07 | Lpneumo | Unknown | FAM | 43.71 | B02 | IAC | Unknown | Cy5 | 27.1 |  |  | 100 | 200 | 100_200 | 29.41 |  |  |  | 200 | 100 | 200_100 | 34.315
A08 | IAC | Unknown | Cy5 | 39.49 | B03 | IAC | Unknown | Cy5 | 28.98 | 29.08 |  | 200 | 200 | 200_200 | 34.115 |  |  |  | 200 | 200 | 200_200 | 34.115
A08 | Lpneumo | Unknown | FAM |  | B04 | IAC | Unknown | Cy5 | 29.18 |  |  | 400 | 200 | 400_200 | 37.825 |  |  |  | 200 | 400 | 200_400 | 33.745
A09 | IAC | Unknown | Cy5 |  | B05 | IAC | Unknown | Cy5 | 34.53 | 34.315 |  | 600 | 200 | 600_200 | 41.1 |  |  |  | 200 | 600 | 200_600 | 33.525
A09 | Lpneumo | Unknown | FAM |  | B06 | IAC | Unknown | Cy5 | 34.1 |  |  | 800 | 200 | 800_200 | 41.04 |  |  |  | 200 | 800 | 200_800 | 33.92
A10 | IAC | Unknown | Cy5 |  | B07 | IAC | Unknown | Cy5 | 38.35 | 38.26 |  | 50 | 400 | 50_400 | 27.19 | 34.8942 | 5.7409 |  | 400 | 50 | 400_50 | 39.205 | 38.0242 | 0.71484
A10 | Lpneumo | Unknown | FAM |  | B08 | IAC | Unknown | Cy5 | 38.17 |  |  | 100 | 400 | 100_400 | 29.53 |  |  |  | 400 | 100 | 400_100 | 38.26
A11 | IAC | Unknown | Cy5 |  | B09 | IAC | Unknown | Cy5 | 42.4 | 41.58 |  | 200 | 400 | 200_400 | 33.745 |  |  |  | 400 | 200 | 400_200 | 37.825
A11 | Lpneumo | Unknown | FAM |  | B10 | IAC | Unknown | Cy5 | 40.76 |  |  | 400 | 400 | 400_400 | 37.33 |  |  |  | 400 | 400 | 400_400 | 37.33
A12 | IAC | Unknown | Cy5 |  | B11 | IAC | Unknown | Cy5 |  | 42.58 |  | 600 | 400 | 600_400 | 40.75 |  |  |  | 400 | 600 | 400_600 | 37.29
A12 | Lpneumo | Unknown | FAM |  | B12 | IAC | Unknown | Cy5 | 42.58 |  |  | 800 | 400 | 800_400 | 40.82 |  |  |  | 400 | 800 | 400_800 | 38.235
B01 | IAC | Unknown | Cy5 | 27.31 | C01 | IAC | Unknown | Cy5 | 27.4 | 27.225 |  | 50 | 600 | 50_600 | 27.115 | 34.7525 | 5.7101 |  | 600 | 50 | 600_50 | 0 | 34.1042 | 16.7118
B01 | Lpneumo | Unknown | FAM | 27.98 | C02 | IAC | Unknown | Cy5 | 27.05 |  |  | 100 | 600 | 100_600 | 29.405 |  |  |  | 600 | 100 | 600_100 | 41.58
B02 | IAC | Unknown | Cy5 | 27.1 | C03 | IAC | Unknown | Cy5 | 29.58 | 29.41 |  | 200 | 600 | 200_600 | 33.525 |  |  |  | 600 | 200 | 600_200 | 41.1
B02 | Lpneumo | Unknown | FAM | 28.02 | C04 | IAC | Unknown | Cy5 | 29.24 |  |  | 400 | 600 | 400_600 | 37.29 |  |  |  | 600 | 400 | 600_400 | 40.75
B03 | IAC | Unknown | Cy5 | 28.98 | C05 | IAC | Unknown | Cy5 | 34.17 | 34.115 |  | 600 | 600 | 600_600 | 40.58 |  |  |  | 600 | 600 | 600_600 | 40.58
B03 | Lpneumo | Unknown | FAM | 28.31 | C06 | IAC | Unknown | Cy5 | 34.06 |  |  | 800 | 600 | 800_600 | 40.6 |  |  |  | 600 | 800 | 600_800 | 40.615
B04 | IAC | Unknown | Cy5 | 29.18 | C07 | IAC | Unknown | Cy5 | 37.79 | 37.825 |  | 50 | 800 | 50_800 | 26.345 | 34.5767 | 6.2519 |  | 800 | 50 | 800_50 | 0 | 34.2367 | 16.7906
B04 | Lpneumo | Unknown | FAM | 28.63 | C08 | IAC | Unknown | Cy5 | 37.86 |  |  | 100 | 800 | 100_800 | 27.965 |  |  |  | 800 | 100 | 800_100 | 42.58
B05 | IAC | Unknown | Cy5 | 34.53 | C09 | IAC | Unknown | Cy5 | 41.64 | 41.1 |  | 200 | 800 | 200_800 | 33.92 |  |  |  | 800 | 200 | 800_200 | 41.04
B05 | Lpneumo | Unknown | FAM | 33.91 | C10 | IAC | Unknown | Cy5 | 40.56 |  |  | 400 | 800 | 400_800 | 38.235 |  |  |  | 800 | 400 | 800_400 | 40.82
B06 | IAC | Unknown | Cy5 | 34.1 | C11 | IAC | Unknown | Cy5 | 41.25 | 41.04 |  | 600 | 800 | 600_800 | 40.615 |  |  |  | 800 | 600 | 800_600 | 40.6
B06 | Lpneumo | Unknown | FAM | 32.98 | C12 | IAC | Unknown | Cy5 | 40.83 |  |  | 800 | 800 | 800_800 | 40.38 |  |  |  | 800 | 800 | 800_800 | 40.38
B07 | IAC | Unknown | Cy5 | 38.35 | D01 | IAC | Unknown | Cy5 | 27.25 | 27.19
B07 | Lpneumo | Unknown | FAM | 43.68 | D02 | IAC | Unknown | Cy5 | 27.13
B08 | IAC | Unknown | Cy5 | 38.17 | D03 | IAC | Unknown | Cy5 | 29.8 | 29.53
B08 | Lpneumo | Unknown | FAM | 42.44 | D04 | IAC | Unknown | Cy5 | 29.26
B09 | IAC | Unknown | Cy5 | 42.4 | D05 | IAC | Unknown | Cy5 | 34.13 | 33.745
B09 | Lpneumo | Unknown | FAM |  | D06 | IAC | Unknown | Cy5 | 33.36
B10 | IAC | Unknown | Cy5 | 40.76 | D07 | IAC | Unknown | Cy5 | 37.07 | 37.33
B10 | Lpneumo | Unknown | FAM |  | D08 | IAC | Unknown | Cy5 | 37.59
B11 | IAC | Unknown | Cy5 |  | D09 | IAC | Unknown | Cy5 | 40.97 | 40.75
B11 | Lpneumo | Unknown | FAM |  | D10 | IAC | Unknown | Cy5 | 40.53
B12 | IAC | Unknown | Cy5 | 42.58 | D11 | IAC | Unknown | Cy5 | 41.13 | 40.82
B12 | Lpneumo | Unknown | FAM |  | D12 | IAC | Unknown | Cy5 | 40.51
C01 | IAC | Unknown | Cy5 | 27.4 | E01 | IAC | Unknown | Cy5 | 27.18 | 27.115
C01 | Lpneumo | Unknown | FAM | 28.07 | E02 | IAC | Unknown | Cy5 | 27.05
C02 | IAC | Unknown | Cy5 | 27.05 | E03 | IAC | Unknown | Cy5 | 29.54 | 29.405
C02 | Lpneumo | Unknown | FAM | 28.09 | E04 | IAC | Unknown | Cy5 | 29.27
C03 | IAC | Unknown | Cy5 | 29.58 | E05 | IAC | Unknown | Cy5 | 33.82 | 33.525
C03 | Lpneumo | Unknown | FAM | 28.48 | E06 | IAC | Unknown | Cy5 | 33.23
C04 | IAC | Unknown | Cy5 | 29.24 | E07 | IAC | Unknown | Cy5 | 37.02 | 37.29
C04 | Lpneumo | Unknown | FAM | 28.69 | E08 | IAC | Unknown | Cy5 | 37.56
C05 | IAC | Unknown | Cy5 | 34.17 | E09 | IAC | Unknown | Cy5 | 40.88 | 40.58
C05 | Lpneumo | Unknown | FAM | 33.27 | E10 | IAC | Unknown | Cy5 | 40.28 |  |  |  |  | IDT DNA Data |  |  |  | Spec Data
C06 | IAC | Unknown | Cy5 | 34.06 | E11 | IAC | Unknown | Cy5 | 40.73 | 40.6 |  | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/ul | Conc (uM) | pM | Copy Number | Copy Number/ul | Log(Copy#) | Cq_mean | Cq_stdev
C06 | Lpneumo | Unknown | FAM | 33.19 | E12 | IAC | Unknown | Cy5 | 40.47 |  |  |  | IDT Source | 10 | 10000000 |  |  | 3.15 | 3150000
C07 | IAC | Unknown | Cy5 | 37.79 | F01 | IAC | Unknown | Cy5 | 26.39 | 26.345 |  | Tube 1 | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 0.0315 | 31500 | 18969300000000 | 37938600000 | 11.0808 | 33.2267 | 0.2916
C07 | Lpneumo | Unknown | FAM | 41.03 | F02 | IAC | Unknown | Cy5 | 26.3 |  |  | Tube 2 | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 0.00315 | 3150 | 1896930000000 | 3793860000 | 10.0808 | 38.25 | 0.12767
C08 | IAC | Unknown | Cy5 | 37.86 | F03 | IAC | Unknown | Cy5 | 27.83 | 27.965 |  | Tube 3 | std_2.1 | 0.005 | 5000 | 3.011e+12 | 6022000000 | 0.001575 | 1575 | 948465000000 | 1896930000 | 9.7797 | 39.47 | 0.088882
C08 | Lpneumo | Unknown | FAM | 41.85 | F04 | IAC | Unknown | Cy5 | 28.1 |  |  | Tube 4 | std_2.2 | 0.0025 | 2500 | 1.5055e+12 | 3011000000 | 0.0007875 | 787.5 | 474232500000 | 948465000 | 9.4787 | 40.5267 | 0.17474
C09 | IAC | Unknown | Cy5 | 41.64 | F05 | IAC | Unknown | Cy5 | 34.44 | 33.92 |  | Tube 5 | std_2.3 | 0.00125 | 1250 | 752750000000.0001 | 1505500000 | 0.00039375 | 393.75 | 237116250000 | 474232500 | 9.1777 | 41.0233 | 0.43593
C09 | Lpneumo | Unknown | FAM |  | F06 | IAC | Unknown | Cy5 | 33.4 |  |  | Tube 6 | std_3 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 0.000315 | 315 | 189693000000 | 379386000 | 9.0808 | 43.52 | 0.21284
C10 | IAC | Unknown | Cy5 | 40.56 | F07 | IAC | Unknown | Cy5 | 38.68 | 38.235 |  | Tube 7 | std_3.1 | 0.0005 | 500 | 301100000000.0001 | 602200000 | 0.0001575 | 157.5 | 94846500000 | 189693000
C10 | Lpneumo | Unknown | FAM |  | F08 | IAC | Unknown | Cy5 | 37.79
C11 | IAC | Unknown | Cy5 | 41.25 | F09 | IAC | Unknown | Cy5 | 40.94 | 40.615 | [11 empty] | m | b
C11 | Lpneumo | Unknown | FAM |  | F10 | IAC | Unknown | Cy5 | 40.29 | [12 empty] | -4.6144 | 84.4642
C12 | IAC | Unknown | Cy5 | 40.83 | F11 | IAC | Unknown | Cy5 | 40.42 | 40.38
C12 | Lpneumo | Unknown | FAM |  | F12 | IAC | Unknown | Cy5 | 40.34 | [12 empty] | PCR Efficiency | 64.71%
D01 | IAC | Unknown | Cy5 | 27.25 | G01 | IAC | Standard | Cy5 | 32.89 | 33.2267 | 0.2916
D01 | Lpneumo | Unknown | FAM | 28.02 | G02 | IAC | Standard | Cy5 | 33.39
D02 | IAC | Unknown | Cy5 | 27.13 | G03 | IAC | Standard | Cy5 | 33.4
D02 | Lpneumo | Unknown | FAM | 27.93 | G04 | IAC | Standard | Cy5 | 38.14 | 38.25 | 0.12767
D03 | IAC | Unknown | Cy5 | 29.8 | G05 | IAC | Standard | Cy5 | 38.39
D03 | Lpneumo | Unknown | FAM | 28.97 | G06 | IAC | Standard | Cy5 | 38.22
D04 | IAC | Unknown | Cy5 | 29.26 | G07 | IAC | Standard | Cy5 | 39.37 | 39.47 | 0.088882
D04 | Lpneumo | Unknown | FAM | 28.79 | G08 | IAC | Standard | Cy5 | 39.5
D05 | IAC | Unknown | Cy5 | 34.13 | G09 | IAC | Standard | Cy5 | 39.54
D05 | Lpneumo | Unknown | FAM | 33.14 | G10 | IAC | Standard | Cy5 | 40.38 | 40.5267 | 0.17474
D06 | IAC | Unknown | Cy5 | 33.36 | G11 | IAC | Standard | Cy5 | 40.72
D06 | Lpneumo | Unknown | FAM | 32.72 | G12 | IAC | Standard | Cy5 | 40.48
D07 | IAC | Unknown | Cy5 | 37.07 | H01 | IAC | Standard | Cy5 | 41.28 | 41.0233 | 0.43593
D07 | Lpneumo | Unknown | FAM | 38.84 | H02 | IAC | Standard | Cy5 | 41.27
D08 | IAC | Unknown | Cy5 | 37.59 | H03 | IAC | Standard | Cy5 | 40.52
D08 | Lpneumo | Unknown | FAM | 40.84 | H04 | IAC | Standard | Cy5 | 43.56 | 43.52 | 0.21284
D09 | IAC | Unknown | Cy5 | 40.97 | H05 | IAC | Standard | Cy5 | 43.29
D09 | Lpneumo | Unknown | FAM |  | H06 | IAC | Standard | Cy5 | 43.71
D10 | IAC | Unknown | Cy5 | 40.53 | H07 | IAC | Standard | Cy5 |  | 43.52 | #DIV/0!
D10 | Lpneumo | Unknown | FAM |  | H08 | IAC | Standard | Cy5 | 43.52
D11 | IAC | Unknown | Cy5 | 41.13 | H09 | IAC | Standard | Cy5
D11 | Lpneumo | Unknown | FAM |  | H10 | IAC | Negative | Cy5 |  | #DIV/0! | #DIV/0!
D12 | IAC | Unknown | Cy5 | 40.51 | H11 | IAC | Negative | Cy5
D12 | Lpneumo | Unknown | FAM |  | H12 | IAC | Negative | Cy5
E01 | IAC | Unknown | Cy5 | 27.18 | A01 | Lpneumo | Unknown | FAM | 27.96 | 28.25
E01 | Lpneumo | Unknown | FAM | 28.14 | A02 | Lpneumo | Unknown | FAM | 28.54
E02 | IAC | Unknown | Cy5 | 27.05 | A03 | Lpneumo | Unknown | FAM | 29.13 | 29.285
E02 | Lpneumo | Unknown | FAM | 28.1 | A04 | Lpneumo | Unknown | FAM | 29.44
E03 | IAC | Unknown | Cy5 | 29.54 | A05 | Lpneumo | Unknown | FAM | 33.58 | 34.175
E03 | Lpneumo | Unknown | FAM | 29.38 | A06 | Lpneumo | Unknown | FAM | 34.77
E04 | IAC | Unknown | Cy5 | 29.27 | A07 | Lpneumo | Unknown | FAM | 43.71 | 43.71
E04 | Lpneumo | Unknown | FAM | 28.82 | A08 | Lpneumo | Unknown | FAM
E05 | IAC | Unknown | Cy5 | 33.82 | A09 | Lpneumo | Unknown | FAM |  | #DIV/0!
E05 | Lpneumo | Unknown | FAM | 33.34 | A10 | Lpneumo | Unknown | FAM
E06 | IAC | Unknown | Cy5 | 33.23 | A11 | Lpneumo | Unknown | FAM |  | #DIV/0!
E06 | Lpneumo | Unknown | FAM | 32.62 | A12 | Lpneumo | Unknown | FAM
E07 | IAC | Unknown | Cy5 | 37.02 | B01 | Lpneumo | Unknown | FAM | 27.98 | 28
E07 | Lpneumo | Unknown | FAM | 39.15 | B02 | Lpneumo | Unknown | FAM | 28.02
E08 | IAC | Unknown | Cy5 | 37.56 | B03 | Lpneumo | Unknown | FAM | 28.31 | 28.47
E08 | Lpneumo | Unknown | FAM | 39.94 | B04 | Lpneumo | Unknown | FAM | 28.63
E09 | IAC | Unknown | Cy5 | 40.88 | B05 | Lpneumo | Unknown | FAM | 33.91 | 33.445
E09 | Lpneumo | Unknown | FAM |  | B06 | Lpneumo | Unknown | FAM | 32.98
E10 | IAC | Unknown | Cy5 | 40.28 | B07 | Lpneumo | Unknown | FAM | 43.68 | 43.06
E10 | Lpneumo | Unknown | FAM |  | B08 | Lpneumo | Unknown | FAM | 42.44
E11 | IAC | Unknown | Cy5 | 40.73 | B09 | Lpneumo | Unknown | FAM |  | #DIV/0!
E11 | Lpneumo | Unknown | FAM |  | B10 | Lpneumo | Unknown | FAM
E12 | IAC | Unknown | Cy5 | 40.47 | B11 | Lpneumo | Unknown | FAM |  | #DIV/0!
E12 | Lpneumo | Unknown | FAM |  | B12 | Lpneumo | Unknown | FAM
F01 | IAC | Unknown | Cy5 | 26.39 | C01 | Lpneumo | Unknown | FAM | 28.07 | 28.08
F01 | Lpneumo | Unknown | FAM | 28.4 | C02 | Lpneumo | Unknown | FAM | 28.09
F02 | IAC | Unknown | Cy5 | 26.3 | C03 | Lpneumo | Unknown | FAM | 28.48 | 28.585
F02 | Lpneumo | Unknown | FAM | 28.47 | C04 | Lpneumo | Unknown | FAM | 28.69
F03 | IAC | Unknown | Cy5 | 27.83 | C05 | Lpneumo | Unknown | FAM | 33.27 | 33.23
F03 | Lpneumo | Unknown | FAM | 28.87 | C06 | Lpneumo | Unknown | FAM | 33.19
F04 | IAC | Unknown | Cy5 | 28.1 | C07 | Lpneumo | Unknown | FAM | 41.03 | 41.44
F04 | Lpneumo | Unknown | FAM | 29.31 | C08 | Lpneumo | Unknown | FAM | 41.85
F05 | IAC | Unknown | Cy5 | 34.44 | C09 | Lpneumo | Unknown | FAM |  | #DIV/0!
F05 | Lpneumo | Unknown | FAM | 34.76 | C10 | Lpneumo | Unknown | FAM
F06 | IAC | Unknown | Cy5 | 33.4 | C11 | Lpneumo | Unknown | FAM |  | #DIV/0!
F06 | Lpneumo | Unknown | FAM | 33.36 | C12 | Lpneumo | Unknown | FAM
F07 | IAC | Unknown | Cy5 | 38.68 | D01 | Lpneumo | Unknown | FAM | 28.02 | 27.975
F07 | Lpneumo | Unknown | FAM | 41.73 | D02 | Lpneumo | Unknown | FAM | 27.93
F08 | IAC | Unknown | Cy5 | 37.79 | D03 | Lpneumo | Unknown | FAM | 28.97 | 28.88
F08 | Lpneumo | Unknown | FAM | 42.11 | D04 | Lpneumo | Unknown | FAM | 28.79
F09 | IAC | Unknown | Cy5 | 40.94 | D05 | Lpneumo | Unknown | FAM | 33.14 | 32.93
F09 | Lpneumo | Unknown | FAM |  | D06 | Lpneumo | Unknown | FAM | 32.72
F10 | IAC | Unknown | Cy5 | 40.29 | D07 | Lpneumo | Unknown | FAM | 38.84 | 39.84
F10 | Lpneumo | Unknown | FAM |  | D08 | Lpneumo | Unknown | FAM | 40.84
F11 | IAC | Unknown | Cy5 | 40.42 | D09 | Lpneumo | Unknown | FAM |  | #DIV/0!
F11 | Lpneumo | Unknown | FAM |  | D10 | Lpneumo | Unknown | FAM
F12 | IAC | Unknown | Cy5 | 40.34 | D11 | Lpneumo | Unknown | FAM |  | #DIV/0!
F12 | Lpneumo | Unknown | FAM |  | D12 | Lpneumo | Unknown | FAM
G01 | IAC | Standard | Cy5 | 32.89 | E01 | Lpneumo | Unknown | FAM | 28.14 | 28.12
G01 | Lpneumo | Unknown | FAM | 31.57 | E02 | Lpneumo | Unknown | FAM | 28.1
G02 | IAC | Standard | Cy5 | 33.39 | E03 | Lpneumo | Unknown | FAM | 29.38 | 29.1
G02 | Lpneumo | Unknown | FAM | 33.02 | E04 | Lpneumo | Unknown | FAM | 28.82
G03 | IAC | Standard | Cy5 | 33.4 | E05 | Lpneumo | Unknown | FAM | 33.34 | 32.98
G03 | Lpneumo | Unknown | FAM | 33.91 | E06 | Lpneumo | Unknown | FAM | 32.62
G04 | IAC | Standard | Cy5 | 38.14 | E07 | Lpneumo | Unknown | FAM | 39.15 | 39.545
G04 | Lpneumo | Unknown | FAM | 34.61 | E08 | Lpneumo | Unknown | FAM | 39.94
G05 | IAC | Standard | Cy5 | 38.39 | E09 | Lpneumo | Unknown | FAM |  | #DIV/0!
G05 | Lpneumo | Unknown | FAM | 35.4 | E10 | Lpneumo | Unknown | FAM
G06 | IAC | Standard | Cy5 | 38.22 | E11 | Lpneumo | Unknown | FAM |  | #DIV/0!
G06 | Lpneumo | Unknown | FAM | 35.4 | E12 | Lpneumo | Unknown | FAM
G07 | IAC | Standard | Cy5 | 39.37 | F01 | Lpneumo | Unknown | FAM | 28.4 | 28.435
G07 | Lpneumo | Unknown | FAM | 35.11 | F02 | Lpneumo | Unknown | FAM | 28.47
G08 | IAC | Standard | Cy5 | 39.5 | F03 | Lpneumo | Unknown | FAM | 28.87 | 29.09
G08 | Lpneumo | Unknown | FAM | 35.45 | F04 | Lpneumo | Unknown | FAM | 29.31
G09 | IAC | Standard | Cy5 | 39.54 | F05 | Lpneumo | Unknown | FAM | 34.76 | 34.06
G09 | Lpneumo | Unknown | FAM | 36.61 | F06 | Lpneumo | Unknown | FAM | 33.36
G10 | IAC | Standard | Cy5 | 40.38 | F07 | Lpneumo | Unknown | FAM | 41.73 | 41.92
G10 | Lpneumo | Unknown | FAM | 35.22 | F08 | Lpneumo | Unknown | FAM | 42.11
G11 | IAC | Standard | Cy5 | 40.72 | F09 | Lpneumo | Unknown | FAM |  | #DIV/0!
G11 | Lpneumo | Unknown | FAM | 36.01 | F10 | Lpneumo | Unknown | FAM
G12 | IAC | Standard | Cy5 | 40.48 | F11 | Lpneumo | Unknown | FAM |  | #DIV/0!
G12 | Lpneumo | Unknown | FAM | 36.18 | F12 | Lpneumo | Unknown | FAM
H01 | IAC | Standard | Cy5 | 41.28 | G01 | Lpneumo | Unknown | FAM | 31.57 | 32.8333
H01 | Lpneumo | Unknown | FAM | 34.99 | G02 | Lpneumo | Unknown | FAM | 33.02
H02 | IAC | Standard | Cy5 | 41.27 | G03 | Lpneumo | Unknown | FAM | 33.91
H02 | Lpneumo | Unknown | FAM | 35.79 | G04 | Lpneumo | Unknown | FAM | 34.61 | 35.1367
H03 | IAC | Standard | Cy5 | 40.52 | G05 | Lpneumo | Unknown | FAM | 35.4
H03 | Lpneumo | Unknown | FAM | 33.83 | G06 | Lpneumo | Unknown | FAM | 35.4
H04 | IAC | Standard | Cy5 | 43.56 | G07 | Lpneumo | Unknown | FAM | 35.11 | 35.7233
H04 | Lpneumo | Unknown | FAM | 35.78 | G08 | Lpneumo | Unknown | FAM | 35.45
H05 | IAC | Standard | Cy5 | 43.29 | G09 | Lpneumo | Unknown | FAM | 36.61
H05 | Lpneumo | Unknown | FAM | 36.09 | G10 | Lpneumo | Unknown | FAM | 35.22 | 35.8033
H06 | IAC | Standard | Cy5 | 43.71 | G11 | Lpneumo | Unknown | FAM | 36.01
H06 | Lpneumo | Unknown | FAM | 36.11 | G12 | Lpneumo | Unknown | FAM | 36.18
H07 | IAC | Standard | Cy5 |  | H01 | Lpneumo | Unknown | FAM | 34.99 | 34.87
H07 | Lpneumo | Unknown | FAM | 35.19 | H02 | Lpneumo | Unknown | FAM | 35.79
H08 | IAC | Standard | Cy5 | 43.52 | H03 | Lpneumo | Unknown | FAM | 33.83
H08 | Lpneumo | Unknown | FAM | 35.87 | H04 | Lpneumo | Unknown | FAM | 35.78 | 35.9933
H09 | IAC | Standard | Cy5 |  | H05 | Lpneumo | Unknown | FAM | 36.09
H09 | Lpneumo | Unknown | FAM | 37.27 | H06 | Lpneumo | Unknown | FAM | 36.11
H10 | IAC | Negative | Cy5 |  | H07 | Lpneumo | Unknown | FAM | 35.19 | 36.11
H10 | Lpneumo | Unknown | FAM | 37.96 | H08 | Lpneumo | Unknown | FAM | 35.87
H11 | IAC | Negative | Cy5 |  | H09 | Lpneumo | Unknown | FAM | 37.27
H11 | Lpneumo | Unknown | FAM | 38.14 | H10 | Lpneumo | Unknown | FAM | 37.96 | 38.1367
H12 | IAC | Negative | Cy5 |  | H11 | Lpneumo | Unknown | FAM | 38.14
H12 | Lpneumo | Unknown | FAM | 38.31 | H12 | Lpneumo | Unknown | FAM | 38.31
Note on P61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is expected concentration from the tube.
Note on R61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on S61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul. Reply: assumes copy number is in 1000ul
Note on V61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on W61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul.
Note on O62: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Note, that this is different than micH std or other standards as I usually have 100x dilution from std_0 to std_1. However, this standard is simply 10x dilution.
Note on R62: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Note on V62: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Note on O64: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: made these because standards below 3 fail. These are 1/2 dilutions. Took 200ul and added it to 200ul. Mixed. Removed 200ul and added to next tube.
Note on Y68: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Remove this because it failed. May be useful for calculating OD.
CONCLUSIONS:
Pattern suggests that a lower concentration, 50nM, in the F primer is optimal for amplification. Pattern is pronounced when FWD held constant. This stair-step pattern suggets the lower concentration is optimal at 50nM even when R conc varies between 50 and 800nM. PCR efficiency much less than 90%. This is probably okay as long as it's consistent. Hypothesese for lower efficinecy:
1) primers below Tm
2) Leg pneumo primer probe interference e.g. competition
3) master mix not optimal for extension

## probe_setup
Create a set of 7, 10-fold std dilutions as in previous experiments e.g. Exp800.05
from a mastermix tube containing everything but DNA, probe and primers
remove XXX ul to a new tube for standard creation
Add 400nm F, R primers, 300nM probe and aliquot to new tubes
Add std tube into S intermediate tubes, mix and dispense on plate
from a mastermix tube containing everything but DNA, probe and primers
add DNA from std_5 or other and mix
remove XXX ul to a new tube for lower F,R creation
remove XXX ul to a new tube for mid F, R creation e.g. 400nm
remove XXX ul to a new tube for upper F, R creation
add F, R pimers to lower F, R tube
add F, R, primers to mid F, R tube
add F, R, primers to upper F, R tube
add water to tube for adjusting
Dispense 18.4*4ul from lower, mid, upper tubes to 6 beg wells
While chillin, make probe dilutions similar to F dilutions
add 1.6*4 into beg well, mix, aliquiot to 3 neighboring wells
Creating other F, R primer concentrations around the empirically determined concentrations to test for F,R-probe efficiencies at diff conc.
| Component | Desired |  | Per 20 uL
| DotAF | 0.4 | uM | 0.08 | uL
| DotAR | 0.4 | uM | 0.08 | uL
| Leg pne DotA MGB probe | 0.5 | uM | 0.1 | uL
| IAC Forward | 0.25 | uM | 0 | uL
| IAC Reverse | 0.25 | uM | 0 | uL
| IAC Probe RED (CY5) | 0.5 | uM | 0 | uL
| Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL
| dNTP Mix | 0.25 | mM | 0.32 | uL
| AptaTaq | 2.5 | Units | 0.05 | uL
| KAPA3G | 1 | Units | 0.033333 | uL
| Mg2+ | 3.5 | mM | 0.07 | uL
| Trehalose | 10 | Percent | 4 | uL
| BSA | 0.5 | mg/mL | 0.2 | uL
| PCR Buffer 10X | 1 | x | 2 | uL
| 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL
| Dextran | 0 | Percent | 0 | uL
| Water |  |  | 0 | uL
|  |  | Total: | 9.1556 | uL
| tube:"mix_bw_tot" | 24 wells for std curve and NTCs
| Component | Desired | Units | PROBE RXN | 115.2
| mmix |  |  | 9.1556 | 1054.72
| IAC_F_primer@10uM | 50nM |  |  | 0
| IAC_R_primer@10uM | 50nM |  |  | 0
| probe@10uM | 500nM |  |  | 0
| Leg pneumo "std_6" |  |  | 2 | 230.4
| water |  |  | 3.6444 | 419.84
| TOTAL |  |  | 14.8 | 1704.96
| tube:"sN_mix" | 24 wells for std curve and NTCs
| Component | Desired | Units | PROBE RXN | 27.6
| mmix |  |  | 14.8 | 408.48
| IAC_F_primer@10uM | 50nM |  | 0.1 | 2.76
| IAC_R_primer@10uM | 50nM |  | 0.1 | 2.76
| probe@10uM | 500nM |  | 1 | 27.6
0
| water |  |  | 2 | 55.2
| TOTAL |  |  | 18 | 496.8
| lower |  |  |  |  | mid |  |  |  |  | upper
| Component | Desired | Units | PROBE RXN | 27.6 | Component | Desired | Units | PROBE RXN | 27.6 | Component | Desired | Units | PROBE RXN | 27.6
| mmix |  |  | 14.8 | 408.48 | mmix |  |  | 14.8 | 408.48 | mmix |  |  | 14.8 | 408.48
| F_primer@10uM | 50nM |  | 0.1 | 2.76 | F_primer@10uM | 100nM |  | 0.2 | 5.52 | F_primer@10uM | 200nM |  | 0.4 | 11.04
| R_primer@10uM | 50nM |  | 0.1 | 2.76 | R_primer@10uM | 100nM |  | 0.2 | 5.52 | R_primer@10uM | 200nM |  | 0.4 | 11.04
| probe@10uM |  |  |  | 0 | probe@10uM |  |  |  | 0 | probe@10uM |  |  |  | 0
[5 empty] | 0 |  |  |  |  | 0 |  |  |  |  | 0
| IAC_DNA "std_1" |  |  | 2 | 55.2 | IAC_DNA "std_1" |  |  | 2 | 55.2 | IAC_DNA "std_1" |  |  | 2 | 55.2
| water |  |  | 1.4 | 38.64 | water |  |  | 1.2 | 33.12 | water |  |  | 0.8 | 22.08
| TOTAL |  |  | 18.4 | 507.84 | TOTAL |  |  | 18.4 | 507.84 | TOTAL |  |  | 18.4 | 507.84
What intermediate (int) concentration of probe is needed such that by adding 1.6ul into 20ul tot vol I obtain a final concentration of 50, 100, 200…etc
| uM | ul | nM | ul
| C1 | V1 | C2 | V2 |  |  |  | OT-2 moves this much to each int_tube (Hi, mid, low) | 494.592
| 0.625 | 1.6 | 50 | 20
| 1.25 | 1.6 | 100 | 20
| 2.5 | 1.6 | 200 | 20
| 5 | 1.6 | 400 | 20
| 7.5 | 1.6 | 600 | 20
| 10 | 1.6 | 800 | 20
Creating probe intermediate concentrations from 10uM in 1.5mL tube
| uM | ul | uM | ul
| C1 | V1 | C2 | V2 |  | 10uM N2_Fwd | Water
| 10 | 3.125 | 0.625 | 50 |  | 3.125 | 46.875
| 10 | 6.25 | 1.25 | 50 |  | 6.25 | 43.75
| 10 | 12.5 | 2.5 | 50 |  | 12.5 | 37.5
| 10 | 25 | 5 | 50 |  | 25 | 25
| 10 | 37.5 | 7.5 | 50 |  | 37.5 | 12.5
| 10 | 50 | 10 | 50 |  | 50 | 0
Creating a primer matrix from concentrations
| USER_INPUTS | Value | Note
| orig_F_conc | 10 | What is the starting F primer concentration? (in uM) | [5 empty] | orig_F_conc = # What is the starting F primer concentration? (in uM)
| orig_R_conc | 10 | What is the starting R primer concentration? (in uM) | [5 empty] | orig_R_conc = # What is the starting R primer concentration? (in uM)
| orig_P_conc | 10 | What is probe starting concentration? (in uM) | [5 empty] | orig_P_conc = # What is probe starting concentration? (in uM)
| P_conc | 500 | What is the probe concentration for the stds? | [5 empty] | P_conc = # What is the probe concentration for the stds?
| P_50 | 50 | Probe at lowest conc. | [5 empty] | P_50 = # Probe at lowest conc.
| P_100 | 100 | Probe at conc | [5 empty] | P_100 = # Probe at conc
| P_200 | 200 | Probe at conc | [5 empty] | P_200 = # Probe at conc
| P_400 | 400 | Probe at conc | [5 empty] | P_400 = # Probe at conc
| P_600 | 600 | Probe at conc | [5 empty] | P_600 = # Probe at conc
| P_800 | 800 | Probe at highest conc. | [5 empty] | P_800 = # Probe at highest conc.
| sample_reps | 24 | How many samples at each F, R conc will be run? | [5 empty] | sample_reps = # How many samples at each F, R conc will be run?
| tot_stds | 21 | How many wells of standards will be run? (in # wells) | [5 empty] | tot_stds = # How many wells of standards will be run? (in # wells)
| tot_NTCs | 3 | How many wells of NTCs will be run? (in # wells) | [5 empty] | tot_NTCs = # How many wells of NTCs will be run? (in # wells)
| tot_samp | 72 | How many samples with varying conditions will be run? (in # wells) | [5 empty] | tot_samp = # How many samples with varying conditions will be run? (in # wells)
| rxn_base | 9.16 | Everything in PCR buffer (Mg2+, dNTPs, polymerase, enhancers, stabilizers, sugars, etc.). (in uL) | [5 empty] | rxn_base = # Everything in PCR buffer (Mg2+, dNTPs, polymerase, enhancers, stabilizers, sugars, etc.). (in uL)
| rxn_base_plus_water | 14.8 | This value set from other experiments. | [5 empty] | rxn_base_plus_water = # This value set from other experiments.
| tot_rxn_vol | 20 | What is total volume of PCR reaction in plate/tube? (in ul) | [5 empty] | tot_rxn_vol = # What is total volume of PCR reaction in plate/tube? (in ul)
| F_upp | 200 | F primer at lowest concentration. (in nM) | [5 empty] | F_upp = # F primer at lowest concentration. (in nM)
| F_mid | 100 | What is the constant F primer concentration for standards? This should be guess or from literature or empirically determined. (in nM) | [5 empty] | F_mid = # What is the constant F primer concentration for standards? This should be guess or from literature or empirically determined. (in nM)
| F_low | 50 | F primer at concentration. (in nM) | [5 empty] | F_low = # F primer at concentration. (in nM)
| R_upp | 200 | R primer at lowest concentration. (in nM) | [5 empty] | R_upp = # R primer at lowest concentration. (in nM)
| R_mid | 100 | What is the constant R primer concentration for standards? This should be guess or from literature or empirically determined. (in nM) | [5 empty] | R_mid = # What is the constant R primer concentration for standards? This should be guess or from literature or empirically determined. (in nM)
| R_low | 50 | R primer at concentration. (in nM) | [5 empty] | R_low = # R primer at concentration. (in nM)
| dna_per_rxn | 2 | How much standard, positive control or NTC to add per well. | [5 empty] | dna_per_rxn = # How much standard, positive control or NTC to add per well.
| std_NTC_reps | 3 | How many standard and NTC replicates? (int) | [5 empty] | std_NTC_reps = # How many standard and NTC replicates? (int)
| P_reps | 12 | How many wells will use probe primer at particular concentration? (int) | [5 empty] | P_reps = # How many wells will use probe primer at particular concentration? (int)
| P_vol_rxn | 1.6 | What vol of probe will be added to reaction? (in ul) | [5 empty] | P_vol_rxn = # What vol of probe will be added to reaction? (in ul)
| P_int_vol | 50 | What is the volume of F intermediate primer in new tube? (in ul) | [5 empty] | P_int_vol = # What is the volume of F intermediate primer in new tube? (in ul)
| percent_waste | 0.2 | What percentage waste? (decimal) | [5 empty] | percent_waste = # What percentage waste? (decimal)
| sN_mix_waste_offset | 0.05 | How much percent_waste offset should sN_mix use? This calculated as percent_waste-sN_mix_overage = percent_waste for sN_mix_overage e.g. (20-7=13%) Should not be 0 otherwise offset = percent_waste. (decimal) | [5 empty] | sN_mix_waste_offset = # How much percent_waste offset should sN_mix use? This calculated as percent_waste-sN_mix_overage = percent_waste for sN_mix_overage e.g. (20-7=13%) Should not be 0 otherwise offset = percent_waste. (decimal)
| samp_int_mix_waste_offset | 0.05 | How much percent_waste offset should R_mix use? This calculated as percent_waste-R_mix_overage = percent_waste for R_mix_overage e.g. (20-=13%) If 0, then offset = percent_waste. (decimal) | [5 empty] | samp_int_mix_waste_offset = # How much percent_waste offset should R_mix use? This calculated as percent_waste-R_mix_overage = percent_waste for R_mix_overage e.g. (20-=13%) If 0, then offset = percent_waste. (decimal)
| std_NTC_waste_offset | 0.07 | How much percent_waste offset should std_NTC use? (decimal) | [5 empty] | std_NTC_waste_offset = # How much percent_waste offset should std_NTC use? (decimal)
| mix_samp_XFT_well_waste_offset | 0.08 | How much should offset be from each sample tube, upper, mid, low to the plate wells? Will receive P, be mixed and then aliquoted to 3 wells. | [5 empty] | mix_samp_XFT_well_waste_offset = # How much should offset be from each sample tube, upper, mid, low to the plate wells? Will receive P, be mixed and then aliquoted to 3 wells.
CALCS
| tot_rxns | 96 | Calc what is total # rxns. (int) | [5 empty] | tot_rxns = # Calc what is total # rxns. (int)
| tot_stds_NTC | 24 | Calc number of standards and nontemplate controls. (int) | [5 empty] | tot_stds_NTC = # Calc number of standards and nontemplate controls. (int)
| rxn_vol_no_dna | 18 | Calc what is volume of PCR rxn with no DNA added. (in ul) | [5 empty] | rxn_vol_no_dna = # Calc what is volume of PCR rxn with no DNA added. (in ul)
| mix_bw_tot | 1704.96 | Tube with mix containing base and water. | [5 empty] | mix_bw_tot = # Tube with mix containing base and water.
| mix_bw_XFR_mix_sn | 408.48 | Transfer this amount from MIX_bw to new tube to receive F, R, P for std and NTC prep | [5 empty] | mix_bw_XFR_mix_sn = # Transfer this amount from MIX_bw to new tube to receive F, R, P for std and NTC prep
| std_vol_F_per_rxn | 0.1 | Calc adding this much F primer to each std rxn | [5 empty] | std_vol_F_per_rxn = # Calc adding this much F primer to each std rxn
| std_vol_R_per_rxn | 0.1 | Calc adding this much R primer to each std rxn | [5 empty] | std_vol_R_per_rxn = # Calc adding this much R primer to each std rxn
| std_vol_P_per_rxn | 1 | Calc adding this much Probe to each std rxn | [5 empty] | std_vol_P_per_rxn = # Calc adding this much Probe to each std rxn
| std_vol_water_per_rxn | 2 | How much water is added to the std_ntc reactions. Should be 18ul to receive 2ul DNA. (in ul) | [5 empty] | std_vol_water_per_rxn = # How much water is added to the std_ntc reactions. Should be 18ul to receive 2ul DNA. (in ul)
| std_vol_F_mix | 2.76 | How much F primer to add to MIX_sn? | [5 empty] | std_vol_F_mix = # How much F primer to add to MIX_sn?
| std_vol_R_mix | 2.76 | How much R primer to add to MIX_sn? | [5 empty] | std_vol_R_mix = # How much R primer to add to MIX_sn?
| std_vol_P_mix | 27.6 | How much P primer to add to MIX_sn? | [5 empty] | std_vol_P_mix = # How much P primer to add to MIX_sn?
| std_vol_water_mix | 55.2 | How much water to add to MIX_sn? | [5 empty] | std_vol_water_mix = # How much water to add to MIX_sn?
| mix_sn_tot | 496.8 | Mix containing base+water+F,R,P. Needs to be aliquoted to std_int tubes | [5 empty] | mix_sn_tot = # Mix containing base+water+F,R,P. Needs to be aliquoted to std_int tubes
| mix_sn_XFR_to_std_int | 61.02 | transfer this amount to std_int tubes | [5 empty] | mix_sn_XFR_to_std_int = # transfer this amount to std_int tubes
| std_dna_XFR_to_std_int | 6.78 | transfer this amount DNA to std_int_tubes to mix and aliquot to 3 wells
| mix_bw_XFR_samp_int | 463.68 | transfer this amount from mix_bw to a new sample intermediate tube for diff values of F, R conc. | [5 empty] | mix_bw_XFR_samp_int = # transfer this amount from mix_bw to a new sample intermediate tube for diff values of F, R conc.
| 10x_dna_per_rxn | 2 | Dilute DNA vol added by 10x and use std_4 instead of std_5 to save vol for P and F, R at upp | [5 empty] | 10x_dna_per_rxn = # Dilute DNA vol added by 10x and use std_4 instead of std_5 to save vol for P and F, R at upp
| dna_XFR_samp_int | 55.2 | Add this DNA to a tube and mix before aliquoting to samp_int | [5 empty] | dna_XFR_samp_int = # Add this DNA to a tube and mix before aliquoting to samp_int
| F_mid_rxn | 0.2 | How much F @ mid conc to add to rxn. | [5 empty] | F_mid_rxn = # How much F @ mid conc to add to rxn.
| F_mid_mix | 5.52 | How much F to add to mix | [5 empty] | F_mid_mix = # How much F to add to mix
| R_mid_rxn | 0.2 | How much R @ mid conc to add to rxn. | [5 empty] | R_mid_rxn = # How much R @ mid conc to add to rxn.
| R_mid_mix | 5.52 | How much R to add to mix | [5 empty] | R_mid_mix = # How much R to add to mix
| F_low_rxn | 0.1 | How much F @ low conc to add to rxn. | [5 empty] | F_low_rxn = # How much F @ low conc to add to rxn.
| F_low_mix | 2.76 | How much F to add to mix | [5 empty] | F_low_mix = # How much F to add to mix
| R_low_rxn | 0.1 | How much R @ low conc to add to rxn. | [5 empty] | R_low_rxn = # How much R @ low conc to add to rxn.
| R_low_mix | 2.76 | How much R to add to mix | [5 empty] | R_low_mix = # How much R to add to mix
| F_upp_rxn | 0.4 | How much F @ upp conc to add to rxn. | [5 empty] | F_upp_rxn = # How much F @ upp conc to add to rxn.
| F_upp_mix | 11.04 | How much F to add to mix | [5 empty] | F_upp_mix = # How much F to add to mix
| R_upp_rxn | 0.4 | How much R @ upp conc to add to rxn. | [5 empty] | R_upp_rxn = # How much R @ upp conc to add to rxn.
| R_upp_mix | 11.04 | How much R to add to mix | [5 empty] | R_upp_mix = # How much R to add to mix
| water_upp_rxn | 0.8 | How much water to add to upp rxn | [5 empty] | water_upp_rxn = # How much water to add to upp rxn
| water_upp_mix | 22.08 | How much water to add to upp mix | [5 empty] | water_upp_mix = # How much water to add to upp mix
| water_low_rxn | 1.4 | How much water to add to low rxn | [5 empty] | water_low_rxn = # How much water to add to low rxn
| water_low_mix | 38.64 | How much water to add to low mix | [5 empty] | water_low_mix = # How much water to add to low mix
| water_mid_rxn | 1.2 | how much water to add to mid rxn | [5 empty] | water_mid_rxn = # how much water to add to mid rxn
| water_mid_mix | 33.12 | how much water to add to mid mix | [5 empty] | water_mid_mix = # how much water to add to mid mix
| MIX_upp_samp_int | 563.04 | How much in upper intermediate sample tube? | [5 empty] | MIX_upp_samp_int = # How much in upper intermediate sample tube?
| MIX_mid_samp_int | 563.04 | How much in middle intermediate sample tube? | [5 empty] | MIX_mid_samp_int = # How much in middle intermediate sample tube?
| MIX_low_samp_int | 563.04 | How much in lower intermediate sample tube? | [5 empty] | MIX_low_samp_int = # How much in lower intermediate sample tube?
| mix_samp_XFR_to_well | 82.432 | How much to add to plate wells?
| P_50_int_conc | 0.625 | What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | P_50_int_conc = # What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM)
| P_100_int_conc | 1.25 | What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | P_100_int_conc = # What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM)
| P_200_int_conc | 2.5 | What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | P_200_int_conc = # What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM)
| P_400_int_conc | 5 | What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | P_400_int_conc = # What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM)
| P_600_int_conc | 7.5 | What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | P_600_int_conc = # What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM)
| P_800_int_conc | 10 | What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | P_800_int_conc = # What intermediate (int) concentration of Probe is needed such that by adding 1.6ul I obtain a final concentration of 50, 100, 200…etc in rxn well? (in uM)
| P_50_int_primer | 3.125 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | P_50_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| P_100_int_primer | 6.25 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | P_100_int_primer = # What amount of water should be added to generate int F primer conc? (in ul)
| P_200_int_primer | 12.5 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | P_200_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| P_400_int_primer | 25 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | P_400_int_primer = # What amount of water should be added to generate int F primer conc? (in ul)
| P_600_int_primer | 37.5 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | P_600_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| P_800_int_primer | 50 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | P_800_int_primer = # What amount of water should be added to generate int F primer conc? (in ul)
| P_50_int_water | 46.875 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | P_50_int_water = # What amount of primer should be added to generate int F primer conc? (in ul)
| P_100_int_water | 43.75 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | P_100_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| P_200_int_water | 37.5 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | P_200_int_water = # What amount of primer should be added to generate int F primer conc? (in ul)
| P_400_int_water | 25 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | P_400_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| P_600_int_water | 12.5 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | P_600_int_water = # What amount of primer should be added to generate int F primer conc? (in ul)
| P_800_int_water | 0 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | P_800_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| p_int_XFR_to_well | 7.168 | What amount of probe intermediate to add to bolus in wells?
Note on E69: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Want to add 1.6ul of probe to equal rxn_mix + 1.6ul = 20ul. Consistent with F primer in Exp800.05.
Note on K69: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Want to add 1.6ul of probe to equal rxn_mix + 1.6ul = 20ul. Consistent with F primer in Exp800.05.
Note on Q69: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Want to add 1.6ul of probe to equal rxn_mix + 1.6ul = 20ul. Consistent with F primer in Exp800.05.
Note on E88: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Exp800.05 uses 100ul total volume. But F primer cheaper than probe. Can this be reduced to 50ul total volume?

## probe_data
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | Ct_mean | Ct_stdDev | F_conc | R_conc | P_conc | Name | Ct_mean | Ct_stdev |  | F_conc | R_conc | P_conc | Name | Ct_mean | Ct_stdev
A01 | IAC | Unknown | Cy5 | 35.37 | A01 | IAC | Unknown | Cy5 | 35.37 | 35.595 | 0.22825 | 200 | 200 | 50 | 200_200_50 | 35.595 | 0.22825 |  | 50 | 50 | 50 | 50_50_50 | 30.14 | 0.38236
A01 | Legpneumo | Unknown | FAM | 34.26 | A02 | IAC | Unknown | Cy5 | 35.69 |  |  | 100 | 100 | 50 | 100_100_50 | 33.6075 | 0.39853 |  | 50 | 50 | 100 | 50_50_100 | 29.5725 | 0.57893
A02 | IAC | Unknown | Cy5 | 35.69 | A03 | IAC | Unknown | Cy5 | 35.87 |  |  | 50 | 50 | 50 | 50_50_50 | 30.14 | 0.38236 |  | 50 | 50 | 200 | 50_50_200 | 28.8425 | 0.1372
A02 | Legpneumo | Unknown | FAM | 34.83 | A04 | IAC | Unknown | Cy5 | 35.45 |  |  | 200 | 200 | 100 | 200_200_100 | 34.77 | 0.32373 |  | 50 | 50 | 400 | 50_50_400 | 28.5875 | 0.27897
A03 | IAC | Unknown | Cy5 | 35.87 | A05 | IAC | Unknown | Cy5 | 33.19 | 33.6075 | 0.39853 | 100 | 100 | 100 | 100_100_100 | 32.8425 | 0.19311 |  | 50 | 50 | 600 | 50_50_600 | 29.145 | 0.23685
A03 | Legpneumo | Unknown | FAM | 34.79 | A06 | IAC | Unknown | Cy5 | 33.55 |  |  | 50 | 50 | 100 | 50_50_100 | 29.5725 | 0.57893 |  | 50 | 50 | 800 | 50_50_800 | 30.2125 | 0.39118
A04 | IAC | Unknown | Cy5 | 35.45 | A07 | IAC | Unknown | Cy5 | 34.15 |  |  | 200 | 200 | 200 | 200_200_200 | 34.1525 | 0.23585 |  | 100 | 100 | 50 | 100_100_50 | 33.6075 | 0.39853
A04 | Legpneumo | Unknown | FAM | 34.79 | A08 | IAC | Unknown | Cy5 | 33.54 |  |  | 100 | 100 | 200 | 100_100_200 | 32.15 | 0.19339 |  | 100 | 100 | 100 | 100_100_100 | 32.8425 | 0.19311
A05 | IAC | Unknown | Cy5 | 33.19 | A09 | IAC | Unknown | Cy5 | 29.61 | 30.14 | 0.38236 | 50 | 50 | 200 | 50_50_200 | 28.8425 | 0.1372 |  | 100 | 100 | 200 | 100_100_200 | 32.15 | 0.19339
A05 | Legpneumo | Unknown | FAM | 30.19 | A10 | IAC | Unknown | Cy5 | 30.36 |  |  | 200 | 200 | 400 | 200_200_400 | 34.0825 | 0.48328 |  | 100 | 100 | 400 | 100_100_400 | 32.46 | 0.28012
A06 | IAC | Unknown | Cy5 | 33.55 | A11 | IAC | Unknown | Cy5 | 30.12 |  |  | 100 | 100 | 400 | 100_100_400 | 32.46 | 0.28012 |  | 100 | 100 | 600 | 100_100_600 | 32.5275 | 0.36317
A06 | Legpneumo | Unknown | FAM | 30.18 | A12 | IAC | Unknown | Cy5 | 30.47 |  |  | 50 | 50 | 400 | 50_50_400 | 28.5875 | 0.27897 |  | 100 | 100 | 800 | 100_100_800 | 32.76 | 0.18403
A07 | IAC | Unknown | Cy5 | 34.15 | B01 | IAC | Unknown | Cy5 | 34.29 | 34.77 | 0.32373 | 200 | 200 | 600 | 200_200_600 | 34.155 | 0.59674 |  | 200 | 200 | 50 | 200_200_50 | 35.595 | 0.22825
A07 | Legpneumo | Unknown | FAM | 30.34 | B02 | IAC | Unknown | Cy5 | 34.99 |  |  | 100 | 100 | 600 | 100_100_600 | 32.5275 | 0.36317 |  | 200 | 200 | 100 | 200_200_100 | 34.77 | 0.32373
A08 | IAC | Unknown | Cy5 | 33.54 | B03 | IAC | Unknown | Cy5 | 34.93 |  |  | 50 | 50 | 600 | 50_50_600 | 29.145 | 0.23685 |  | 200 | 200 | 200 | 200_200_200 | 34.1525 | 0.23585
A08 | Legpneumo | Unknown | FAM | 30.26 | B04 | IAC | Unknown | Cy5 | 34.87 |  |  | 200 | 200 | 800 | 200_200_800 | 33.9375 | 0.14127 |  | 200 | 200 | 400 | 200_200_400 | 34.0825 | 0.48328
A09 | IAC | Unknown | Cy5 | 29.61 | B05 | IAC | Unknown | Cy5 | 32.61 | 32.8425 | 0.19311 | 100 | 100 | 800 | 100_100_800 | 32.76 | 0.18403 |  | 200 | 200 | 600 | 200_200_600 | 34.155 | 0.59674
A09 | Legpneumo | Unknown | FAM | 29.5 | B06 | IAC | Unknown | Cy5 | 32.79 |  |  | 50 | 50 | 800 | 50_50_800 | 30.2125 | 0.39118 |  | 200 | 200 | 800 | 200_200_800 | 33.9375 | 0.14127
A10 | IAC | Unknown | Cy5 | 30.36 | B07 | IAC | Unknown | Cy5 | 33.07
A10 | Legpneumo | Unknown | FAM | 29.6 | B08 | IAC | Unknown | Cy5 | 32.9
A11 | IAC | Unknown | Cy5 | 30.12 | B09 | IAC | Unknown | Cy5 | 28.75 | 29.5725 | 0.57893
A11 | Legpneumo | Unknown | FAM | 29.29 | B10 | IAC | Unknown | Cy5 | 29.66
A12 | IAC | Unknown | Cy5 | 30.47 | B11 | IAC | Unknown | Cy5 | 30.1
A12 | Legpneumo | Unknown | FAM | 29.69 | B12 | IAC | Unknown | Cy5 | 29.78
B01 | IAC | Unknown | Cy5 | 34.29 | C01 | IAC | Unknown | Cy5 | 33.85 | 34.1525 | 0.23585
B01 | Legpneumo | Unknown | FAM | 33.83 | C02 | IAC | Unknown | Cy5 | 34.13
B02 | IAC | Unknown | Cy5 | 34.99 | C03 | IAC | Unknown | Cy5 | 34.42
B02 | Legpneumo | Unknown | FAM | 34.53 | C04 | IAC | Unknown | Cy5 | 34.21
B03 | IAC | Unknown | Cy5 | 34.93 | C05 | IAC | Unknown | Cy5 | 31.99 | 32.15 | 0.19339
B03 | Legpneumo | Unknown | FAM | 34.48 | C06 | IAC | Unknown | Cy5 | 31.99
B04 | IAC | Unknown | Cy5 | 34.87 | C07 | IAC | Unknown | Cy5 | 32.38
B04 | Legpneumo | Unknown | FAM | 34.29 | C08 | IAC | Unknown | Cy5 | 32.24
B05 | IAC | Unknown | Cy5 | 32.61 | C09 | IAC | Unknown | Cy5 | 28.98 | 28.8425 | 0.1372
B05 | Legpneumo | Unknown | FAM | 30.23 | C10 | IAC | Unknown | Cy5 | 28.94
B06 | IAC | Unknown | Cy5 | 32.79 | C11 | IAC | Unknown | Cy5 | 28.71
B06 | Legpneumo | Unknown | FAM | 30.45 | C12 | IAC | Unknown | Cy5 | 28.74
B07 | IAC | Unknown | Cy5 | 33.07 | D01 | IAC | Unknown | Cy5 | 33.65 | 34.0825 | 0.48328
B07 | Legpneumo | Unknown | FAM | 30.57 | D02 | IAC | Unknown | Cy5 | 34.02
B08 | IAC | Unknown | Cy5 | 32.9 | D03 | IAC | Unknown | Cy5 | 33.89
B08 | Legpneumo | Unknown | FAM | 30.23 | D04 | IAC | Unknown | Cy5 | 34.77
B09 | IAC | Unknown | Cy5 | 28.75 | D05 | IAC | Unknown | Cy5 | 32.25 | 32.46 | 0.28012
B09 | Legpneumo | Unknown | FAM | 29.38 | D06 | IAC | Unknown | Cy5 | 32.19
B10 | IAC | Unknown | Cy5 | 29.66 | D07 | IAC | Unknown | Cy5 | 32.74
B10 | Legpneumo | Unknown | FAM | 29.3 | D08 | IAC | Unknown | Cy5 | 32.66
B11 | IAC | Unknown | Cy5 | 30.1 | D09 | IAC | Unknown | Cy5 | 28.24 | 28.5875 | 0.27897
B11 | Legpneumo | Unknown | FAM | 30.13 | D10 | IAC | Unknown | Cy5 | 28.52
B12 | IAC | Unknown | Cy5 | 29.78 | D11 | IAC | Unknown | Cy5 | 28.9
B12 | Legpneumo | Unknown | FAM | 29.75 | D12 | IAC | Unknown | Cy5 | 28.69
C01 | IAC | Unknown | Cy5 | 33.85 | E01 | IAC | Unknown | Cy5 | 33.45 | 34.155 | 0.59674
C01 | Legpneumo | Unknown | FAM | 33.76 | E02 | IAC | Unknown | Cy5 | 33.87
C02 | IAC | Unknown | Cy5 | 34.13 | E03 | IAC | Unknown | Cy5 | 34.65
C02 | Legpneumo | Unknown | FAM | 33.98 | E04 | IAC | Unknown | Cy5 | 34.65
C03 | IAC | Unknown | Cy5 | 34.42 | E05 | IAC | Unknown | Cy5 | 32.63 | 32.5275 | 0.36317
C03 | Legpneumo | Unknown | FAM | 34.35 | E06 | IAC | Unknown | Cy5 | 32.01
C04 | IAC | Unknown | Cy5 | 34.21 | E07 | IAC | Unknown | Cy5 | 32.86
C04 | Legpneumo | Unknown | FAM | 34.15 | E08 | IAC | Unknown | Cy5 | 32.61
C05 | IAC | Unknown | Cy5 | 31.99 | E09 | IAC | Unknown | Cy5 | 29.34 | 29.145 | 0.23685
C05 | Legpneumo | Unknown | FAM | 30.21 | E10 | IAC | Unknown | Cy5 | 29.04
C06 | IAC | Unknown | Cy5 | 31.99 | E11 | IAC | Unknown | Cy5 | 28.86
C06 | Legpneumo | Unknown | FAM | 30 | E12 | IAC | Unknown | Cy5 | 29.34
C07 | IAC | Unknown | Cy5 | 32.38 | F01 | IAC | Unknown | Cy5 | 33.79 | 33.9375 | 0.14127
C07 | Legpneumo | Unknown | FAM | 30.3 | F02 | IAC | Unknown | Cy5 | 33.92
C08 | IAC | Unknown | Cy5 | 32.24 | F03 | IAC | Unknown | Cy5 | 33.91
C08 | Legpneumo | Unknown | FAM | 30.26 | F04 | IAC | Unknown | Cy5 | 34.13
C09 | IAC | Unknown | Cy5 | 28.98 | F05 | IAC | Unknown | Cy5 | 32.66 | 32.76 | 0.18403
C09 | Legpneumo | Unknown | FAM | 29.58 | F06 | IAC | Unknown | Cy5 | 32.58
C10 | IAC | Unknown | Cy5 | 28.94 | F07 | IAC | Unknown | Cy5 | 33
C10 | Legpneumo | Unknown | FAM | 29.25 | F08 | IAC | Unknown | Cy5 | 32.8
C11 | IAC | Unknown | Cy5 | 28.71 | F09 | IAC | Unknown | Cy5 | 30.67 | 30.2125 | 0.39118
C11 | Legpneumo | Unknown | FAM | 28.96 | F10 | IAC | Unknown | Cy5 | 29.76
C12 | IAC | Unknown | Cy5 | 28.74 | F11 | IAC | Unknown | Cy5 | 30.36
C12 | Legpneumo | Unknown | FAM | 29.1 | F12 | IAC | Unknown | Cy5 | 30.06
D01 | IAC | Unknown | Cy5 | 33.65 | G01 | IAC | Standard | Cy5 | 30.46 | 30.18 | 0.45078 | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/ul | Conc (uM) | pM | Copy Number | Copy Number/ul | Log(Copy#) | Cq_mean | Cq_stdev
D01 | Legpneumo | Unknown | FAM | 33.47 | G02 | IAC | Standard | Cy5 | 30.42 |  |  |  | IDT Source | 10 | 10000000 |  |  | 3.15 | 3150000
D02 | IAC | Unknown | Cy5 | 34.02 | G03 | IAC | Standard | Cy5 | 29.66 |  |  | Tube 1 | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 0.0315 | 31500 | 18969300000000 | 37938600000 | 11.0808 | 30.18 | 0.45078
D02 | Legpneumo | Unknown | FAM | 33.64 | G04 | IAC | Standard | Cy5 | 34.55 | 34.7133 | 1.4519 | Tube 2 | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 0.00315 | 3150 | 1896930000000 | 3793860000 | 10.0808 | 34.7133 | 1.4519
D03 | IAC | Unknown | Cy5 | 33.89 | G05 | IAC | Standard | Cy5 | 36.24 |  |  | Tube 3 | std_2.1 | 0.005 | 5000 | 3.011e+12 | 6022000000 | 0.001575 | 1575 | 948465000000 | 1896930000 | 9.7797 | 39.16 | 0.151
D03 | Legpneumo | Unknown | FAM | 33.81 | G06 | IAC | Standard | Cy5 | 33.35 |  |  | Tube 4 | std_2.2 | 0.0025 | 2500 | 1.5055e+12 | 3011000000 | 0.0007875 | 787.5 | 474232500000 | 948465000 | 9.4787 | 40.58 | 0.30414
D04 | IAC | Unknown | Cy5 | 34.77 | G07 | IAC | Standard | Cy5 | 39.18 | 39.16 | 0.151 | Tube 5 | std_2.3 | 0.00125 | 1250 | 752750000000.0001 | 1505500000 | 0.00039375 | 393.75 | 237116250000 | 474232500 | 9.1777 | 41.5133 | 0.43822
D04 | Legpneumo | Unknown | FAM | 34.41 | G08 | IAC | Standard | Cy5 | 39 |  |  | Tube 6 | std_3 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 0.000315 | 315 | 189693000000 | 379386000 | 9.0808 | 42.5 | 0.087178
D05 | IAC | Unknown | Cy5 | 32.25 | G09 | IAC | Standard | Cy5 | 39.3 |  |  | Tube 7 | std_3.1 | 0.0005 | 500 | 301100000000.0001 | 602200000 | 0.0001575 | 157.5 | 94846500000 | 189693000
D05 | Legpneumo | Unknown | FAM | 29.79 | G10 | IAC | Standard | Cy5 | 40.43 | 40.58 | 0.30414
D06 | IAC | Unknown | Cy5 | 32.19 | G11 | IAC | Standard | Cy5 | 40.93
D06 | Legpneumo | Unknown | FAM | 30.27 | G12 | IAC | Standard | Cy5 | 40.38
D07 | IAC | Unknown | Cy5 | 32.74 | H01 | IAC | Standard | Cy5 | 41.39 | 41.5133 | 0.43822 | [10 empty] | m | b
D07 | Legpneumo | Unknown | FAM | 30.21 | H02 | IAC | Standard | Cy5 | 41.15 | [12 empty] | -6.3123 | 99.8404
D08 | IAC | Unknown | Cy5 | 32.66 | H03 | IAC | Standard | Cy5 | 42
D08 | Legpneumo | Unknown | FAM | 30.37 | H04 | IAC | Standard | Cy5 | 42.56 | 42.5 | 0.087178 | [10 empty] | PCR Efficiency | 44.02%
D09 | IAC | Unknown | Cy5 | 28.24 | H05 | IAC | Standard | Cy5 | 42.54
D09 | Legpneumo | Unknown | FAM | 28.62 | H06 | IAC | Standard | Cy5 | 42.4
D10 | IAC | Unknown | Cy5 | 28.52 | H07 | IAC | Standard | Cy5 |  | 43.39 | 1.4001
D10 | Legpneumo | Unknown | FAM | 28.95 | H08 | IAC | Standard | Cy5 | 44.38
D11 | IAC | Unknown | Cy5 | 28.9 | H09 | IAC | Standard | Cy5
D11 | Legpneumo | Unknown | FAM | 29.31 | H10 | IAC | Negative | Cy5
D12 | IAC | Unknown | Cy5 | 28.69 | H11 | IAC | Negative | Cy5
D12 | Legpneumo | Unknown | FAM | 29.08 | H12 | IAC | Negative | Cy5
E01 | IAC | Unknown | Cy5 | 33.45 | A01 | Legpneumo | Unknown | FAM | 34.26
E01 | Legpneumo | Unknown | FAM | 32.54 | A02 | Legpneumo | Unknown | FAM | 34.83
E02 | IAC | Unknown | Cy5 | 33.87 | A03 | Legpneumo | Unknown | FAM | 34.79
E02 | Legpneumo | Unknown | FAM | 33.67 | A04 | Legpneumo | Unknown | FAM | 34.79
E03 | IAC | Unknown | Cy5 | 34.65 | A05 | Legpneumo | Unknown | FAM | 30.19
E03 | Legpneumo | Unknown | FAM | 33.95 | A06 | Legpneumo | Unknown | FAM | 30.18
E04 | IAC | Unknown | Cy5 | 34.65 | A07 | Legpneumo | Unknown | FAM | 30.34
E04 | Legpneumo | Unknown | FAM | 34.01 | A08 | Legpneumo | Unknown | FAM | 30.26
E05 | IAC | Unknown | Cy5 | 32.63 | A09 | Legpneumo | Unknown | FAM | 29.5
E05 | Legpneumo | Unknown | FAM | 30.19 | A10 | Legpneumo | Unknown | FAM | 29.6
E06 | IAC | Unknown | Cy5 | 32.01 | A11 | Legpneumo | Unknown | FAM | 29.29
E06 | Legpneumo | Unknown | FAM | 30.3 | A12 | Legpneumo | Unknown | FAM | 29.69
E07 | IAC | Unknown | Cy5 | 32.86 | B01 | Legpneumo | Unknown | FAM | 33.83
E07 | Legpneumo | Unknown | FAM | 30.47 | B02 | Legpneumo | Unknown | FAM | 34.53
E08 | IAC | Unknown | Cy5 | 32.61 | B03 | Legpneumo | Unknown | FAM | 34.48
E08 | Legpneumo | Unknown | FAM | 30.24 | B04 | Legpneumo | Unknown | FAM | 34.29
E09 | IAC | Unknown | Cy5 | 29.34 | B05 | Legpneumo | Unknown | FAM | 30.23
E09 | Legpneumo | Unknown | FAM | 29.62 | B06 | Legpneumo | Unknown | FAM | 30.45
E10 | IAC | Unknown | Cy5 | 29.04 | B07 | Legpneumo | Unknown | FAM | 30.57
E10 | Legpneumo | Unknown | FAM | 29.12 | B08 | Legpneumo | Unknown | FAM | 30.23
E11 | IAC | Unknown | Cy5 | 28.86 | B09 | Legpneumo | Unknown | FAM | 29.38
E11 | Legpneumo | Unknown | FAM | 29.16 | B10 | Legpneumo | Unknown | FAM | 29.3
E12 | IAC | Unknown | Cy5 | 29.34 | B11 | Legpneumo | Unknown | FAM | 30.13
E12 | Legpneumo | Unknown | FAM | 29.4 | B12 | Legpneumo | Unknown | FAM | 29.75
F01 | IAC | Unknown | Cy5 | 33.79 | C01 | Legpneumo | Unknown | FAM | 33.76
F01 | Legpneumo | Unknown | FAM | 33.55 | C02 | Legpneumo | Unknown | FAM | 33.98
F02 | IAC | Unknown | Cy5 | 33.92 | C03 | Legpneumo | Unknown | FAM | 34.35
F02 | Legpneumo | Unknown | FAM | 33.74 | C04 | Legpneumo | Unknown | FAM | 34.15
F03 | IAC | Unknown | Cy5 | 33.91 | C05 | Legpneumo | Unknown | FAM | 30.21
F03 | Legpneumo | Unknown | FAM | 33.7 | C06 | Legpneumo | Unknown | FAM | 30
F04 | IAC | Unknown | Cy5 | 34.13 | C07 | Legpneumo | Unknown | FAM | 30.3
F04 | Legpneumo | Unknown | FAM | 33.97 | C08 | Legpneumo | Unknown | FAM | 30.26
F05 | IAC | Unknown | Cy5 | 32.66 | C09 | Legpneumo | Unknown | FAM | 29.58
F05 | Legpneumo | Unknown | FAM | 30.3 | C10 | Legpneumo | Unknown | FAM | 29.25
F06 | IAC | Unknown | Cy5 | 32.58 | C11 | Legpneumo | Unknown | FAM | 28.96
F06 | Legpneumo | Unknown | FAM | 30 | C12 | Legpneumo | Unknown | FAM | 29.1
F07 | IAC | Unknown | Cy5 | 33 | D01 | Legpneumo | Unknown | FAM | 33.47
F07 | Legpneumo | Unknown | FAM | 30.41 | D02 | Legpneumo | Unknown | FAM | 33.64
F08 | IAC | Unknown | Cy5 | 32.8 | D03 | Legpneumo | Unknown | FAM | 33.81
F08 | Legpneumo | Unknown | FAM | 30.35 | D04 | Legpneumo | Unknown | FAM | 34.41
F09 | IAC | Unknown | Cy5 | 30.67 | D05 | Legpneumo | Unknown | FAM | 29.79
F09 | Legpneumo | Unknown | FAM | 30.82 | D06 | Legpneumo | Unknown | FAM | 30.27
F10 | IAC | Unknown | Cy5 | 29.76 | D07 | Legpneumo | Unknown | FAM | 30.21
F10 | Legpneumo | Unknown | FAM | 29.9 | D08 | Legpneumo | Unknown | FAM | 30.37
F11 | IAC | Unknown | Cy5 | 30.36 | D09 | Legpneumo | Unknown | FAM | 28.62
F11 | Legpneumo | Unknown | FAM | 30.03 | D10 | Legpneumo | Unknown | FAM | 28.95
F12 | IAC | Unknown | Cy5 | 30.06 | D11 | Legpneumo | Unknown | FAM | 29.31
F12 | Legpneumo | Unknown | FAM | 30.02 | D12 | Legpneumo | Unknown | FAM | 29.08
G01 | IAC | Standard | Cy5 | 30.46 | E01 | Legpneumo | Unknown | FAM | 32.54
G01 | Legpneumo | Unknown | FAM | 28.27 | E02 | Legpneumo | Unknown | FAM | 33.67
G02 | IAC | Standard | Cy5 | 30.42 | E03 | Legpneumo | Unknown | FAM | 33.95
G02 | Legpneumo | Unknown | FAM | 28.37 | E04 | Legpneumo | Unknown | FAM | 34.01
G03 | IAC | Standard | Cy5 | 29.66 | E05 | Legpneumo | Unknown | FAM | 30.19
G03 | Legpneumo | Unknown | FAM | 28.5 | E06 | Legpneumo | Unknown | FAM | 30.3
G04 | IAC | Standard | Cy5 | 34.55 | E07 | Legpneumo | Unknown | FAM | 30.47
G04 | Legpneumo | Unknown | FAM | 29.06 | E08 | Legpneumo | Unknown | FAM | 30.24
G05 | IAC | Standard | Cy5 | 36.24 | E09 | Legpneumo | Unknown | FAM | 29.62
G05 | Legpneumo | Unknown | FAM | 29.41 | E10 | Legpneumo | Unknown | FAM | 29.12
G06 | IAC | Standard | Cy5 | 33.35 | E11 | Legpneumo | Unknown | FAM | 29.16
G06 | Legpneumo | Unknown | FAM | 29.45 | E12 | Legpneumo | Unknown | FAM | 29.4
G07 | IAC | Standard | Cy5 | 39.18 | F01 | Legpneumo | Unknown | FAM | 33.55
G07 | Legpneumo | Unknown | FAM | 29 | F02 | Legpneumo | Unknown | FAM | 33.74
G08 | IAC | Standard | Cy5 | 39 | F03 | Legpneumo | Unknown | FAM | 33.7
G08 | Legpneumo | Unknown | FAM | 29.04 | F04 | Legpneumo | Unknown | FAM | 33.97
G09 | IAC | Standard | Cy5 | 39.3 | F05 | Legpneumo | Unknown | FAM | 30.3
G09 | Legpneumo | Unknown | FAM | 29.22 | F06 | Legpneumo | Unknown | FAM | 30
G10 | IAC | Standard | Cy5 | 40.43 | F07 | Legpneumo | Unknown | FAM | 30.41
G10 | Legpneumo | Unknown | FAM | 28.91 | F08 | Legpneumo | Unknown | FAM | 30.35
G11 | IAC | Standard | Cy5 | 40.93 | F09 | Legpneumo | Unknown | FAM | 30.82
G11 | Legpneumo | Unknown | FAM | 29.33 | F10 | Legpneumo | Unknown | FAM | 29.9
G12 | IAC | Standard | Cy5 | 40.38 | F11 | Legpneumo | Unknown | FAM | 30.03
G12 | Legpneumo | Unknown | FAM | 29.68 | F12 | Legpneumo | Unknown | FAM | 30.02
H01 | IAC | Standard | Cy5 | 41.39 | G01 | Legpneumo | Unknown | FAM | 28.27
H01 | Legpneumo | Unknown | FAM | 29.11 | G02 | Legpneumo | Unknown | FAM | 28.37
H02 | IAC | Standard | Cy5 | 41.15 | G03 | Legpneumo | Unknown | FAM | 28.5
H02 | Legpneumo | Unknown | FAM | 29.23 | G04 | Legpneumo | Unknown | FAM | 29.06
H03 | IAC | Standard | Cy5 | 42 | G05 | Legpneumo | Unknown | FAM | 29.41
H03 | Legpneumo | Unknown | FAM | 29.28 | G06 | Legpneumo | Unknown | FAM | 29.45
H04 | IAC | Standard | Cy5 | 42.56 | G07 | Legpneumo | Unknown | FAM | 29
H04 | Legpneumo | Unknown | FAM | 29.2 | G08 | Legpneumo | Unknown | FAM | 29.04
H05 | IAC | Standard | Cy5 | 42.54 | G09 | Legpneumo | Unknown | FAM | 29.22
H05 | Legpneumo | Unknown | FAM | 29.31 | G10 | Legpneumo | Unknown | FAM | 28.91
H06 | IAC | Standard | Cy5 | 42.4 | G11 | Legpneumo | Unknown | FAM | 29.33
H06 | Legpneumo | Unknown | FAM | 29.36 | G12 | Legpneumo | Unknown | FAM | 29.68
H07 | IAC | Standard | Cy5 |  | H01 | Legpneumo | Unknown | FAM | 29.11
H07 | Legpneumo | Unknown | FAM | 28.94 | H02 | Legpneumo | Unknown | FAM | 29.23
H08 | IAC | Standard | Cy5 | 44.38 | H03 | Legpneumo | Unknown | FAM | 29.28
H08 | Legpneumo | Unknown | FAM | 29.11 | H04 | Legpneumo | Unknown | FAM | 29.2
H09 | IAC | Standard | Cy5 |  | H05 | Legpneumo | Unknown | FAM | 29.31
H09 | Legpneumo | Unknown | FAM | 29.67 | H06 | Legpneumo | Unknown | FAM | 29.36
H10 | IAC | Negative | Cy5 |  | H07 | Legpneumo | Unknown | FAM | 28.94
H10 | Legpneumo | Unknown | FAM | 29.2 | H08 | Legpneumo | Unknown | FAM | 29.11
H11 | IAC | Negative | Cy5 |  | H09 | Legpneumo | Unknown | FAM | 29.67
H11 | Legpneumo | Unknown | FAM | 29.24 | H10 | Legpneumo | Unknown | FAM | 29.2
H12 | IAC | Negative | Cy5 |  | H11 | Legpneumo | Unknown | FAM | 29.24
H12 | Legpneumo | Unknown | FAM | 29.36 | H12 | Legpneumo | Unknown | FAM | 29.36
Note on Q75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is expected concentration from the tube.
Note on S75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on T75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul. Reply: assumes copy number is in 1000ul
Note on W75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on X75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul.
Note on P76: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Note, that this is different than micH std or other standards as I usually have 100x dilution from std_0 to std_1. However, this standard is simply 10x dilution.
Note on S76: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Note on W76: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Note on P78: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: made these because standards below 3 fail. These are 1/2 dilutions. Took 200ul and added it to 200ul. Mixed. Removed 200ul and added to next tube.
Note on Z82: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Remove this because it failed. May be useful for calculating OD.
Note: The "create.probe.matrix" program added 2x probe volume to wells. This was a bug that wasn't caught until halfway through program. Error is found in else clause L570. Performing double aspirant volume like in L556. The excess volume was only added to the low F, R intermediate tube. So, in this case, only 50nM volumes affected. Still think data informative. Yes, previous assays (micH, Legionella pneumophila) also affected. Still, trenda data can be helpful.
Conclusions:
CONCLUSIONS
Natural depression around 200-400nM when F,W = 50nM. Actual concentration is more like 400-800nm (see Note below). Current concentration of probe is 500nM. So, use 400nM.

## HI_std_vs_opt
STD |  |  |  |  | OPT |  |  |  |  | Well | Ct_STD | Ct_STD_MEAN | Ct_STD_STDEV | Ct_OPT | Ct_OPT_MEAN | Ct_OPT_STDEV |  | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/ul | Log(Copy Number/ul) | Ct_STD_MEAN | Ct_STD_STDEV | 2*stdev | Ct_OPT_MEAN | Ct_OPT_STDEV | 2*stdev | Diff
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | A01 | 32.85 | 33.0645 | 0.11527 | 24.52 | 24.9145 | 0.090262 |  | Tube 1 | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808 | 33.0645 | 0.11527 | 0.23054 | 24.9145 | 0.090262 | 0.18052 | -8.15
A01 | IAC | Unknown | Cy5 | 32.85 | A01 | IAC | Unknown | Cy5 | 24.52 | A02 | 33.07 |  |  | 24.99 |  |  |  | Tube 2 | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | 37.22 | 0.30486 | 0.60972 | 29.4827 | 0.39241 | 0.78481 | -7.7373
A02 | IAC | Unknown | Cy5 | 33.07 | A02 | IAC | Unknown | Cy5 | 24.99 | A03 | 33.11 |  |  | 24.96 |  |  |  | Tube 3 | std_2.1 | 0.005 | 5000 | 3.011e+12 | 6022000000 | 9.7797 | 39.1589 | 0.55777 | 1.1155 | 31.4436 | 0.24299 | 0.48599 | -7.7153
A03 | IAC | Unknown | Cy5 | 33.11 | A03 | IAC | Unknown | Cy5 | 24.96 | A04 | 33.13 |  |  | 24.88 |  |  |  | Tube 4 | std_2.2 | 0.0025 | 2500 | 1.5055e+12 | 3011000000 | 9.4787 |  |  | 0 | 35.3373 | 0.94022 | 1.8804
A04 | IAC | Unknown | Cy5 | 33.13 | A04 | IAC | Unknown | Cy5 | 24.88 | A05 | 33.06 |  |  | 24.93 |  |  |  | Tube 5 | std_2.3 | 0.00125 | 1250 | 752750000000.0001 | 1505500000 | 9.1777
A05 | IAC | Unknown | Cy5 | 33.06 | A05 | IAC | Unknown | Cy5 | 24.93 | A06 | 32.94 |  |  | 24.83 |  |  |  | Tube 6 | std_3 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 9.0808 | [5 empty] | upper | 37.2177
A06 | IAC | Unknown | Cy5 | 32.94 | A06 | IAC | Unknown | Cy5 | 24.83 | A07 | 33.29 |  |  | 25.13 |  |  |  | Tube 7 | std_4 | 0.0001 | 100 | 301100000000.0001 | 602200000 | 8.7797 | [5 empty] | lower | 33.4568
A07 | IAC | Unknown | Cy5 | 33.29 | A07 | IAC | Unknown | Cy5 | 25.13 | A08 | 32.97 |  |  | 24.84
A08 | IAC | Unknown | Cy5 | 32.97 | A08 | IAC | Unknown | Cy5 | 24.84 | A09 | 33.1 |  |  | 24.91
A09 | IAC | Unknown | Cy5 | 33.1 | A09 | IAC | Unknown | Cy5 | 24.91 | A10 | 33.16 |  |  | 24.92
A10 | IAC | Unknown | Cy5 | 33.16 | A10 | IAC | Unknown | Cy5 | 24.92 | A11 | 32.87 |  |  | 24.82
A11 | IAC | Unknown | Cy5 | 32.87 | A11 | IAC | Unknown | Cy5 | 24.82 | A12 | 33.01 |  |  | 24.85
A12 | IAC | Unknown | Cy5 | 33.01 | A12 | IAC | Unknown | Cy5 | 24.85 | B01 | 36.84 | 37.22 | 0.30486 | 28.65 | 29.4827 | 0.39241
B01 | IAC | Unknown | Cy5 | 36.84 | B01 | IAC | Unknown | Cy5 | 28.65 | B02 | 37.22 |  |  | 29.22 |  |  | C1 | V1 | C2 | V2
B02 | IAC | Unknown | Cy5 | 37.22 | B02 | IAC | Unknown | Cy5 | 29.22 | B03 | 36.83 |  |  | 29.32 |  |  | 12044000000 | x | 3011000000 | 2
B03 | IAC | Unknown | Cy5 | 36.83 | B03 | IAC | Unknown | Cy5 | 29.32 | B04 | 36.79 |  |  | 29.31 |  |  |  | 0.5
B04 | IAC | Unknown | Cy5 | 36.79 | B04 | IAC | Unknown | Cy5 | 29.31 | B05 | 37.31 |  |  | 29.43
B05 | IAC | Unknown | Cy5 | 37.31 | B05 | IAC | Unknown | Cy5 | 29.43 | B06 | 37.29 |  |  | 29.51 |  |  | Determining how much from tube 2 IAC, I need to add to get to same concentration of "std_2.2".
B06 | IAC | Unknown | Cy5 | 37.29 | B06 | IAC | Unknown | Cy5 | 29.51 | B07 | 37.85 |  |  | 29.53 |  |  | The equivalent of 0.5ul from std_2 is the same as 2ul std_2_2
B07 | IAC | Unknown | Cy5 | 37.85 | B07 | IAC | Unknown | Cy5 | 29.53 | B08 | 37.34 |  |  | 29.02
B08 | IAC | Unknown | Cy5 | 37.34 | B08 | IAC | Unknown | Cy5 | 29.02 | B09 | 36.9 |  |  | 30.57
B09 | IAC | Unknown | Cy5 | 36.9 | B09 | IAC | Unknown | Cy5 | 30.57 | B10 | 37.36 |  |  | 29.4
B10 | IAC | Unknown | Cy5 | 37.36 | B10 | IAC | Unknown | Cy5 | 29.4 | B11 | 37.41 |  |  | 29.55
B11 | IAC | Unknown | Cy5 | 37.41 | B11 | IAC | Unknown | Cy5 | 29.55 | B12 | 37.12 |  |  | 29.45
B12 | IAC | Unknown | Cy5 | 37.12 | B12 | IAC | Unknown | Cy5 | 29.45 | C01 | 38.82 | 39.1589 | 0.55777 | 30.22 | 31.4436 | 0.24299
C01 | IAC | Unknown | Cy5 | 38.82 | C01 | IAC | Unknown | Cy5 | 30.22 | C02 | 39.3 |  |  | 31.29
C02 | IAC | Unknown | Cy5 | 39.3 | C02 | IAC | Unknown | Cy5 | 31.29 | C03 | 38.97 |  |  | 31.57
C03 | IAC | Unknown | Cy5 | 38.97 | C03 | IAC | Unknown | Cy5 | 31.57 | C04 | 38.59 |  |  | 31.41
C04 | IAC | Unknown | Cy5 | 38.59 | C04 | IAC | Unknown | Cy5 | 31.41 | C05 | 38.33 |  |  | 31.81
C05 | IAC | Unknown | Cy5 | 38.33 | C05 | IAC | Unknown | Cy5 | 31.81 | C06 | 39.78 |  |  | 31.52
C06 | IAC | Unknown | Cy5 | 39.78 | C06 | IAC | Unknown | Cy5 | 31.52 | C07 |  |  |  | 31.67
C07 | IAC | Unknown | Cy5 |  | C07 | IAC | Unknown | Cy5 | 31.67 | C08 | 39.77 |  |  | 31.39 |  |  | TubeName | Copy Number/ul | Log(Copy Number/ul) | Ct_STD_MEAN | Ct_STD_STDEV | 2*stdev | Ct_OPT_MEAN | Ct_OPT_STDEV | 2*stdev
C08 | IAC | Unknown | Cy5 | 39.77 | C08 | IAC | Unknown | Cy5 | 31.39 | C09 | 39.88 |  |  | 31.39 |  |  | std_1 | 120440000000 | 11.0808 | 31.5445 | 0.95961 | 1.9192 | 28.6309 | 0.1916 | 0.38319
C09 | IAC | Unknown | Cy5 | 39.88 | C09 | IAC | Unknown | Cy5 | 31.39 | C10 |  |  |  | 31.69 |  |  | std_2 | 12044000000 | 10.0808 | 34.0182 | 1.5515 | 3.1029 | 28.7864 | 0.36806 | 0.73611
C10 | IAC | Unknown | Cy5 |  | C10 | IAC | Unknown | Cy5 | 31.69 | C11 | 38.99 |  |  | 31.15 |  |  | std_2.1 | 6022000000 | 9.7797 | 34.0409 | 1.1303 | 2.2605 | 28.6327 | 0.15447 | 0.30895
C11 | IAC | Unknown | Cy5 | 38.99 | C11 | IAC | Unknown | Cy5 | 31.15 | C12 | 38.82 |  |  | 30.99 |  |  | std_2.2 | 3011000000 | 9.4787 | 33.4173 | 0.87373 | 1.7475 | 28.5664 | 0.21855 | 0.43711
C12 | IAC | Unknown | Cy5 | 38.82 | C12 | IAC | Unknown | Cy5 | 30.99 | D01 |  | #DIV/0! | #DIV/0! | 32.91 | 35.3373 | 0.94022 | std_3 | 1204400000 | 9.0808 | 33.2273 | 0.85107 | 1.7021 | 28.4464 | 0.13655 | 0.2731
D01 | IAC | Unknown | Cy5 |  | D01 | IAC | Unknown | Cy5 | 32.91 | D02 |  |  |  | 36.11 |  |  | std_4 | 602200000 | 8.7797 | 33.1164 | 1.0791 | 2.1583 | 28.3455 | 0.15175 | 0.30349
D02 | IAC | Unknown | Cy5 |  | D02 | IAC | Unknown | Cy5 | 36.11 | D03 |  |  |  | 35.36
D03 | IAC | Unknown | Cy5 |  | D03 | IAC | Unknown | Cy5 | 35.36 | D04 |  |  |  | 37.21
D04 | IAC | Unknown | Cy5 |  | D04 | IAC | Unknown | Cy5 | 37.21 | D05 |  |  |  | 35.41 |  |  | Legionella pneumophila positive control
D05 | IAC | Unknown | Cy5 |  | D05 | IAC | Unknown | Cy5 | 35.41 | D06 |  |  |  | 34.89 |  |  | Tube | Conc (uM) | Copy Number | copies/2ul_rxn
D06 | IAC | Unknown | Cy5 |  | D06 | IAC | Unknown | Cy5 | 34.89 | D07 |  |  |  | 35.51 |  |  | Orig | 100 | 2.4088e+15 | 1.2044e+14
D07 | IAC | Unknown | Cy5 |  | D07 | IAC | Unknown | Cy5 | 35.51 | D08 |  |  |  | 35.3 |  |  | std_0 | 1 | 6.022e+14 | 1.2044e+12
D08 | IAC | Unknown | Cy5 |  | D08 | IAC | Unknown | Cy5 | 35.3 | D09 |  |  |  | 35.4 |  |  | std_1 | 0.01 | 6.022e+12 | 12044000000
D09 | IAC | Unknown | Cy5 |  | D09 | IAC | Unknown | Cy5 | 35.4 | D10 |  |  |  | 34 |  |  | std_2 | 0.001 | 602200000000.0001 | 1204400000
D10 | IAC | Unknown | Cy5 |  | D10 | IAC | Unknown | Cy5 | 34 | D11 |  |  |  | 35.77 |  |  | std_3 | 0.0001 | 60220000000 | 120440000
D11 | IAC | Unknown | Cy5 |  | D11 | IAC | Unknown | Cy5 | 35.77 | D12 |  |  |  | 33.75 |  |  | std_4 | 1e-05 | 6022000000 | 12044000
D12 | IAC | Unknown | Cy5 |  | D12 | IAC | Unknown | Cy5 | 33.75 | E01 | [6 empty] | std_5 | 1e-06 | 602200000 | 1204400
E01 | IAC | Unknown | Cy5 |  | E01 | IAC | Unknown | Cy5 |  | E02 | [6 empty] | std_6 | 1e-07 | 60220000 | 120440
E02 | IAC | Unknown | Cy5 |  | E02 | IAC | Unknown | Cy5 |  | E03 | [6 empty] | std_7 | 1e-08 | 6022000 | 12044
E03 | IAC | Unknown | Cy5 |  | E03 | IAC | Unknown | Cy5 |  | E04 | [6 empty] | std_8 | 1e-09 | 602200 | 1204.4
E04 | IAC | Unknown | Cy5 |  | E04 | IAC | Unknown | Cy5 |  | E05 | [6 empty] | std_9 | 1e-10 | 60220 | 120.44
E05 | IAC | Unknown | Cy5 |  | E05 | IAC | Unknown | Cy5 |  | E06 | [6 empty] | std_10 | 1e-11 | 6022 | 12.044
E06 | IAC | Unknown | Cy5 |  | E06 | IAC | Unknown | Cy5 |  | E07 | [6 empty] | std_11 | 1e-12 | 602.2 | 1.2044
E07 | IAC | Unknown | Cy5 |  | E07 | IAC | Unknown | Cy5 |  | E08 | [6 empty] | std_12 | 1e-13 | 60.22 | 0.12044
E08 | IAC | Unknown | Cy5 |  | E08 | IAC | Unknown | Cy5 |  | E09 | [6 empty] | std_13 | 1e-14 | 6.022 | 0.012044
E09 | IAC | Unknown | Cy5 |  | E09 | IAC | Unknown | Cy5 |  | E10 | [6 empty] | std_14 | 1e-15 | 0.6022 | 0.0012044
E10 | IAC | Unknown | Cy5 |  | E10 | IAC | Unknown | Cy5 |  | E11 | [6 empty] | std_15 | 1e-16 | 0.06022 | 0.00012044
E11 | IAC | Unknown | Cy5 |  | E11 | IAC | Unknown | Cy5 |  | E12
E12 | IAC | Unknown | Cy5 |  | E12 | IAC | Unknown | Cy5 |  | F01
F01 | IAC | Unknown | Cy5 |  | F01 | IAC | Unknown | Cy5 |  | F02
F02 | IAC | Unknown | Cy5 |  | F02 | IAC | Unknown | Cy5 |  | F03
F03 | IAC | Unknown | Cy5 |  | F03 | IAC | Unknown | Cy5 |  | F04
F04 | IAC | Unknown | Cy5 |  | F04 | IAC | Unknown | Cy5 |  | F05
F05 | IAC | Unknown | Cy5 |  | F05 | IAC | Unknown | Cy5 |  | F06
F06 | IAC | Unknown | Cy5 |  | F06 | IAC | Unknown | Cy5 |  | F07 | [20 empty] | Next Steps |  | ﻿Make big batch of Leg pneumo master mix with optimized concentrations. Test/Re-Test the following
F07 | IAC | Unknown | Cy5 |  | F07 | IAC | Unknown | Cy5 |  | F08 | [7 empty] | TubeName | std_1 | std_2 | std_2.1 | std_2.2 | std_3 | std_4 | [6 empty] | Liquid Master Mix in Tube
F08 | IAC | Unknown | Cy5 |  | F08 | IAC | Unknown | Cy5 |  | F09 | [7 empty] | STD | 32.51 | 32.57 | 33.01 | 33.49 | 33.69 | 32.89 | [6 empty] | Assay | Estimated Reactions | Notes
F09 | IAC | Unknown | Cy5 |  | F09 | IAC | Unknown | Cy5 |  | F10 | [8 empty] | 30.69 | 32.56 | 32.57 | 32.11 | 32.75 | 31.96 | [6 empty] | Leg pneumo std curve | 100
F10 | IAC | Unknown | Cy5 |  | F10 | IAC | Unknown | Cy5 |  | F11 | [8 empty] | 31.41 | 31.88 | 34.47 | 32.75 | 33.5 | 32.75 | [6 empty] | LOD Determination | 200-400
F11 | IAC | Unknown | Cy5 |  | F11 | IAC | Unknown | Cy5 |  | F12 | [8 empty] | 31.86 | 34.52 | 33.75 | 33.18 | 33.86 | 32.4 | [6 empty] | Sample Detection (Leg pneumo gDNA, tablet from ieLab) | 100 | 15ul formulation
F12 | IAC | Unknown | Cy5 |  | F12 | IAC | Unknown | Cy5 |  | G01 | 32.05 |  |  | 24.69 |  |  |  |  | 32.37 | 33.24 | 35.43 | 34.91 | 33.76 | 33.22
G01 | IAC | Standard | Cy5 | 32.05 | G01 | IAC | Standard | Cy5 | 24.69 | G02 | 32.77 |  |  | 24.87 |  |  |  |  | 31.47 | 34.08 | 32.77 | 32.12 | 31.77 | 32.65
G02 | IAC | Standard | Cy5 | 32.77 | G02 | IAC | Standard | Cy5 | 24.87 | G03 | 32.86 |  |  | 24.91 |  |  |  |  | 31.08 | 33.07 | 35.2 | 33.82 | 32.75 | 31.3 | [6 empty] | Lyophilized Master Mix in 8-well strip cap tubes
G03 | IAC | Standard | Cy5 | 32.86 | G03 | IAC | Standard | Cy5 | 24.91 | G04 | 36.71 |  |  | 29.13 |  |  |  |  | 30.04 | 34.53 | 35.92 | 33.36 | 34.76 | 33.66 | [6 empty] | Leg pneumo std curve | 100
G04 | IAC | Standard | Cy5 | 36.71 | G04 | IAC | Standard | Cy5 | 29.13 | G05 | 36.91 |  |  | 29.41 |  |  |  |  | 33.13 | 34.93 | 33.65 | 33.46 | 33.59 | 34.61 | [6 empty] | Sample Detection (Leg pneumo gDNA, tablet from ieLab) | 100
G05 | IAC | Standard | Cy5 | 36.91 | G05 | IAC | Standard | Cy5 | 29.41 | G06 | 37.09 |  |  | 29.49 |  |  |  |  | 30.34 | 37.11 | 34.4 | 34.45 | 32.18 | 34.7 | [6 empty] | Q16 validation | 200
G06 | IAC | Standard | Cy5 | 37.09 | G06 | IAC | Standard | Cy5 | 29.49 | G07 | 39.1 |  |  | 31.39 |  |  |  |  | 32.09 | 35.71 | 32.52 | 33.94 | 32.89 | 34.14
G07 | IAC | Standard | Cy5 | 39.1 | G07 | IAC | Standard | Cy5 | 31.39 | G08 | 38.54 |  |  | 31.11
G08 | IAC | Standard | Cy5 | 38.54 | G08 | IAC | Standard | Cy5 | 31.11 | G09 | 38.92 |  |  | 31.23 |  |  |  | OPT | std_1 | std_2 | std_2.1 | std_2.2 | std_3 | std_4 |  | std_1 | 28.57 | 28.58 | 28.5 | 28.64 | 28.41 | 28.92 | 28.5 | 28.82 | 28.87 | 28.35 | 28.78 | std_1=[28.57,28.58,28.5,28.64,28.41,28.92,28.5,28.82,28.87,28.35,28.78]
G09 | IAC | Standard | Cy5 | 38.92 | G09 | IAC | Standard | Cy5 | 31.23 | G10 |  |  |  | 35.07 |  |  |  |  | 28.57 | 28.73 | 28.64 | 28.63 | 28.24 | 28.3 |  | std_2 | 28.73 | 28.59 | 28.64 | 28.87 | 28.67 | 28.87 | 28.56 | 29.78 | 28.49 | 28.98 | 28.47 | std_2=[28.73,28.59,28.64,28.87,28.67,28.87,28.56,29.78,28.49,28.98,28.47]
G10 | IAC | Standard | Cy5 |  | G10 | IAC | Standard | Cy5 | 35.07 | G11 |  |  |  | 36.12 |  |  |  |  | 28.58 | 28.59 | 28.66 | 28.39 | 28.43 | 28.22 |  | std_2.1 | 28.64 | 28.66 | 28.56 | 28.83 | 28.49 | 28.65 | 28.62 | 28.82 | 28.84 | 28.36 | 28.49 | std_2.1=[28.64,28.66,28.56,28.83,28.49,28.65,28.62,28.82,28.84,28.36,28.49]
G11 | IAC | Standard | Cy5 |  | G11 | IAC | Standard | Cy5 | 36.12 | G12 |  |  |  | 35.83 |  |  |  |  | 28.5 | 28.64 | 28.56 | 28.7 | 28.36 | 28.39 |  | std_2.2 | 28.63 | 28.39 | 28.7 | 28.5 | 29.14 | 28.39 | 28.57 | 28.44 | 28.45 | 28.62 | 28.4 | std_2.2=[28.63,28.39,28.7,28.5,29.14,28.39,28.57,28.44,28.45,28.62,28.4]
G12 | IAC | Standard | Cy5 |  | G12 | IAC | Standard | Cy5 | 35.83 | H01 |  |  |  | 37.02 |  |  |  |  | 28.64 | 28.87 | 28.83 | 28.5 | 28.58 | 28.35 |  | std_3 | 28.24 | 28.43 | 28.36 | 28.58 | 28.44 | 28.67 | 28.3 | 28.56 | 28.45 | 28.31 | 28.57 | std_3=[28.24,28.43,28.36,28.58,28.44,28.67,28.3,28.56,28.45,28.31,28.57]
H01 | IAC | Standard | Cy5 |  | H01 | IAC | Standard | Cy5 | 37.02 | H02 |  |  |  | 37.07 |  |  |  |  | 28.41 | 28.67 | 28.49 | 29.14 | 28.44 | 28.09 |  | std_4 | 28.3 | 28.22 | 28.39 | 28.35 | 28.09 | 28.39 | 28.24 | 28.7 | 28.37 | 28.42 | 28.33 | std_4=[28.3,28.22,28.39,28.35,28.09,28.39,28.24,28.7,28.37,28.42,28.33]
H02 | IAC | Standard | Cy5 |  | H02 | IAC | Standard | Cy5 | 37.07 | H03 |  |  |  | 36.72 |  |  |  |  | 28.92 | 28.87 | 28.65 | 28.39 | 28.67 | 28.39
H03 | IAC | Standard | Cy5 |  | H03 | IAC | Standard | Cy5 | 36.72 | H04 | [8 empty] | 28.5 | 28.56 | 28.62 | 28.57 | 28.3 | 28.24
H04 | IAC | Standard | Cy5 |  | H04 | IAC | Standard | Cy5 |  | H05 | [8 empty] | 28.82 | 29.78 | 28.82 | 28.44 | 28.56 | 28.7
H05 | IAC | Standard | Cy5 |  | H05 | IAC | Standard | Cy5 |  | H06 |  |  |  | 39.86 |  |  |  |  | 28.87 | 28.49 | 28.84 | 28.45 | 28.45 | 28.37
H06 | IAC | Standard | Cy5 |  | H06 | IAC | Standard | Cy5 | 39.86 | H07 | [8 empty] | 28.35 | 28.98 | 28.36 | 28.62 | 28.31 | 28.42
H07 | IAC | Standard | Cy5 |  | H07 | IAC | Standard | Cy5 |  | H08 | [8 empty] | 28.78 | 28.47 | 28.49 | 28.4 | 28.57 | 28.33
H08 | IAC | Standard | Cy5 |  | H08 | IAC | Standard | Cy5 |  | H09
H09 | IAC | Standard | Cy5 |  | H09 | IAC | Standard | Cy5 |  | H10
H10 | IAC | Negative | Cy5 |  | H10 | IAC | Negative | Cy5 |  | H11 |  | Leg pneumo = "std_5" |  |  | Leg pneumo = "std_5" |  |  |  | Group | Cq
H11 | IAC | Negative | Cy5 |  | H11 | IAC | Negative | Cy5 |  | H12 |  | Ct_STD_MEAN for Leg pneumo | Ct_STD_STDEV for Leg pneumo |  | Ct_STD_MEAN for Leg pneumo | Ct_STD_STDEV for Leg pneumo |  |  | std_1 | 28.57
H12 | IAC | Negative | Cy5 |  | H12 | IAC | Negative | Cy5 |  | A01 | 32.92 | 31.5445 | 0.95961 | 28.09 | 28.6309 | 0.1916 |  |  | std_1 | 28.58
A01 | Legpneumo | Unknown | FAM | 32.92 | A01 | Leg pneumo | Unknown | FAM | 28.09 | A02 | 32.51 |  |  | 28.57 |  |  |  |  | std_1 | 28.5
A02 | Legpneumo | Unknown | FAM | 32.51 | A02 | Leg pneumo | Unknown | FAM | 28.57 | A03 | 30.69 |  |  | 28.58 |  |  |  |  | std_1 | 28.64
A03 | Legpneumo | Unknown | FAM | 30.69 | A03 | Leg pneumo | Unknown | FAM | 28.58 | A04 | 31.41 |  |  | 28.5 |  |  |  |  | std_1 | 28.41
A04 | Legpneumo | Unknown | FAM | 31.41 | A04 | Leg pneumo | Unknown | FAM | 28.5 | A05 | 31.86 |  |  | 28.64 |  |  |  |  | std_1 | 28.92
A05 | Legpneumo | Unknown | FAM | 31.86 | A05 | Leg pneumo | Unknown | FAM | 28.64 | A06 | 32.37 |  |  | 28.41 |  |  |  |  | std_1 | 28.5
A06 | Legpneumo | Unknown | FAM | 32.37 | A06 | Leg pneumo | Unknown | FAM | 28.41 | A07 | 31.47 |  |  | 28.92 |  |  |  |  | std_1 | 28.82
A07 | Legpneumo | Unknown | FAM | 31.47 | A07 | Leg pneumo | Unknown | FAM | 28.92 | A08 | 31.08 |  |  | 28.5 |  |  |  |  | std_1 | 28.87
A08 | Legpneumo | Unknown | FAM | 31.08 | A08 | Leg pneumo | Unknown | FAM | 28.5 | A09 | 30.04 |  |  | 28.82 |  |  |  |  | std_1 | 28.35
A09 | Legpneumo | Unknown | FAM | 30.04 | A09 | Leg pneumo | Unknown | FAM | 28.82 | A10 | 33.13 |  |  | 28.87 |  |  |  |  | std_1 | 28.78
A10 | Legpneumo | Unknown | FAM | 33.13 | A10 | Leg pneumo | Unknown | FAM | 28.87 | A11 | 30.34 |  |  | 28.35 |  |  |  |  | std_2 | 28.73
A11 | Legpneumo | Unknown | FAM | 30.34 | A11 | Leg pneumo | Unknown | FAM | 28.35 | A12 | 32.09 |  |  | 28.78 |  |  |  |  | std_2 | 28.59
A12 | Legpneumo | Unknown | FAM | 32.09 | A12 | Leg pneumo | Unknown | FAM | 28.78 | B01 | 35.14 | 34.0182 | 1.5515 | 28.12 | 28.7864 | 0.36806 |  |  | std_2 | 28.64
B01 | Legpneumo | Unknown | FAM | 35.14 | B01 | Leg pneumo | Unknown | FAM | 28.12 | B02 | 32.57 |  |  | 28.73 |  |  |  |  | std_2 | 28.87
B02 | Legpneumo | Unknown | FAM | 32.57 | B02 | Leg pneumo | Unknown | FAM | 28.73 | B03 | 32.56 |  |  | 28.59 |  |  |  |  | std_2 | 28.67
B03 | Legpneumo | Unknown | FAM | 32.56 | B03 | Leg pneumo | Unknown | FAM | 28.59 | B04 | 31.88 |  |  | 28.64 |  |  |  |  | std_2 | 28.87
B04 | Legpneumo | Unknown | FAM | 31.88 | B04 | Leg pneumo | Unknown | FAM | 28.64 | B05 | 34.52 |  |  | 28.87 |  |  |  |  | std_2 | 28.56
B05 | Legpneumo | Unknown | FAM | 34.52 | B05 | Leg pneumo | Unknown | FAM | 28.87 | B06 | 33.24 |  |  | 28.67 |  |  |  |  | std_2 | 29.78
B06 | Legpneumo | Unknown | FAM | 33.24 | B06 | Leg pneumo | Unknown | FAM | 28.67 | B07 | 34.08 |  |  | 28.87 |  |  |  |  | std_2 | 28.49
B07 | Legpneumo | Unknown | FAM | 34.08 | B07 | Leg pneumo | Unknown | FAM | 28.87 | B08 | 33.07 |  |  | 28.56 |  |  |  |  | std_2 | 28.98
B08 | Legpneumo | Unknown | FAM | 33.07 | B08 | Leg pneumo | Unknown | FAM | 28.56 | B09 | 34.53 |  |  | 29.78 |  |  |  |  | std_2 | 28.47
B09 | Legpneumo | Unknown | FAM | 34.53 | B09 | Leg pneumo | Unknown | FAM | 29.78 | B10 | 34.93 |  |  | 28.49 |  |  |  |  | std_2.1 | 28.64
B10 | Legpneumo | Unknown | FAM | 34.93 | B10 | Leg pneumo | Unknown | FAM | 28.49 | B11 | 37.11 |  |  | 28.98 |  |  |  |  | std_2.1 | 28.66
B11 | Legpneumo | Unknown | FAM | 37.11 | B11 | Leg pneumo | Unknown | FAM | 28.98 | B12 | 35.71 |  |  | 28.47 |  |  |  |  | std_2.1 | 28.56
B12 | Legpneumo | Unknown | FAM | 35.71 | B12 | Leg pneumo | Unknown | FAM | 28.47 | C01 | 33.85 | 34.0409 | 1.1303 | 27.95 | 28.6327 | 0.15447 |  |  | std_2.1 | 28.83
C01 | Legpneumo | Unknown | FAM | 33.85 | C01 | Leg pneumo | Unknown | FAM | 27.95 | C02 | 33.28 |  |  | 28.64 |  |  |  |  | std_2.1 | 28.49
C02 | Legpneumo | Unknown | FAM | 33.28 | C02 | Leg pneumo | Unknown | FAM | 28.64 | C03 | 33.01 |  |  | 28.66 |  |  |  |  | std_2.1 | 28.65
C03 | Legpneumo | Unknown | FAM | 33.01 | C03 | Leg pneumo | Unknown | FAM | 28.66 | C04 | 32.57 |  |  | 28.56 |  |  |  |  | std_2.1 | 28.62
C04 | Legpneumo | Unknown | FAM | 32.57 | C04 | Leg pneumo | Unknown | FAM | 28.56 | C05 | 34.47 |  |  | 28.83 |  |  |  |  | std_2.1 | 28.82
C05 | Legpneumo | Unknown | FAM | 34.47 | C05 | Leg pneumo | Unknown | FAM | 28.83 | C06 | 33.75 |  |  | 28.49 |  |  |  |  | std_2.1 | 28.84
C06 | Legpneumo | Unknown | FAM | 33.75 | C06 | Leg pneumo | Unknown | FAM | 28.49 | C07 | 35.43 |  |  | 28.65 |  |  |  |  | std_2.1 | 28.36
C07 | Legpneumo | Unknown | FAM | 35.43 | C07 | Leg pneumo | Unknown | FAM | 28.65 | C08 | 32.77 |  |  | 28.62 |  |  |  |  | std_2.1 | 28.49
C08 | Legpneumo | Unknown | FAM | 32.77 | C08 | Leg pneumo | Unknown | FAM | 28.62 | C09 | 35.2 |  |  | 28.82 |  |  |  |  | std_2.2 | 28.63
C09 | Legpneumo | Unknown | FAM | 35.2 | C09 | Leg pneumo | Unknown | FAM | 28.82 | C10 | 35.92 |  |  | 28.84 |  |  |  |  | std_2.2 | 28.39
C10 | Legpneumo | Unknown | FAM | 35.92 | C10 | Leg pneumo | Unknown | FAM | 28.84 | C11 | 33.65 |  |  | 28.36 |  |  |  |  | std_2.2 | 28.7
C11 | Legpneumo | Unknown | FAM | 33.65 | C11 | Leg pneumo | Unknown | FAM | 28.36 | C12 | 34.4 |  |  | 28.49 |  |  |  |  | std_2.2 | 28.5
C12 | Legpneumo | Unknown | FAM | 34.4 | C12 | Leg pneumo | Unknown | FAM | 28.49 | D01 | 32.52 | 33.4173 | 0.87373 | 27.85 | 28.5664 | 0.21855 |  |  | std_2.2 | 29.14
D01 | Legpneumo | Unknown | FAM | 32.52 | D01 | Leg pneumo | Unknown | FAM | 27.85 | D02 | 33.49 |  |  | 28.63 |  |  |  |  | std_2.2 | 28.39
D02 | Legpneumo | Unknown | FAM | 33.49 | D02 | Leg pneumo | Unknown | FAM | 28.63 | D03 | 32.11 |  |  | 28.39 |  |  |  |  | std_2.2 | 28.57
D03 | Legpneumo | Unknown | FAM | 32.11 | D03 | Leg pneumo | Unknown | FAM | 28.39 | D04 | 32.75 |  |  | 28.7 |  |  |  |  | std_2.2 | 28.44
D04 | Legpneumo | Unknown | FAM | 32.75 | D04 | Leg pneumo | Unknown | FAM | 28.7 | D05 | 33.18 |  |  | 28.5 |  |  |  |  | std_2.2 | 28.45
D05 | Legpneumo | Unknown | FAM | 33.18 | D05 | Leg pneumo | Unknown | FAM | 28.5 | D06 | 34.91 |  |  | 29.14 |  |  |  |  | std_2.2 | 28.62
D06 | Legpneumo | Unknown | FAM | 34.91 | D06 | Leg pneumo | Unknown | FAM | 29.14 | D07 | 32.12 |  |  | 28.39 |  |  |  |  | std_2.2 | 28.4
D07 | Legpneumo | Unknown | FAM | 32.12 | D07 | Leg pneumo | Unknown | FAM | 28.39 | D08 | 33.82 |  |  | 28.57 |  |  |  |  | std_3 | 28.24
D08 | Legpneumo | Unknown | FAM | 33.82 | D08 | Leg pneumo | Unknown | FAM | 28.57 | D09 | 33.36 |  |  | 28.44 |  |  |  |  | std_3 | 28.43
D09 | Legpneumo | Unknown | FAM | 33.36 | D09 | Leg pneumo | Unknown | FAM | 28.44 | D10 | 33.46 |  |  | 28.45 |  |  |  |  | std_3 | 28.36
D10 | Legpneumo | Unknown | FAM | 33.46 | D10 | Leg pneumo | Unknown | FAM | 28.45 | D11 | 34.45 |  |  | 28.62 |  |  |  |  | std_3 | 28.58
D11 | Legpneumo | Unknown | FAM | 34.45 | D11 | Leg pneumo | Unknown | FAM | 28.62 | D12 | 33.94 |  |  | 28.4 |  |  |  |  | std_3 | 28.44
D12 | Legpneumo | Unknown | FAM | 33.94 | D12 | Leg pneumo | Unknown | FAM | 28.4 | E01 | 30.9 | 33.2273 | 0.85107 | 27.64 | 28.4464 | 0.13655 |  |  | std_3 | 28.67
E01 | Legpneumo | Unknown | FAM | 30.9 | E01 | Leg pneumo | Unknown | FAM | 27.64 | E02 | 33.69 |  |  | 28.24 |  |  |  |  | std_3 | 28.3
E02 | Legpneumo | Unknown | FAM | 33.69 | E02 | Leg pneumo | Unknown | FAM | 28.24 | E03 | 32.75 |  |  | 28.43 |  |  |  |  | std_3 | 28.56
E03 | Legpneumo | Unknown | FAM | 32.75 | E03 | Leg pneumo | Unknown | FAM | 28.43 | E04 | 33.5 |  |  | 28.36 |  |  |  |  | std_3 | 28.45
E04 | Legpneumo | Unknown | FAM | 33.5 | E04 | Leg pneumo | Unknown | FAM | 28.36 | E05 | 33.86 |  |  | 28.58 |  |  |  |  | std_3 | 28.31
E05 | Legpneumo | Unknown | FAM | 33.86 | E05 | Leg pneumo | Unknown | FAM | 28.58 | E06 | 33.76 |  |  | 28.44 |  |  |  |  | std_3 | 28.57
E06 | Legpneumo | Unknown | FAM | 33.76 | E06 | Leg pneumo | Unknown | FAM | 28.44 | E07 | 31.77 |  |  | 28.67 |  |  |  |  | std_4 | 28.3
E07 | Legpneumo | Unknown | FAM | 31.77 | E07 | Leg pneumo | Unknown | FAM | 28.67 | E08 | 32.75 |  |  | 28.3 |  |  |  |  | std_4 | 28.22
E08 | Legpneumo | Unknown | FAM | 32.75 | E08 | Leg pneumo | Unknown | FAM | 28.3 | E09 | 34.76 |  |  | 28.56 |  |  |  |  | std_4 | 28.39
E09 | Legpneumo | Unknown | FAM | 34.76 | E09 | Leg pneumo | Unknown | FAM | 28.56 | E10 | 33.59 |  |  | 28.45 |  |  |  |  | std_4 | 28.35
E10 | Legpneumo | Unknown | FAM | 33.59 | E10 | Leg pneumo | Unknown | FAM | 28.45 | E11 | 32.18 |  |  | 28.31 |  |  |  |  | std_4 | 28.09
E11 | Legpneumo | Unknown | FAM | 32.18 | E11 | Leg pneumo | Unknown | FAM | 28.31 | E12 | 32.89 |  |  | 28.57 |  |  |  |  | std_4 | 28.39
E12 | Legpneumo | Unknown | FAM | 32.89 | E12 | Leg pneumo | Unknown | FAM | 28.57 | F01 | 34.25 | 33.1164 | 1.0791 | 27.79 | 28.3455 | 0.15175 |  |  | std_4 | 28.24
F01 | Legpneumo | Unknown | FAM | 34.25 | F01 | Leg pneumo | Unknown | FAM | 27.79 | F02 | 32.89 |  |  | 28.3 |  |  |  |  | std_4 | 28.7
F02 | Legpneumo | Unknown | FAM | 32.89 | F02 | Leg pneumo | Unknown | FAM | 28.3 | F03 | 31.96 |  |  | 28.22 |  |  |  |  | std_4 | 28.37
F03 | Legpneumo | Unknown | FAM | 31.96 | F03 | Leg pneumo | Unknown | FAM | 28.22 | F04 | 32.75 |  |  | 28.39 |  |  |  |  | std_4 | 28.42
F04 | Legpneumo | Unknown | FAM | 32.75 | F04 | Leg pneumo | Unknown | FAM | 28.39 | F05 | 32.4 |  |  | 28.35 |  |  |  |  | std_4 | 28.33
F05 | Legpneumo | Unknown | FAM | 32.4 | F05 | Leg pneumo | Unknown | FAM | 28.35 | F06 | 33.22 |  |  | 28.09
F06 | Legpneumo | Unknown | FAM | 33.22 | F06 | Leg pneumo | Unknown | FAM | 28.09 | F07 | 32.65 |  |  | 28.39
F07 | Legpneumo | Unknown | FAM | 32.65 | F07 | Leg pneumo | Unknown | FAM | 28.39 | F08 | 31.3 |  |  | 28.24
F08 | Legpneumo | Unknown | FAM | 31.3 | F08 | Leg pneumo | Unknown | FAM | 28.24 | F09 | 33.66 |  |  | 28.7
F09 | Legpneumo | Unknown | FAM | 33.66 | F09 | Leg pneumo | Unknown | FAM | 28.7 | F10 | 34.61 |  |  | 28.37
F10 | Legpneumo | Unknown | FAM | 34.61 | F10 | Leg pneumo | Unknown | FAM | 28.37 | F11 | 34.7 |  |  | 28.42
F11 | Legpneumo | Unknown | FAM | 34.7 | F11 | Leg pneumo | Unknown | FAM | 28.42 | F12 | 34.14 |  |  | 28.33
F12 | Legpneumo | Unknown | FAM | 34.14 | F12 | Leg pneumo | Unknown | FAM | 28.33 | G01 | 29.74 |  |  | 27.84
G01 | Legpneumo | Unknown | FAM | 29.74 | G01 | Leg pneumo | Unknown | FAM | 27.84 | G02 | 31.07 |  |  | 28.22
G02 | Legpneumo | Unknown | FAM | 31.07 | G02 | Leg pneumo | Unknown | FAM | 28.22 | G03 | 31.05 |  |  | 28.12
G03 | Legpneumo | Unknown | FAM | 31.05 | G03 | Leg pneumo | Unknown | FAM | 28.12 | G04 | 33.31 |  |  | 28.59
G04 | Legpneumo | Unknown | FAM | 33.31 | G04 | Leg pneumo | Unknown | FAM | 28.59 | G05 | 34.59 |  |  | 28.61
G05 | Legpneumo | Unknown | FAM | 34.59 | G05 | Leg pneumo | Unknown | FAM | 28.61 | G06 | 33.15 |  |  | 28.73
G06 | Legpneumo | Unknown | FAM | 33.15 | G06 | Leg pneumo | Unknown | FAM | 28.73 | G07 | 31.22 |  |  | 28.19
G07 | Legpneumo | Unknown | FAM | 31.22 | G07 | Leg pneumo | Unknown | FAM | 28.19 | G08 | 33.68 |  |  | 28.15
G08 | Legpneumo | Unknown | FAM | 33.68 | G08 | Leg pneumo | Unknown | FAM | 28.15 | G09 | 34.23 |  |  | 28.42
G09 | Legpneumo | Unknown | FAM | 34.23 | G09 | Leg pneumo | Unknown | FAM | 28.42 | G10 | 33.07 |  |  | 28.18
G10 | Legpneumo | Unknown | FAM | 33.07 | G10 | Leg pneumo | Unknown | FAM | 28.18 | G11 | 35.05 |  |  | 28.55
G11 | Legpneumo | Unknown | FAM | 35.05 | G11 | Leg pneumo | Unknown | FAM | 28.55 | G12 | 34.59 |  |  | 28.34
G12 | Legpneumo | Unknown | FAM | 34.59 | G12 | Leg pneumo | Unknown | FAM | 28.34 | H01 | 31.39 |  |  | 28.11
H01 | Legpneumo | Unknown | FAM | 31.39 | H01 | Leg pneumo | Unknown | FAM | 28.11 | H02 | 33.73 |  |  | 28.33
H02 | Legpneumo | Unknown | FAM | 33.73 | H02 | Leg pneumo | Unknown | FAM | 28.33 | H03 | 31.94 |  |  | 28.32
H03 | Legpneumo | Unknown | FAM | 31.94 | H03 | Leg pneumo | Unknown | FAM | 28.32 | H04 | 34.13 |  |  | 28.32
H04 | Legpneumo | Unknown | FAM | 34.13 | H04 | Leg pneumo | Unknown | FAM | 28.32 | H05 | 31.32 |  |  | 28.31
H05 | Legpneumo | Unknown | FAM | 31.32 | H05 | Leg pneumo | Unknown | FAM | 28.31 | H06 | 33.36 |  |  | 28.24
H06 | Legpneumo | Unknown | FAM | 33.36 | H06 | Leg pneumo | Unknown | FAM | 28.24 | H07 | 32.26 |  |  | 28.1
H07 | Legpneumo | Unknown | FAM | 32.26 | H07 | Leg pneumo | Unknown | FAM | 28.1 | H08 | 32.34 |  |  | 28.13
H08 | Legpneumo | Unknown | FAM | 32.34 | H08 | Leg pneumo | Unknown | FAM | 28.13 | H09 | 32.94 |  |  | 28.35
H09 | Legpneumo | Unknown | FAM | 32.94 | H09 | Leg pneumo | Unknown | FAM | 28.35 | H10 | 34.01 |  |  | 28.22
H10 | Legpneumo | Unknown | FAM | 34.01 | H10 | Leg pneumo | Unknown | FAM | 28.22 | H11 | 33.52 |  |  | 28.08
H11 | Legpneumo | Unknown | FAM | 33.52 | H11 | Leg pneumo | Unknown | FAM | 28.08 | H12 | 34.41 |  |  | 28.28
H12 | Legpneumo | Unknown | FAM | 34.41 | H12 | Leg pneumo | Unknown | FAM | 28.28
CONCLUSIONS:
Using the optimized (OPT) IAC master mix (F,R,P = 50, 50, 400nM) results in a 7-8 decrease in Cq compared to using the previous standard (STD) mix (F,R,P = 250, 250, 500nM).
Error bars represent 95% confidence interval. For copy number 1.20e11, the error bars are so small that the marker covers them. Each marker represents the average of 11 replicates.
The OPT mix not only results in lower Cq for the same concentration, but the OPT mix was more sensitive and detected a 1/10 concentration that the STD mix did not.
The IAC positive control is not linear, but it is binary (yes/no). Thus, the IAC positive control should not be used to quantify unknown concentrations of IAC positive control, but can be used as positive control for a reaction.
If a 100 samples are run with an IAC positive control concentration of 3x109copies/ul, 95 of the samples will have a Cq between 37 and 33.
THOUGHTS:
This data suggests a reason Greg and I didn't previously observe the IAC is due to our use at concentrations less than 3e9 copies/ul.
Surprised how many copies/ul are needed for amplification. Lowest limit of amplification is about 3e9 or about three million times higher than 1e3, the IAC concentration in Legionella spp assay.
A higher copy number may be beneficial as the IAC reaction may be resistant to low levels of contamination.
Important to optimize esp in multiplexed reactions because never know interplay.
Another reasons to optimize: Save 400% on lower F, R costs. Save 25% on P.
Hard to believe that decreasing the F and R primer by 1/5 results in a lower Cq and an increased sensitivity compared to the STD mix. This data suggests that the optimal concentration of IAC F, R, P is 50, 50, 400nM. The concentration that should be used is 3e9 copies/ul.
CONCLUSIONS:
Across different IAC pos control concentrations, Leg pneumo positive control consistent. That is, Leg pneumo reactions amplify independently of whether IAC pos control is at high or low concentration.
The OPT master mix contributes to stability (decreased std dev) of the amplification of Legionella pneumophila (1.2x106 copies/ul) compared to the stability of the STD master mix.
Error bars represent 95% confidence interval (2*std dev).
CONCLUSIONS
From 95% CI we may assume that there is no difference. However, One way Anova sheet "ANOVA" and Welch Anova show difference between the means. However, just because there is a difference doesn't mean it's practical.
Jupyter workbook in this folder.

## ANOVA
Anova: Single Factor
SUMMARY
Groups | Count | Sum | Average | Variance
std_1 | 11 | 314.94 | 28.6309 | 0.036709
std_2 | 11 | 316.65 | 28.7864 | 0.13547
std_2.1 | 11 | 314.96 | 28.6327 | 0.023862
std_2.2 | 11 | 314.23 | 28.5664 | 0.047765
std_3 | 11 | 312.91 | 28.4464 | 0.018645
std_4 | 11 | 311.8 | 28.3455 | 0.023027
ANOVA
Source of Variation | SS | df | MS | F | P-value | F crit
Between Groups | 1.3217 | 5 | 0.26434 | 5.5558 | 0.00028604 | 2.3683
Within Groups | 2.8547 | 60 | 0.047579
Total | 4.1764 | 65

## LO_std_vs_opt
Well | Target | Property | Dye | Ct | Ct_STD_Mean | Ct_STD_STDEV | Well | Target | Property | Dye | Ct | Ct_OPT_Mean | Ct_OPT_STDEV |  |  |  | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/ul | Log(Copy Number/ul) | Ct_STD_MEAN | Ct_STD_STDEV | 2*stdev | Ct_OPT_MEAN | Ct_OPT_STDEV | 2*stdev | Diff | Diff (hi-lo)
A01 | IAC | Unknown | Cy5 | 33.53 | 33.37 | 0.25235 | A01 | IAC | Unknown | Cy5 | 26.01 | 26.5427 | 0.090784 |  |  |  | Tube 1 | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808 | 33.37 | 0.25235 | 0.5047 | 26.5427 | 0.090784 | 0.18157 | -6.8273 | -1.6282
A02 | IAC | Unknown | Cy5 | 33.9 |  |  | A02 | IAC | Unknown | Cy5 | 26.55 | [5 empty] | Tube 2 | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | 37.8309 | 0.37972 | 0.75944 | 30.93 | 0.2286 | 0.45721 | -6.9009 | -1.4473
A03 | IAC | Unknown | Cy5 | 33.31 |  |  | A03 | IAC | Unknown | Cy5 | 26.62 | [5 empty] | Tube 3 | std_2.1 | 0.005 | 5000 | 3.011e+12 | 6022000000 | 9.7797 |  |  | 0 | 32.9809 | 0.30972 | 0.61945 |  | -1.5373
A04 | IAC | Unknown | Cy5 | 33.34 |  |  | A04 | IAC | Unknown | Cy5 | 26.52 | [5 empty] | Tube 4 | std_2.2 | 0.0025 | 2500 | 1.5055e+12 | 3011000000 | 9.4787 |  |  | 0 | 34.8536 | 0.16378 | 0.32757 |  | 0.48364
A05 | IAC | Unknown | Cy5 | 33.2 |  |  | A05 | IAC | Unknown | Cy5 | 26.56 | [5 empty] | Tube 5 | std_2.3 | 0.00125 | 1250 | 752750000000.0001 | 1505500000 | 9.1777 |  |  |  | 37.2845 | 0.87985 | 1.7597
A06 | IAC | Unknown | Cy5 | 33.04 |  |  | A06 | IAC | Unknown | Cy5 | 26.45 | [5 empty] | Tube 6 | std_3 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 9.0808
A07 | IAC | Unknown | Cy5 | 33.48 |  |  | A07 | IAC | Unknown | Cy5 | 26.69 | [5 empty] | Tube 7 | std_4 | 0.0001 | 100 | 301100000000.0001 | 602200000 | 8.7797
A08 | IAC | Unknown | Cy5 | 33.73 |  |  | A08 | IAC | Unknown | Cy5 | 26.41
A09 | IAC | Unknown | Cy5 | 33.39 |  |  | A09 | IAC | Unknown | Cy5 | 26.58
A10 | IAC | Unknown | Cy5 | 33.3 |  |  | A10 | IAC | Unknown | Cy5 | 26.61
A11 | IAC | Unknown | Cy5 | 33.16 |  |  | A11 | IAC | Unknown | Cy5 | 26.4
A12 | IAC | Unknown | Cy5 | 33.22 |  |  | A12 | IAC | Unknown | Cy5 | 26.58
B01 | IAC | Unknown | Cy5 | 37.21 | 37.8309 | 0.37972 | B01 | IAC | Unknown | Cy5 | 30.25 | 30.93 | 0.2286
B02 | IAC | Unknown | Cy5 | 37.78 |  |  | B02 | IAC | Unknown | Cy5 | 30.88
B03 | IAC | Unknown | Cy5 | 37.32 |  |  | B03 | IAC | Unknown | Cy5 | 30.57
B04 | IAC | Unknown | Cy5 | 37.57 |  |  | B04 | IAC | Unknown | Cy5 | 30.75
B05 | IAC | Unknown | Cy5 | 37.84 |  |  | B05 | IAC | Unknown | Cy5 | 30.95
B06 | IAC | Unknown | Cy5 | 37.81 |  |  | B06 | IAC | Unknown | Cy5 | 30.81
B07 | IAC | Unknown | Cy5 | 38.68 |  |  | B07 | IAC | Unknown | Cy5 | 31.36
B08 | IAC | Unknown | Cy5 | 37.9 |  |  | B08 | IAC | Unknown | Cy5 | 30.82
B09 | IAC | Unknown | Cy5 | 37.55 |  |  | B09 | IAC | Unknown | Cy5 | 30.8
B10 | IAC | Unknown | Cy5 | 37.54 |  |  | B10 | IAC | Unknown | Cy5 | 31.07
B11 | IAC | Unknown | Cy5 | 38.3 |  |  | B11 | IAC | Unknown | Cy5 | 31.26
B12 | IAC | Unknown | Cy5 | 37.85 |  |  | B12 | IAC | Unknown | Cy5 | 30.96
C01 | IAC | Unknown | Cy5 |  |  |  | C01 | IAC | Unknown | Cy5 | 31.94 | 32.9809 | 0.30972
C02 | IAC | Unknown | Cy5 |  |  |  | C02 | IAC | Unknown | Cy5 | 32.94
C03 | IAC | Unknown | Cy5 |  |  |  | C03 | IAC | Unknown | Cy5 | 32.9
C04 | IAC | Unknown | Cy5 |  |  |  | C04 | IAC | Unknown | Cy5 | 33.48 |  |  | TubeName | Copy Number/ul | Log(Copy Number/ul) | Ct_STD_MEAN | Ct_STD_STDEV | 2*stdev | Ct_OPT_MEAN | Ct_OPT_STDEV | 2*stdev
C05 | IAC | Unknown | Cy5 |  |  |  | C05 | IAC | Unknown | Cy5 | 33.62 |  |  | std_1 | 120440000000 | 11.0808 | 12.705 | 0.51619 | 1.0324 | 37.5975 | 1.0812 | 2.1624
C06 | IAC | Unknown | Cy5 |  |  |  | C06 | IAC | Unknown | Cy5 | 33 |  |  | std_2 | 12044000000 | 10.0808 | 11.1333 | 8.6691 | 17.3381 | 35.6373 | 0.84798 | 1.696
C07 | IAC | Unknown | Cy5 |  |  |  | C07 | IAC | Unknown | Cy5 | 32.97 |  |  | std_2.1 | 6022000000 | 9.7797 | 15.4 | #DIV/0! | #DIV/0! | 35.4218 | 0.48979 | 0.97958
C08 | IAC | Unknown | Cy5 |  |  |  | C08 | IAC | Unknown | Cy5 | 32.63 |  |  | std_2.2 | 3011000000 | 9.4787 | 10.6533 | 2.117 | 0 | 36.444 | 0.85575 | 1.7115
C09 | IAC | Unknown | Cy5 |  |  |  | C09 | IAC | Unknown | Cy5 | 32.85 |  |  | std_3 | 1204400000 | 9.0808 | 12.63 | #DIV/0! | #DIV/0! | 35.8518 | 0.82064 | 1.6413
C10 | IAC | Unknown | Cy5 |  |  |  | C10 | IAC | Unknown | Cy5 | 32.97 |  |  | std_4 | 602200000 | 8.7797 | #DIV/0! | #DIV/0! | #DIV/0! | 37.262 | 0.99572 | 1.9914
C11 | IAC | Unknown | Cy5 |  |  |  | C11 | IAC | Unknown | Cy5 | 32.79 | [5 empty] | from OPT HI |  |  | from this assay (LO OPT)
C12 | IAC | Unknown | Cy5 |  |  |  | C12 | IAC | Unknown | Cy5 | 32.64 |  |  | TubeName | Copy Number/ul | Log(Copy Number/ul) | Ct_OPT_MEAN | Ct_OPT_STDEV | 2*stdev | Ct_OPT_MEAN | Ct_OPT_STDEV | 2*stdev
D01 | IAC | Unknown | Cy5 |  |  |  | D01 | IAC | Unknown | Cy5 | 34.11 | 34.8536 | 0.16378 | std_1 | 120440000000 | 11.0808 | 28.6309 | 0.1916 | 0.38319 | 37.5975 | 1.0812 | 2.1624
D02 | IAC | Unknown | Cy5 |  |  |  | D02 | IAC | Unknown | Cy5 | 35.04 |  |  | std_2 | 12044000000 | 10.0808 | 28.7864 | 0.36806 | 0.73611 | 35.6373 | 0.84798 | 1.696
D03 | IAC | Unknown | Cy5 |  |  |  | D03 | IAC | Unknown | Cy5 | 34.64 |  |  | std_2.1 | 6022000000 | 9.7797 | 28.6327 | 0.15447 | 0.30895 | 35.4218 | 0.48979 | 0.97958
D04 | IAC | Unknown | Cy5 |  |  |  | D04 | IAC | Unknown | Cy5 | 35.06 |  |  | std_2.2 | 3011000000 | 9.4787 | 28.5664 | 0.21855 | 0.43711 | 36.444 | 0.85575 | 1.7115
D05 | IAC | Unknown | Cy5 |  |  |  | D05 | IAC | Unknown | Cy5 | 34.78 |  |  | std_3 | 1204400000 | 9.0808 | 28.4464 | 0.13655 | 0.2731 | 35.8518 | 0.82064 | 1.6413
D06 | IAC | Unknown | Cy5 |  |  |  | D06 | IAC | Unknown | Cy5 | 35 |  |  | std_4 | 602200000 | 8.7797 | 28.3455 | 0.15175 | 0.30349 | 37.262 | 0.99572 | 1.9914
D07 | IAC | Unknown | Cy5 |  |  |  | D07 | IAC | Unknown | Cy5 | 34.78
D08 | IAC | Unknown | Cy5 |  |  |  | D08 | IAC | Unknown | Cy5 | 34.99
D09 | IAC | Unknown | Cy5 |  |  |  | D09 | IAC | Unknown | Cy5 | 34.71
D10 | IAC | Unknown | Cy5 |  |  |  | D10 | IAC | Unknown | Cy5 | 34.6
D11 | IAC | Unknown | Cy5 |  |  |  | D11 | IAC | Unknown | Cy5 | 34.95
D12 | IAC | Unknown | Cy5 |  |  |  | D12 | IAC | Unknown | Cy5 | 34.84
E01 | IAC | Unknown | Cy5 |  |  |  | E01 | IAC | Unknown | Cy5 | 36.03 | 37.2845 | 0.87985
E02 | IAC | Unknown | Cy5 |  |  |  | E02 | IAC | Unknown | Cy5 | 36.98
E03 | IAC | Unknown | Cy5 |  |  |  | E03 | IAC | Unknown | Cy5 | 37.71
E04 | IAC | Unknown | Cy5 |  |  |  | E04 | IAC | Unknown | Cy5 | 39.25
E05 | IAC | Unknown | Cy5 |  |  |  | E05 | IAC | Unknown | Cy5 | 37.05
E06 | IAC | Unknown | Cy5 |  |  |  | E06 | IAC | Unknown | Cy5 | 36.53
E07 | IAC | Unknown | Cy5 |  |  |  | E07 | IAC | Unknown | Cy5 | 36.93
E08 | IAC | Unknown | Cy5 |  |  |  | E08 | IAC | Unknown | Cy5 | 36.89
E09 | IAC | Unknown | Cy5 |  |  |  | E09 | IAC | Unknown | Cy5 | 36.9
E10 | IAC | Unknown | Cy5 |  |  |  | E10 | IAC | Unknown | Cy5 | 37.13
E11 | IAC | Unknown | Cy5 |  |  |  | E11 | IAC | Unknown | Cy5 | 36.26
E12 | IAC | Unknown | Cy5 |  |  |  | E12 | IAC | Unknown | Cy5 | 38.5
F01 | IAC | Unknown | Cy5 |  |  |  | F01 | IAC | Unknown | Cy5
F02 | IAC | Unknown | Cy5 |  |  |  | F02 | IAC | Unknown | Cy5
F03 | IAC | Unknown | Cy5 |  |  |  | F03 | IAC | Unknown | Cy5
F04 | IAC | Unknown | Cy5 |  |  |  | F04 | IAC | Unknown | Cy5
F05 | IAC | Unknown | Cy5 |  |  |  | F05 | IAC | Unknown | Cy5
F06 | IAC | Unknown | Cy5 |  |  |  | F06 | IAC | Unknown | Cy5
F07 | IAC | Unknown | Cy5 |  |  |  | F07 | IAC | Unknown | Cy5 |  |  |  |  | TubeName | std_1 | std_2 | std_2.1 | std_2.2 | std_3 | std_4
F08 | IAC | Unknown | Cy5 |  |  |  | F08 | IAC | Unknown | Cy5 |  |  |  |  | Copy Number/ul | 120440000000 | 12044000000 | 6022000000 | 3011000000 | 1204400000 | 602200000
F09 | IAC | Unknown | Cy5 |  |  |  | F09 | IAC | Unknown | Cy5 | [5 empty] | 36.52 | 34.6 | 35.2 | 36.78 | 35.12 | 37.65
F10 | IAC | Unknown | Cy5 |  |  |  | F10 | IAC | Unknown | Cy5 | [5 empty] | 38.57 | 34.98 | 35.05 | 36.33 | 35.75 | 37.52
F11 | IAC | Unknown | Cy5 |  |  |  | F11 | IAC | Unknown | Cy5 | [5 empty] | 37.68 | 34.26 | 35.56 | 37.19 | 35.89 | 36.35
F12 | IAC | Unknown | Cy5 |  |  |  | F12 | IAC | Unknown | Cy5 | [5 empty] | 37.35 | 36.47 | 35.64 | 38.4 | 35.79 | 38.63
G01 | IAC | Unknown | Cy5 | 32.7 |  |  | G01 | IAC | Unknown | Cy5 | 26.21 |  |  |  |  | 36.53 | 35.56 | 35.75 | 35.73 | 37.23 | 35.57
G02 | IAC | Unknown | Cy5 | 32.91 |  |  | G02 | IAC | Unknown | Cy5 | 26.51 |  |  |  |  | 36.37 | 37.24 | 36.6 | 36.35 | 36.28 | 37.68
G03 | IAC | Unknown | Cy5 | 32.97 |  |  | G03 | IAC | Unknown | Cy5 | 26.52 |  |  |  |  | 39.07 | 35.32 | 35.3 | 35.42 | 35.14 | 36.6
G04 | IAC | Unknown | Cy5 | 36.88 |  |  | G04 | IAC | Unknown | Cy5 | 30.76 |  |  |  |  | 38.69 | 35.87 | 35.2 | 36.28 | 35.62 | 37.2
G05 | IAC | Unknown | Cy5 | 37.5 |  |  | G05 | IAC | Unknown | Cy5 | 31.32 | [5 empty] | 36.22 | 35.56 | 35.81 | 37.36 | 38.74
G06 | IAC | Unknown | Cy5 | 37.32 |  |  | G06 | IAC | Unknown | Cy5 | 30.91 | [5 empty] | 35.75 | 34.88 | 36.15 | 35.34 | 36.68
G07 | IAC | Unknown | Cy5 |  |  |  | G07 | IAC | Unknown | Cy5 | 33.11 | [5 empty] | 35.74 | 34.9 |  | 34.85
G08 | IAC | Unknown | Cy5 | 39.65 |  |  | G08 | IAC | Unknown | Cy5 | 32.65
G09 | IAC | Unknown | Cy5 | 39.69 |  |  | G09 | IAC | Unknown | Cy5 | 32.76
G10 | IAC | Unknown | Cy5 |  |  |  | G10 | IAC | Unknown | Cy5 | 34.43
G11 | IAC | Unknown | Cy5 |  |  |  | G11 | IAC | Unknown | Cy5 | 35.15
G12 | IAC | Unknown | Cy5 |  |  |  | G12 | IAC | Unknown | Cy5 | 34.72
H01 | IAC | Unknown | Cy5 |  |  |  | H01 | IAC | Unknown | Cy5 | 35.45
H02 | IAC | Unknown | Cy5 |  |  |  | H02 | IAC | Unknown | Cy5 | 35.01
H03 | IAC | Unknown | Cy5 |  |  |  | H03 | IAC | Unknown | Cy5 | 34.91
H04 | IAC | Unknown | Cy5 |  |  |  | H04 | IAC | Unknown | Cy5 | 37.02
H05 | IAC | Unknown | Cy5 |  |  |  | H05 | IAC | Unknown | Cy5 | 36.88
H06 | IAC | Unknown | Cy5 |  |  |  | H06 | IAC | Unknown | Cy5 | 37.49
H07 | IAC | Unknown | Cy5 |  |  |  | H07 | IAC | Unknown | Cy5
H08 | IAC | Unknown | Cy5 |  |  |  | H08 | IAC | Unknown | Cy5
H09 | IAC | Unknown | Cy5 |  |  |  | H09 | IAC | Unknown | Cy5
H10 | IAC | Unknown | Cy5 |  |  |  | H10 | IAC | Negative | Cy5
H11 | IAC | Unknown | Cy5 |  |  |  | H11 | IAC | Negative | Cy5
H12 | IAC | Unknown | Cy5 |  |  |  | H12 | IAC | Negative | Cy5
A01 | Leg pneumo | Unknown | FAM |  | 12.705 | 0.51619 | A01 | Leg pneumo | Unknown | FAM | 35.9 | 37.5975 | 1.0812
A02 | Leg pneumo | Unknown | FAM |  |  |  | A02 | Leg pneumo | Unknown | FAM | 36.52
A03 | Leg pneumo | Unknown | FAM |  |  |  | A03 | Leg pneumo | Unknown | FAM | 38.57
A04 | Leg pneumo | Unknown | FAM |  |  |  | A04 | Leg pneumo | Unknown | FAM | 37.68
A05 | Leg pneumo | Unknown | FAM |  |  |  | A05 | Leg pneumo | Unknown | FAM | 37.35
A06 | Leg pneumo | Unknown | FAM |  |  |  | A06 | Leg pneumo | Unknown | FAM
A07 | Leg pneumo | Unknown | FAM | 13.07 |  |  | A07 | Leg pneumo | Unknown | FAM | 36.53
A08 | Leg pneumo | Unknown | FAM |  |  |  | A08 | Leg pneumo | Unknown | FAM | 36.37
A09 | Leg pneumo | Unknown | FAM |  |  |  | A09 | Leg pneumo | Unknown | FAM
A10 | Leg pneumo | Unknown | FAM |  |  |  | A10 | Leg pneumo | Unknown | FAM | 39.07
A11 | Leg pneumo | Unknown | FAM | 12.34 |  |  | A11 | Leg pneumo | Unknown | FAM | 38.69
A12 | Leg pneumo | Unknown | FAM |  |  |  | A12 | Leg pneumo | Unknown | FAM
B01 | Leg pneumo | Unknown | FAM |  | 11.1333 | 8.6691 | B01 | Leg pneumo | Unknown | FAM | 35.52 | 35.6373 | 0.84798
B02 | Leg pneumo | Unknown | FAM |  |  |  | B02 | Leg pneumo | Unknown | FAM | 34.6
B03 | Leg pneumo | Unknown | FAM | 3.41 |  |  | B03 | Leg pneumo | Unknown | FAM | 34.98
B04 | Leg pneumo | Unknown | FAM | 9.48 |  |  | B04 | Leg pneumo | Unknown | FAM | 34.26
B05 | Leg pneumo | Unknown | FAM | 20.51 |  |  | B05 | Leg pneumo | Unknown | FAM | 36.47
B06 | Leg pneumo | Unknown | FAM |  |  |  | B06 | Leg pneumo | Unknown | FAM | 35.56
B07 | Leg pneumo | Unknown | FAM |  |  |  | B07 | Leg pneumo | Unknown | FAM | 37.24
B08 | Leg pneumo | Unknown | FAM |  |  |  | B08 | Leg pneumo | Unknown | FAM | 35.32
B09 | Leg pneumo | Unknown | FAM |  |  |  | B09 | Leg pneumo | Unknown | FAM | 35.87
B10 | Leg pneumo | Unknown | FAM |  |  |  | B10 | Leg pneumo | Unknown | FAM | 36.22
B11 | Leg pneumo | Unknown | FAM |  |  |  | B11 | Leg pneumo | Unknown | FAM | 35.75
B12 | Leg pneumo | Unknown | FAM |  |  |  | B12 | Leg pneumo | Unknown | FAM | 35.74
C01 | Leg pneumo | Unknown | FAM |  | 15.4 | #DIV/0! | C01 | Leg pneumo | Unknown | FAM | 34.66 | 35.4218 | 0.48979
C02 | Leg pneumo | Unknown | FAM | 15.4 |  |  | C02 | Leg pneumo | Unknown | FAM | 35.2
C03 | Leg pneumo | Unknown | FAM |  |  |  | C03 | Leg pneumo | Unknown | FAM | 35.05
C04 | Leg pneumo | Unknown | FAM |  |  |  | C04 | Leg pneumo | Unknown | FAM | 35.56
C05 | Leg pneumo | Unknown | FAM |  |  |  | C05 | Leg pneumo | Unknown | FAM | 35.64
C06 | Leg pneumo | Unknown | FAM |  |  |  | C06 | Leg pneumo | Unknown | FAM | 35.75
C07 | Leg pneumo | Unknown | FAM |  |  |  | C07 | Leg pneumo | Unknown | FAM | 36.6
C08 | Leg pneumo | Unknown | FAM |  |  |  | C08 | Leg pneumo | Unknown | FAM | 35.3
C09 | Leg pneumo | Unknown | FAM |  |  |  | C09 | Leg pneumo | Unknown | FAM | 35.2
C10 | Leg pneumo | Unknown | FAM |  |  |  | C10 | Leg pneumo | Unknown | FAM | 35.56
C11 | Leg pneumo | Unknown | FAM |  |  |  | C11 | Leg pneumo | Unknown | FAM | 34.88
C12 | Leg pneumo | Unknown | FAM |  |  |  | C12 | Leg pneumo | Unknown | FAM | 34.9
D01 | Leg pneumo | Unknown | FAM |  | 10.6533 | 2.117 | D01 | Leg pneumo | Unknown | FAM | 35.67 | 36.444 | 0.85575
D02 | Leg pneumo | Unknown | FAM |  |  |  | D02 | Leg pneumo | Unknown | FAM | 36.78
D03 | Leg pneumo | Unknown | FAM | 8.21 |  |  | D03 | Leg pneumo | Unknown | FAM | 36.33
D04 | Leg pneumo | Unknown | FAM |  |  |  | D04 | Leg pneumo | Unknown | FAM | 37.19
D05 | Leg pneumo | Unknown | FAM |  |  |  | D05 | Leg pneumo | Unknown | FAM
D06 | Leg pneumo | Unknown | FAM | 11.81 |  |  | D06 | Leg pneumo | Unknown | FAM | 38.4
D07 | Leg pneumo | Unknown | FAM | 11.94 |  |  | D07 | Leg pneumo | Unknown | FAM | 35.73
D08 | Leg pneumo | Unknown | FAM |  |  |  | D08 | Leg pneumo | Unknown | FAM | 36.35
D09 | Leg pneumo | Unknown | FAM |  |  |  | D09 | Leg pneumo | Unknown | FAM | 35.42
D10 | Leg pneumo | Unknown | FAM |  |  |  | D10 | Leg pneumo | Unknown | FAM | 36.28
D11 | Leg pneumo | Unknown | FAM |  |  |  | D11 | Leg pneumo | Unknown | FAM | 35.81
D12 | Leg pneumo | Unknown | FAM |  |  |  | D12 | Leg pneumo | Unknown | FAM | 36.15
E01 | Leg pneumo | Unknown | FAM |  | 12.63 | #DIV/0! | E01 | Leg pneumo | Unknown | FAM | 34.38 | 35.8518 | 0.82064
E02 | Leg pneumo | Unknown | FAM |  |  |  | E02 | Leg pneumo | Unknown | FAM | 35.12
E03 | Leg pneumo | Unknown | FAM |  |  |  | E03 | Leg pneumo | Unknown | FAM | 35.75
E04 | Leg pneumo | Unknown | FAM |  |  |  | E04 | Leg pneumo | Unknown | FAM | 35.89
E05 | Leg pneumo | Unknown | FAM |  |  |  | E05 | Leg pneumo | Unknown | FAM | 35.79
E06 | Leg pneumo | Unknown | FAM |  |  |  | E06 | Leg pneumo | Unknown | FAM | 37.23
E07 | Leg pneumo | Unknown | FAM | 12.63 |  |  | E07 | Leg pneumo | Unknown | FAM | 36.28
E08 | Leg pneumo | Unknown | FAM |  |  |  | E08 | Leg pneumo | Unknown | FAM | 35.14
E09 | Leg pneumo | Unknown | FAM |  |  |  | E09 | Leg pneumo | Unknown | FAM | 35.62
E10 | Leg pneumo | Unknown | FAM |  |  |  | E10 | Leg pneumo | Unknown | FAM | 37.36
E11 | Leg pneumo | Unknown | FAM |  |  |  | E11 | Leg pneumo | Unknown | FAM | 35.34
E12 | Leg pneumo | Unknown | FAM |  |  |  | E12 | Leg pneumo | Unknown | FAM | 34.85
F01 | Leg pneumo | Unknown | FAM |  | #DIV/0! | #DIV/0! | F01 | Leg pneumo | Unknown | FAM | 37.77 | 37.262 | 0.99572
F02 | Leg pneumo | Unknown | FAM |  |  |  | F02 | Leg pneumo | Unknown | FAM | 37.65
F03 | Leg pneumo | Unknown | FAM |  |  |  | F03 | Leg pneumo | Unknown | FAM | 37.52
F04 | Leg pneumo | Unknown | FAM |  |  |  | F04 | Leg pneumo | Unknown | FAM | 36.35
F05 | Leg pneumo | Unknown | FAM |  |  |  | F05 | Leg pneumo | Unknown | FAM | 38.63
F06 | Leg pneumo | Unknown | FAM |  |  |  | F06 | Leg pneumo | Unknown | FAM | 35.57
F07 | Leg pneumo | Unknown | FAM |  |  |  | F07 | Leg pneumo | Unknown | FAM | 37.68
F08 | Leg pneumo | Unknown | FAM |  |  |  | F08 | Leg pneumo | Unknown | FAM | 36.6
F09 | Leg pneumo | Unknown | FAM |  |  |  | F09 | Leg pneumo | Unknown | FAM | 37.2
F10 | Leg pneumo | Unknown | FAM |  |  |  | F10 | Leg pneumo | Unknown | FAM
F11 | Leg pneumo | Unknown | FAM |  |  |  | F11 | Leg pneumo | Unknown | FAM | 38.74
F12 | Leg pneumo | Unknown | FAM |  |  |  | F12 | Leg pneumo | Unknown | FAM | 36.68
G01 | Leg pneumo | Unknown | FAM |  |  |  | G01 | Leg pneumo | Unknown | FAM | 38.88
G02 | Leg pneumo | Unknown | FAM |  |  |  | G02 | Leg pneumo | Unknown | FAM | 39.71
G03 | Leg pneumo | Unknown | FAM |  |  |  | G03 | Leg pneumo | Unknown | FAM | 38.04
G04 | Leg pneumo | Unknown | FAM | 13.23 |  |  | G04 | Leg pneumo | Unknown | FAM | 35.66
G05 | Leg pneumo | Unknown | FAM |  |  |  | G05 | Leg pneumo | Unknown | FAM | 37
G06 | Leg pneumo | Unknown | FAM |  |  |  | G06 | Leg pneumo | Unknown | FAM | 37.01
G07 | Leg pneumo | Unknown | FAM |  |  |  | G07 | Leg pneumo | Unknown | FAM | 35.31
G08 | Leg pneumo | Unknown | FAM |  |  |  | G08 | Leg pneumo | Unknown | FAM | 35.03
G09 | Leg pneumo | Unknown | FAM |  |  |  | G09 | Leg pneumo | Unknown | FAM | 35.75
G10 | Leg pneumo | Unknown | FAM |  |  |  | G10 | Leg pneumo | Unknown | FAM | 35.83
G11 | Leg pneumo | Unknown | FAM |  |  |  | G11 | Leg pneumo | Unknown | FAM | 36.92
G12 | Leg pneumo | Unknown | FAM |  |  |  | G12 | Leg pneumo | Unknown | FAM | 35.47
H01 | Leg pneumo | Unknown | FAM |  |  |  | H01 | Leg pneumo | Unknown | FAM | 35.24
H02 | Leg pneumo | Unknown | FAM | 12.42 |  |  | H02 | Leg pneumo | Unknown | FAM | 35.2
H03 | Leg pneumo | Unknown | FAM |  |  |  | H03 | Leg pneumo | Unknown | FAM
H04 | Leg pneumo | Unknown | FAM |  |  |  | H04 | Leg pneumo | Unknown | FAM | 34.81
H05 | Leg pneumo | Unknown | FAM |  |  |  | H05 | Leg pneumo | Unknown | FAM | 37.48
H06 | Leg pneumo | Unknown | FAM |  |  |  | H06 | Leg pneumo | Unknown | FAM | 36.29
H07 | Leg pneumo | Unknown | FAM |  |  |  | H07 | Leg pneumo | Unknown | FAM
H08 | Leg pneumo | Unknown | FAM | 39.61 |  |  | H08 | Leg pneumo | Unknown | FAM | 37.56
H09 | Leg pneumo | Unknown | FAM |  |  |  | H09 | Leg pneumo | Unknown | FAM | 35.79
H10 | Leg pneumo | Unknown | FAM |  |  |  | H10 | Leg pneumo | Unknown | FAM | 38.11
H11 | Leg pneumo | Unknown | FAM |  |  |  | H11 | Leg pneumo | Unknown | FAM
H12 | Leg pneumo | Unknown | FAM | 11.79 |  |  | H12 | Leg pneumo | Unknown | FAM
Note on S29: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Poor fluorescence in these wells. Not convinced by any of the Cqs.
CONCLUSIONS
IAC Cq decreases ~1.5 Cq when Leg pneumo decreases from 1e6 copies/ul to ~10 copies/ul. This is statistically significant only at high concentrations of IAC. Leg pneumo concentration insignificant when IAC used at recommended concentration: 3e9 copies/ul.

## Why_so_high_mmix
This master mix given to Raquell
from this file: "C:\Users\HarleyKing\OneDrive - LuminUltra Technologies Ltd\Documents\Exp800-automation\Exp800.09 Optimizing Legionella pneumophila\Legionella recipes wet 04_12_21.xlsx"
OPTIMIZED RECIPE | [6 empty] | 20210430_Testing OPT recipe vs LEG recipe
Component | Desired |  | Per 20 uL |  | Multiple Reactions |  | Component | Desired | Units | PROBE RXN | 115.2 | Component | Desired |  | Per 20 uL
DotAF | 0.4 | uM | 0.08 | uL | 21 | uL | mmix |  |  | 9.3956 | 1082.368 | DotAF | 0.4 | uM | 0.08 | uL
DotAR | 0.4 | uM | 0.08 | uL | 21 | uL | Leg pneumo "std_5" |  |  | 0.2 | 23.04 | DotAR | 0.4 | uM | 0.08 | uL
Leg pne DotA MGB probe | 0.5 | uM | 0 | uL | 0 | uL | [5 empty] | Leg pne DotA MGB probe | 0.5 | uM | 0.1 | uL
IAC Forward | 0.05 | uM | 0.01 | uL | 2.625 | uL | water |  |  | 8.4044 | 968.192 | IAC Forward | 0.25 | uM | 0 | uL
IAC Reverse | 0.05 | uM | 0.01 | uL | 2.625 | uL | TOTAL |  |  | 18 | 2073.6 | IAC Reverse | 0.25 | uM | 0 | uL
IAC Probe RED (CY5) | 0.4 | uM | 0.08 | uL | 21 | uL | [5 empty] | IAC Probe RED (CY5) | 0.5 | uM | 0 | uL
Deer_IAC_Ultramer (1 x 10e3 copies/uL) |  |  | 0 | uL | 0 | uL | Used same IAC standards set from previous experiments. No more vol in std_2.1 and std_2.2 |  |  |  |  | Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL
dNTP Mix | 0.25 | mM | 0.32 | uL | 84 | uL | Row | IAC Std |  |  |  | dNTP Mix | 0.25 | mM | 0.32 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 13.125 | uL | A | std_1 |  |  |  | AptaTaq | 2.5 | Units | 0.05 | uL
KAPA3G | 1 | Units | 0.033333 | uL | 8.75 | uL | B | std_2 |  |  |  | KAPA3G | 1 | Units | 0.033333 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 18.375 | uL | C | std_2.3 |  |  |  | Mg2+ | 3.5 | mM | 0.07 | uL
Trehalose | 10 | Percent | 4 | uL | 1050 | uL | D | std_3 |  |  |  | Trehalose | 10 | Percent | 4 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 52.5 | uL | E | std_5 | Meant to add std_4 here but didn't update OT-2 settings |  |  | BSA | 0.5 | mg/mL | 0.2 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 525 | uL | F | std_6 |  |  |  | PCR Buffer 10X | 1 | x | 2 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 583.3333 | uL | [5 empty] | 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL
Water |  |  | 0.24 | uL | 63 | uL | [5 empty] | Dextran | 0 | Percent | 0 | uL
[12 empty] | Water |  |  | 0 | uL
|  | Total: | 9.3956 | uL | 2466.3333 | uL
[14 empty] | Total: | 9.1556 | uL
LEG spp. RECIPE
Component | Desired |  | Per 20 uL |  | Multiple Reactions
LegGF | 1 | uM | 0.2 | uL | 52.5 | uL
LegGR | 1 | uM | 0.2 | uL | 52.5 | uL
Leg spp MGB probe | 0.5 | uM | 0.1 | uL | 0 | uL
IAC Forward | 0.05 | uM | 0.01 | uL | 2.625 | uL
IAC Reverse | 0.05 | uM | 0.01 | uL | 2.625 | uL
IAC Probe RED (CY5) | 0.4 | uM | 0.08 | uL | 21 | uL
Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL | 0 | uL
dNTP Mix | 0.25 | mM | 0.32 | uL | 84 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 13.125 | uL
KAPA3G | 1 | Units | 0.033333 | uL | 8.75 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 18.375 | uL
Trehalose | 10 | Percent | 4 | uL | 1050 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 52.5 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 525 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 583.3333 | uL
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 9.3956 | uL | 2466.3333 | uL
20210507
Master mixes made by Raquel after 4/26 have a problem. Diagnosed to be dNTP and Mg2+ concentration.
She made a Leg spp mastermix that I also believe needs dNTP and Mg adjustment.
Component | Desired | Units | PROBE RXN | 115.2
mmix |  |  | 9.3956 | 1082.368
Leg spp "std_5" |  |  | 0.2 | 23.04
dNTP | 0.25 | mM | 0.8 | 92.16
Mg2+ | 3.5 | mM | 2.8 | 322.56
water |  |  | 4.8044 | 553.472
TOTAL |  |  | 18 | 2073.6
This worked splendidly well.
Didn't observe Leg spp. Fluorescence but expected, too. Did Raquel leave out this probe, too?
Repeated same program and mastermix above with Leg pneumo mix.
Note on D8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This probe is empty. It has been re-ordered but didn't arrive in time.
Note on A66: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel provided me with 10ul of undiluted Leg spp positive control.
Note on D67: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I don't know why we multiple by 32 instead of 20.

## std_vs_opt
Well | Target | Property | Dye | Ct | Ct_Mean | Ct_StDev | Well | Target | Property | Dye | Ct | Ct_Mean | Ct_StDev
A01 | IAC | Unknown | CY5 |  |  |  | A01 | IAC | Unknown | CY5 | 24.62 | 25.4436 | 0.21049 | LEG spp. RECIPE | [8 empty] | OPTIMIZED RECIPE
A02 | IAC | Unknown | CY5 |  | #DIV/0! |  | A02 | IAC | Unknown | CY5 | 25.72 |  |  | Component | Desired |  | Per 20 uL |  |  | Multiple Reactions |  |  | Component | Desired |  | Per 20 uL |  | Multiple Reactions
A03 | IAC | Unknown | CY5 |  |  |  | A03 | IAC | Unknown | CY5 | 25.64 |  |  | LegGF | 1 | uM | 0.2 | uL |  | 52.5 | uL |  | DotAF | 0.4 | uM | 0.08 | uL | 21 | uL
A04 | IAC | Unknown | CY5 |  |  |  | A04 | IAC | Unknown | CY5 | 25.64 |  |  | LegGR | 1 | uM | 0.2 | uL |  | 52.5 | uL |  | DotAR | 0.4 | uM | 0.08 | uL | 21 | uL
A05 | IAC | Unknown | CY5 |  |  |  | A05 | IAC | Unknown | CY5 | 25.64 |  |  | Leg spp MGB probe | 0.5 | uM | 0.1 | uL |  | 0 | uL |  | Leg pne DotA MGB probe | 0.5 | uM | 0 | uL | 0 | uL
A06 | IAC | Unknown | CY5 |  |  |  | A06 | IAC | Unknown | CY5 | 25.31 |  |  | IAC Forward | 0.05 | uM | 0.01 | uL |  | 2.625 | uL |  | IAC Forward | 0.05 | uM | 0.01 | uL | 2.625 | uL
A07 | IAC | Unknown | CY5 |  |  |  | A07 | IAC | Unknown | CY5 | 25.63 |  |  | IAC Reverse | 0.05 | uM | 0.01 | uL |  | 2.625 | uL |  | IAC Reverse | 0.05 | uM | 0.01 | uL | 2.625 | uL
A08 | IAC | Unknown | CY5 |  |  |  | A08 | IAC | Unknown | CY5 | 25.14 |  |  | IAC Probe RED (CY5) | 0.4 | uM | 0.08 | uL |  | 21 | uL |  | IAC Probe RED (CY5) | 0.4 | uM | 0.08 | uL | 21 | uL
A09 | IAC | Unknown | CY5 |  |  |  | A09 | IAC | Unknown | CY5 | 25.34 |  |  | Deer_IAC_Ultramer (1 x 10e3 copies/uL) | 100 |  | 0 | uL |  | 0 | uL |  | Deer_IAC_Ultramer (1 x 10e3 copies/uL) |  |  | 0 | uL | 0 | uL
A10 | IAC | Unknown | CY5 |  |  |  | A10 | IAC | Unknown | CY5 | 25.34 |  |  | dNTP Mix | 0.25 | mM | 0.32 | uL |  | 84 | uL |  | dNTP Mix | 0.25 | mM | 0.32 | uL | 84 | uL
A11 | IAC | Unknown | CY5 |  |  |  | A11 | IAC | Unknown | CY5 | 25.21 |  |  | AptaTaq | 2.5 | Units | 0.05 | uL |  | 13.125 | uL |  | AptaTaq | 2.5 | Units | 0.05 | uL | 13.125 | uL
A12 | IAC | Unknown | CY5 |  |  |  | A12 | IAC | Unknown | CY5 | 25.27 |  |  | KAPA3G | 1 | Units | 0.033333 | uL |  | 8.75 | uL |  | KAPA3G | 1 | Units | 0.033333 | uL | 8.75 | uL
B01 | IAC | Unknown | CY5 | 7.83 | 9.4864 | 0.45123 | B01 | IAC | Unknown | CY5 | 28.99 | 30.0682 | 0.16491 | Mg2+ | 3.5 | mM | 0.07 | uL |  | 18.375 | uL |  | Mg2+ | 3.5 | mM | 0.07 | uL | 18.375 | uL
B02 | IAC | Unknown | CY5 | 8.97 |  |  | B02 | IAC | Unknown | CY5 | 30.21 |  |  | Trehalose | 10 | Percent | 4 | uL |  | 1050 | uL |  | Trehalose | 10 | Percent | 4 | uL | 1050 | uL
B03 | IAC | Unknown | CY5 | 10.06 |  |  | B03 | IAC | Unknown | CY5 | 29.99 |  |  | BSA | 0.5 | mg/mL | 0.2 | uL |  | 52.5 | uL |  | BSA | 0.5 | mg/mL | 0.2 | uL | 52.5 | uL
B04 | IAC | Unknown | CY5 | 9.98 |  |  | B04 | IAC | Unknown | CY5 | 30.06 |  |  | PCR Buffer 10X | 1 | x | 2 | uL |  | 525 | uL |  | PCR Buffer 10X | 1 | x | 2 | uL | 525 | uL
B05 | IAC | Unknown | CY5 | 9.77 |  |  | B05 | IAC | Unknown | CY5 | 30.05 |  |  | 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL |  | 583.3333 | uL |  | 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 583.3333 | uL
B06 | IAC | Unknown | CY5 | 9.78 |  |  | B06 | IAC | Unknown | CY5 | 30.04 |  |  | Water |  |  | 0 | uL |  | 0 | uL |  | Water |  |  | 0.24 | uL | 63 | uL
B07 | IAC | Unknown | CY5 | 9.88 |  |  | B07 | IAC | Unknown | CY5 | 30.42
B08 | IAC | Unknown | CY5 | 9.49 |  |  | B08 | IAC | Unknown | CY5 | 29.83 |  |  |  |  | Total: | 9.3956 | uL |  | 2466.3333 | uL |  |  |  | Total: | 9.3956 | uL | 2466.3333 | uL
B09 | IAC | Unknown | CY5 | 9.5 |  |  | B09 | IAC | Unknown | CY5 | 29.87
B10 | IAC | Unknown | CY5 | 8.83 |  |  | B10 | IAC | Unknown | CY5 | 30.16
B11 | IAC | Unknown | CY5 | 9.21 |  |  | B11 | IAC | Unknown | CY5 | 30.15 |  |  |  | IDT DNA Data (IAC) |  |  |  |  | Leg spp |  |  | Leg pneumo
B12 | IAC | Unknown | CY5 | 8.88 |  |  | B12 | IAC | Unknown | CY5 | 29.97 |  |  | TubeName | Conc (uM) | pM | Copy Number | Copy Number/rxn | Log(Copy#) | Ct_Mean | Ct_STDEV | 2*STDEV | Ct_Mean | Ct_STDEV | 2*STDEV | Diff
C01 | IAC | Unknown | CY5 | 13.48 | 15.1618 | 0.244 | C01 | IAC | Unknown | CY5 | 37.24 | 38.9864 | 1.2785 | IDT Source | 10 | 10000000
C02 | IAC | Unknown | CY5 | 15.12 |  |  | C02 | IAC | Unknown | CY5 | 38.87 |  |  | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808 | too high |  |  | 25.4436 | 0.21049 | 0.42098
C03 | IAC | Unknown | CY5 | 15.47 |  |  | C03 | IAC | Unknown | CY5 | 38.74 |  |  | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | 9.4864 | 0.45123 | 0.90245 | 30.0682 | 0.16491 | 0.32983 | 20.5818
C04 | IAC | Unknown | CY5 | 15.38 |  |  | C04 | IAC | Unknown | CY5 | 40.61 |  |  | std_3 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 9.0808 | 15.1618 | 0.244 | 0.488 | 38.9864 | 1.2785 | 2.557 | 23.8245
C05 | IAC | Unknown | CY5 | 15.4 |  |  | C05 | IAC | Unknown | CY5 | 39.23 |  |  | std_4 | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808 | 22.0545 | 0.72825 | 1.4565
C06 | IAC | Unknown | CY5 | 15.25 |  |  | C06 | IAC | Unknown | CY5 | 38.08 |  |  | std_5 | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808 | 27.5018 | 0.28389 | 0.56779
C07 | IAC | Unknown | CY5 | 15.31 |  |  | C07 | IAC | Unknown | CY5 | 38.84 |  |  | std_6 | 1e-06 | 1 | 602200000 | 1204400 | 6.0808 | 28.3445 | 0.35551 | 0.71102
C08 | IAC | Unknown | CY5 | 15 |  |  | C08 | IAC | Unknown | CY5 | 39.31 |  |  | std_7 | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808
C09 | IAC | Unknown | CY5 | 15.21 |  |  | C09 | IAC | Unknown | CY5 | 37.99 |  |  | std_8 | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808
C10 | IAC | Unknown | CY5 | 15.13 |  |  | C10 | IAC | Unknown | CY5 | 41.66 |  |  | std_9 | 1e-09 | 0.001 | 602200 | 1204.4 | 3.0808
C11 | IAC | Unknown | CY5 | 14.8 |  |  | C11 | IAC | Unknown | CY5 | 36.9 |  |  | std_10 | 1e-10 | 0.0001 | 60220 | 120.44 | 2.0808
C12 | IAC | Unknown | CY5 | 14.71 |  |  | C12 | IAC | Unknown | CY5 | 38.62 |  |  | std_11 | 1e-11 | 1e-05 | 6022 | 12.044 | 1.0808
D01 | IAC | Unknown | CY5 | 19.89 | 22.0545 | 0.72825 | D01 | IAC | Unknown | CY5 |  |  |  | std_12 | 1e-12 | 1e-06 | 602.2 | 1.2044 | 0.080771
D02 | IAC | Unknown | CY5 | 22.07 |  |  | D02 | IAC | Unknown | CY5 |  |  |  | std_13 | 1e-13 | 1e-07 | 60.22 | 0.12044 | -0.91923
D03 | IAC | Unknown | CY5 | 22.92 |  |  | D03 | IAC | Unknown | CY5 |  |  |  | std_14 | 1e-14 | 1e-08 | 6.022 | 0.012044 | -1.9192
D04 | IAC | Unknown | CY5 | 23.51 |  |  | D04 | IAC | Unknown | CY5 |  |  |  | std_15 | 1e-15 | 1e-09 | 0.6022 | 0.0012044 | -2.9192
D05 | IAC | Unknown | CY5 | 22.49 |  |  | D05 | IAC | Unknown | CY5 |  |  |  | std_16 | 1e-16 | 1e-10 | 0.06022 | 0.00012044 | -3.9192
D06 | IAC | Unknown | CY5 | 22.45 |  |  | D06 | IAC | Unknown | CY5 |  |  |  | std_17 | 1e-17 | 1e-11 | 0.006022 | 1.2044e-05 | -4.9192
D07 | IAC | Unknown | CY5 | 21.47 |  |  | D07 | IAC | Unknown | CY5 |  |  |  | std_18 | 1e-18 | 1e-12 | 0.0006022 | 1.2044e-06 | -5.9192
D08 | IAC | Unknown | CY5 | 21.86 |  |  | D08 | IAC | Unknown | CY5
D09 | IAC | Unknown | CY5 | 21.9 |  |  | D09 | IAC | Unknown | CY5
D10 | IAC | Unknown | CY5 | 21.43 |  |  | D10 | IAC | Unknown | CY5
D11 | IAC | Unknown | CY5 | 21.28 |  |  | D11 | IAC | Unknown | CY5 |  |  |  | c1 | v1 | c2 | v2
D12 | IAC | Unknown | CY5 | 21.22 |  |  | D12 | IAC | Unknown | CY5 |  |  |  | 602200000000.0001 | 10 | 4014666666.6667 | 1500
E01 | IAC | Unknown | CY5 | 26.37 | 27.5018 | 0.28389 | E01 | IAC | Unknown | CY5
E02 | IAC | Unknown | CY5 | 27.63 |  |  | E02 | IAC | Unknown | CY5
E03 | IAC | Unknown | CY5 | 27.42 |  |  | E03 | IAC | Unknown | CY5
E04 | IAC | Unknown | CY5 | 26.99 |  |  | E04 | IAC | Unknown | CY5
E05 | IAC | Unknown | CY5 | 27.42 |  |  | E05 | IAC | Unknown | CY5
E06 | IAC | Unknown | CY5 | 27.51 |  |  | E06 | IAC | Unknown | CY5
E07 | IAC | Unknown | CY5 | 27.8 |  |  | E07 | IAC | Unknown | CY5
E08 | IAC | Unknown | CY5 | 27.27 |  |  | E08 | IAC | Unknown | CY5
E09 | IAC | Unknown | CY5 | 27.61 |  |  | E09 | IAC | Unknown | CY5
E10 | IAC | Unknown | CY5 | 27.79 |  |  | E10 | IAC | Unknown | CY5
E11 | IAC | Unknown | CY5 | 27.17 |  |  | E11 | IAC | Unknown | CY5
E12 | IAC | Unknown | CY5 | 27.91 |  |  | E12 | IAC | Unknown | CY5
F01 | IAC | Unknown | CY5 | 27.06 | 28.3445 | 0.35551 | F01 | IAC | Unknown | CY5
F02 | IAC | Unknown | CY5 | 28.36 |  |  | F02 | IAC | Unknown | CY5
F03 | IAC | Unknown | CY5 | 28.53 |  |  | F03 | IAC | Unknown | CY5
F04 | IAC | Unknown | CY5 | 28.29 |  |  | F04 | IAC | Unknown | CY5
F05 | IAC | Unknown | CY5 | 29.14 |  |  | F05 | IAC | Unknown | CY5
F06 | IAC | Unknown | CY5 | 28.18 |  |  | F06 | IAC | Unknown | CY5
F07 | IAC | Unknown | CY5 | 28.51 |  |  | F07 | IAC | Unknown | CY5
F08 | IAC | Unknown | CY5 | 27.78 |  |  | F08 | IAC | Unknown | CY5
F09 | IAC | Unknown | CY5 | 28.39 |  |  | F09 | IAC | Unknown | CY5
F10 | IAC | Unknown | CY5 | 28.5 |  |  | F10 | IAC | Unknown | CY5
F11 | IAC | Unknown | CY5 | 28.17 |  |  | F11 | IAC | Unknown | CY5
F12 | IAC | Unknown | CY5 | 27.94 |  |  | F12 | IAC | Unknown | CY5
G01 | IAC | Unknown | CY5 |  | #DIV/0! | #DIV/0! | G01 | IAC | Unknown | CY5 | 24.98 | 25.2 | 0.20664
G02 | IAC | Unknown | CY5 |  |  |  | G02 | IAC | Unknown | CY5 | 25.39
G03 | IAC | Unknown | CY5 |  |  |  | G03 | IAC | Unknown | CY5 | 25.23
G04 | IAC | Unknown | CY5 | 9.48 | 9.6633 | 0.16503 | G04 | IAC | Unknown | CY5 | 29.69 | 29.7767 | 0.075719
G05 | IAC | Unknown | CY5 | 9.71 |  |  | G05 | IAC | Unknown | CY5 | 29.83
G06 | IAC | Unknown | CY5 | 9.8 |  |  | G06 | IAC | Unknown | CY5 | 29.81
G07 | IAC | Unknown | CY5 | 14.85 | 14.8267 | 0.097125 | G07 | IAC | Unknown | CY5 | 37.82 | 37.66 | 0.14
G08 | IAC | Unknown | CY5 | 14.72 |  |  | G08 | IAC | Unknown | CY5 | 37.6
G09 | IAC | Unknown | CY5 | 14.91 |  |  | G09 | IAC | Unknown | CY5 | 37.56
G10 | IAC | Unknown | CY5 | 21.32 | 21.7833 | 0.44613 | G10 | IAC | Unknown | CY5
G11 | IAC | Unknown | CY5 | 22.21 |  |  | G11 | IAC | Unknown | CY5
G12 | IAC | Unknown | CY5 | 21.82 |  |  | G12 | IAC | Unknown | CY5
H01 | IAC | Unknown | CY5 | 26.53 | 26.8933 | 0.38109 | H01 | IAC | Unknown | CY5
H02 | IAC | Unknown | CY5 | 27.29 |  |  | H02 | IAC | Unknown | CY5
H03 | IAC | Unknown | CY5 | 26.86 |  |  | H03 | IAC | Unknown | CY5
H04 | IAC | Unknown | CY5 | 28.02 | 28.0333 | 0.3302 | H04 | IAC | Unknown | CY5
H05 | IAC | Unknown | CY5 | 27.71 |  |  | H05 | IAC | Unknown | CY5
H06 | IAC | Unknown | CY5 | 28.37 |  |  | H06 | IAC | Unknown | CY5
H07 | IAC | Unknown | CY5 | 28.85 | 28.89 | 0.045826 | H07 | IAC | Unknown | CY5
H08 | IAC | Unknown | CY5 | 28.88 |  |  | H08 | IAC | Unknown | CY5
H09 | IAC | Unknown | CY5 | 28.94 |  |  | H09 | IAC | Unknown | CY5
H10 | IAC | Unknown | CY5 | 29.07 | 29.0867 | 0.047258 | H10 | IAC | Unknown | CY5
H11 | IAC | Unknown | CY5 | 29.05 |  |  | H11 | IAC | Unknown | CY5
H12 | IAC | Unknown | CY5 | 29.14 |  |  | H12 | IAC | Unknown | CY5
A01 | Leg spp | Unknown | FAM |  |  |  | A01 | Leg pneumo | Unknown | FAM
A02 | Leg spp | Unknown | FAM |  |  |  | A02 | Leg pneumo | Unknown | FAM
A03 | Leg spp | Unknown | FAM | 11.78 |  |  | A03 | Leg pneumo | Unknown | FAM
A04 | Leg spp | Unknown | FAM |  |  |  | A04 | Leg pneumo | Unknown | FAM
A05 | Leg spp | Unknown | FAM |  |  |  | A05 | Leg pneumo | Unknown | FAM
A06 | Leg spp | Unknown | FAM |  |  |  | A06 | Leg pneumo | Unknown | FAM
A07 | Leg spp | Unknown | FAM |  |  |  | A07 | Leg pneumo | Unknown | FAM
A08 | Leg spp | Unknown | FAM |  |  |  | A08 | Leg pneumo | Unknown | FAM
A09 | Leg spp | Unknown | FAM |  |  |  | A09 | Leg pneumo | Unknown | FAM | 7.26
A10 | Leg spp | Unknown | FAM |  |  |  | A10 | Leg pneumo | Unknown | FAM
A11 | Leg spp | Unknown | FAM |  |  |  | A11 | Leg pneumo | Unknown | FAM
A12 | Leg spp | Unknown | FAM |  |  |  | A12 | Leg pneumo | Unknown | FAM
B01 | Leg spp | Unknown | FAM |  |  |  | B01 | Leg pneumo | Unknown | FAM
B02 | Leg spp | Unknown | FAM |  |  |  | B02 | Leg pneumo | Unknown | FAM
B03 | Leg spp | Unknown | FAM |  |  |  | B03 | Leg pneumo | Unknown | FAM
B04 | Leg spp | Unknown | FAM |  |  |  | B04 | Leg pneumo | Unknown | FAM
B05 | Leg spp | Unknown | FAM |  |  |  | B05 | Leg pneumo | Unknown | FAM | [10 empty] | Neil says
B06 | Leg spp | Unknown | FAM |  |  |  | B06 | Leg pneumo | Unknown | FAM | [10 empty] | This is really crazy. Out of curiosity, I took the primers and probes and plugged them into Thermo's multiple primer analyzer: https://www.thermofisher.com/us/en/home/brands/thermo-scientific/molecular-biology/molecular-biology...
B07 | Leg spp | Unknown | FAM |  |  |  | B07 | Leg pneumo | Unknown | FAM | [10 empty] | One significant interaction came up between the forward Legionella primer and the IAC forward:
B08 | Leg spp | Unknown | FAM |  |  |  | B08 | Leg pneumo | Unknown | FAM
B09 | Leg spp | Unknown | FAM |  |  |  | B09 | Leg pneumo | Unknown | FAM | [10 empty] | While we are waiting for more Leg probe, would it make sense to run a couple tests to confirm which primer(s) may be causing the issue? For example, it may be interesting to know if only one of the Legionella primers is causing the issue (i.e. perhaps by running complete IAC assay with each Legionella primer together and individually). That could potentially suggest a way to fix the issue by changing the IAC sequence slightly?
B10 | Leg spp | Unknown | FAM |  |  |  | B10 | Leg pneumo | Unknown | FAM
B11 | Leg spp | Unknown | FAM |  |  |  | B11 | Leg pneumo | Unknown | FAM
B12 | Leg spp | Unknown | FAM |  |  |  | B12 | Leg pneumo | Unknown | FAM
C01 | Leg spp | Unknown | FAM |  |  |  | C01 | Leg pneumo | Unknown | FAM
C02 | Leg spp | Unknown | FAM |  |  |  | C02 | Leg pneumo | Unknown | FAM
C03 | Leg spp | Unknown | FAM | 37.38 |  |  | C03 | Leg pneumo | Unknown | FAM
C04 | Leg spp | Unknown | FAM |  |  |  | C04 | Leg pneumo | Unknown | FAM
C05 | Leg spp | Unknown | FAM |  |  |  | C05 | Leg pneumo | Unknown | FAM
C06 | Leg spp | Unknown | FAM |  |  |  | C06 | Leg pneumo | Unknown | FAM
C07 | Leg spp | Unknown | FAM |  |  |  | C07 | Leg pneumo | Unknown | FAM
C08 | Leg spp | Unknown | FAM |  |  |  | C08 | Leg pneumo | Unknown | FAM
C09 | Leg spp | Unknown | FAM |  |  |  | C09 | Leg pneumo | Unknown | FAM
C10 | Leg spp | Unknown | FAM |  |  |  | C10 | Leg pneumo | Unknown | FAM
C11 | Leg spp | Unknown | FAM |  |  |  | C11 | Leg pneumo | Unknown | FAM
C12 | Leg spp | Unknown | FAM |  |  |  | C12 | Leg pneumo | Unknown | FAM
D01 | Leg spp | Unknown | FAM |  |  |  | D01 | Leg pneumo | Unknown | FAM
D02 | Leg spp | Unknown | FAM |  |  |  | D02 | Leg pneumo | Unknown | FAM
D03 | Leg spp | Unknown | FAM |  |  |  | D03 | Leg pneumo | Unknown | FAM
D04 | Leg spp | Unknown | FAM |  |  |  | D04 | Leg pneumo | Unknown | FAM
D05 | Leg spp | Unknown | FAM |  |  |  | D05 | Leg pneumo | Unknown | FAM
D06 | Leg spp | Unknown | FAM |  |  |  | D06 | Leg pneumo | Unknown | FAM
D07 | Leg spp | Unknown | FAM |  |  |  | D07 | Leg pneumo | Unknown | FAM
D08 | Leg spp | Unknown | FAM |  |  |  | D08 | Leg pneumo | Unknown | FAM
D09 | Leg spp | Unknown | FAM |  |  |  | D09 | Leg pneumo | Unknown | FAM
D10 | Leg spp | Unknown | FAM |  |  |  | D10 | Leg pneumo | Unknown | FAM
D11 | Leg spp | Unknown | FAM |  |  |  | D11 | Leg pneumo | Unknown | FAM
D12 | Leg spp | Unknown | FAM |  |  |  | D12 | Leg pneumo | Unknown | FAM
E01 | Leg spp | Unknown | FAM |  |  |  | E01 | Leg pneumo | Unknown | FAM
E02 | Leg spp | Unknown | FAM |  |  |  | E02 | Leg pneumo | Unknown | FAM
E03 | Leg spp | Unknown | FAM |  |  |  | E03 | Leg pneumo | Unknown | FAM
E04 | Leg spp | Unknown | FAM |  |  |  | E04 | Leg pneumo | Unknown | FAM
E05 | Leg spp | Unknown | FAM |  |  |  | E05 | Leg pneumo | Unknown | FAM
E06 | Leg spp | Unknown | FAM |  |  |  | E06 | Leg pneumo | Unknown | FAM
E07 | Leg spp | Unknown | FAM |  |  |  | E07 | Leg pneumo | Unknown | FAM
E08 | Leg spp | Unknown | FAM |  |  |  | E08 | Leg pneumo | Unknown | FAM
E09 | Leg spp | Unknown | FAM |  |  |  | E09 | Leg pneumo | Unknown | FAM
E10 | Leg spp | Unknown | FAM |  |  |  | E10 | Leg pneumo | Unknown | FAM
E11 | Leg spp | Unknown | FAM |  |  |  | E11 | Leg pneumo | Unknown | FAM
E12 | Leg spp | Unknown | FAM |  |  |  | E12 | Leg pneumo | Unknown | FAM
F01 | Leg spp | Unknown | FAM |  |  |  | F01 | Leg pneumo | Unknown | FAM
F02 | Leg spp | Unknown | FAM |  |  |  | F02 | Leg pneumo | Unknown | FAM
F03 | Leg spp | Unknown | FAM |  |  |  | F03 | Leg pneumo | Unknown | FAM
F04 | Leg spp | Unknown | FAM |  |  |  | F04 | Leg pneumo | Unknown | FAM
F05 | Leg spp | Unknown | FAM |  |  |  | F05 | Leg pneumo | Unknown | FAM
F06 | Leg spp | Unknown | FAM |  |  |  | F06 | Leg pneumo | Unknown | FAM
F07 | Leg spp | Unknown | FAM |  |  |  | F07 | Leg pneumo | Unknown | FAM
F08 | Leg spp | Unknown | FAM |  |  |  | F08 | Leg pneumo | Unknown | FAM | 36.17
F09 | Leg spp | Unknown | FAM |  |  |  | F09 | Leg pneumo | Unknown | FAM
F10 | Leg spp | Unknown | FAM |  |  |  | F10 | Leg pneumo | Unknown | FAM
F11 | Leg spp | Unknown | FAM |  |  |  | F11 | Leg pneumo | Unknown | FAM
F12 | Leg spp | Unknown | FAM |  |  |  | F12 | Leg pneumo | Unknown | FAM
G01 | Leg spp | Unknown | FAM |  |  |  | G01 | Leg pneumo | Unknown | FAM
G02 | Leg spp | Unknown | FAM |  |  |  | G02 | Leg pneumo | Unknown | FAM
G03 | Leg spp | Unknown | FAM |  |  |  | G03 | Leg pneumo | Unknown | FAM
G04 | Leg spp | Unknown | FAM |  |  |  | G04 | Leg pneumo | Unknown | FAM
G05 | Leg spp | Unknown | FAM |  |  |  | G05 | Leg pneumo | Unknown | FAM
G06 | Leg spp | Unknown | FAM |  |  |  | G06 | Leg pneumo | Unknown | FAM
G07 | Leg spp | Unknown | FAM |  |  |  | G07 | Leg pneumo | Unknown | FAM
G08 | Leg spp | Unknown | FAM |  |  |  | G08 | Leg pneumo | Unknown | FAM
G09 | Leg spp | Unknown | FAM |  |  |  | G09 | Leg pneumo | Unknown | FAM
G10 | Leg spp | Unknown | FAM |  |  |  | G10 | Leg pneumo | Unknown | FAM
G11 | Leg spp | Unknown | FAM |  |  |  | G11 | Leg pneumo | Unknown | FAM
G12 | Leg spp | Unknown | FAM |  |  |  | G12 | Leg pneumo | Unknown | FAM
H01 | Leg spp | Unknown | FAM |  |  |  | H01 | Leg pneumo | Unknown | FAM
H02 | Leg spp | Unknown | FAM |  |  |  | H02 | Leg pneumo | Unknown | FAM
H03 | Leg spp | Unknown | FAM |  |  |  | H03 | Leg pneumo | Unknown | FAM
H04 | Leg spp | Unknown | FAM |  |  |  | H04 | Leg pneumo | Unknown | FAM
H05 | Leg spp | Unknown | FAM |  |  |  | H05 | Leg pneumo | Unknown | FAM
H06 | Leg spp | Unknown | FAM |  |  |  | H06 | Leg pneumo | Unknown | FAM
H07 | Leg spp | Unknown | FAM |  |  |  | H07 | Leg pneumo | Unknown | FAM
H08 | Leg spp | Unknown | FAM |  |  |  | H08 | Leg pneumo | Unknown | FAM
H09 | Leg spp | Unknown | FAM |  |  |  | H09 | Leg pneumo | Unknown | FAM
H10 | Leg spp | Unknown | FAM |  |  |  | H10 | Leg pneumo | Unknown | FAM
H11 | Leg spp | Unknown | FAM |  |  |  | H11 | Leg pneumo | Unknown | FAM
H12 | Leg spp | Unknown | FAM |  |  |  | H12 | Leg pneumo | Unknown | FAM
Note on AC6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This probe is empty. It has been re-ordered but didn't arrive in time.
Note on R26: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is expected concentration from the tube.
Note on T26: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: don't have details or tech sheet regarding this original tube.
Note on U26: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel and production use copy number per ul. Reply: assumes copy number is in 1000ul
Note on Q27: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Note, that this is different than micH std or other standards as I usually have 100x dilution from std_0 to std_1. However, this standard is simply 10x dilution.
Note on T27: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In 10ul of the source, there are this many molecules. Reply: conc uM = umol/L = 10e-6101e6*avogadro's number
Background: The internal amplification control (IAC) serves as a control in PCR. Ideally, the IAC should cross at a late Cq, greater than 30. In reactions containing Leg pneumophila F, R, P and pos control, the IAC pos control needs to be at a higher-than-expected concentration to generate fluorescence. Why?
We designed an experiment using two master mixes. The question we hoped to answer: is the IAC positive control high because there's a problem with the master mix or is it due to Leg pneumo in the master mix?
Raquel created two master mixes. Both mixes contained IAC F, R and P at the same optimized concentration (50, 50, 400). Both reactions included the same IAC positive control concentrations.
The "LEG" mastermix contained Leg spp F, R and P.
The "OPT" mastermix contained Leg pneumophila F, R and P.
IAC Cq is increased by ~20 cycles in reactions containing Leg pneumo F, R and positive controls. The IAC behaves as expected in reactions containing Leg spp F, R, P and positive controls. This data suggests that the IAC performance is not affected by reagents in the master mix, but is affected by the presence of Leg pneumophila F, R primer or positive control.
This is crazy! How can a synthetically generated construct like IAC interfere with Leg pneumophila primer/probe or positive control? It's hard to believe, but the data suggests that there is a difference when the IAC is mixed with Leg pneumo.
Okay, decision time.
Decision 1: Select a new IAC for Leg pneumo. Repeat optimizations. Push deadline back 2-3 weeks. Ordering = 1st week. Experiments = 2nd week.
Decision 2: Use IAC with Leg pneumo at a higher concentration. This has a few benefits: 1) More resistant to IAC contamination. 2) higher conc = greater confidence in Cq i.e. (at 3x109copies/ul, 95% of the samples will have a Cq between 37 and 33) There are a few drawbacks: 1) per rxn cost increase 2) unknown effects like crowding out issues with Leg pneumo samples?
I recommend Decision 2: use the IAC at a higher conc with Leg pneumo.

## interference
Using this master mix to test IAC amplification inhibited by presence of Leg pneumo F.
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 800
Stock Concentrations |  |  | Overage = |  |  | 5 | %
dNTP Mix | 25 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Dextran | 25 | Percent
Component | Desired |  | Per 20 uL |  | Multiple Reactions
dNTP Mix | 0.25 | mM | 0.32 | uL | 257.92 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 40.3 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 56.42 | uL
Trehalose | 10 | Percent | 4 | uL | 3224 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 161.2 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 1612 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 1791.1111 | uL
dextran | 3 | Percent | 2.4 | ul | 1934.4 | ul
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 11.2622 | uL | 9077.3511 | uL
mastermix
Component | Desired | Units | PROBE RXN | 115.2 | [6 empty] | Component | Desired | Units | PROBE RXN | 9.6
mastermix |  |  | 11.2622 | 1297.408 | [6 empty] | mastermix |  |  | 11.2622 | 108.1173
IAC F_primer@10uM |  | 0.25 | 0.5 | 57.6 | [6 empty] | IAC F_primer@10uM |  | 0.3 | 0.6 | 5.76
IAC R_primer@10uM |  | 0.25 | 0.5 | 57.6 | [6 empty] | IAC R_primer@10uM |  | 0.3 | 0.6 | 5.76
IAC probe@10uM |  | 0.5 | 1 | 115.2 | [6 empty] | IAC probe@10uM |  | 0.3 | 0.6 | 5.76
IAC pos control (std_5) |  |  |  | 0 | [6 empty] | sample |  |  | 2 | 19.2
water |  |  | 4.7378 | 545.792 | [6 empty] | water |  |  | 4.9378 | 47.4027
TOTAL |  |  | 6.7378 | 776.192 | [6 empty] | TOTAL |  |  | 20 | 192
procedure
Use program "qPCR test samples with std curve.py".
Allow program to create std curve using IAC stds 2..7.
This will use "18*3*1.14"= 61.56ul * 8 = 492.48ul from mastermix
to leftover mastermix add 10ul IAC "std_2" and homogenize about 10/1500 = 0.67% error in mmix, which is acceptable.
make 6 "sample" tubes containing increasing concentrations of F, R primer.
Program requires "2ul*12*1.1" = 26.4ul total sample
From 10uM F, R Leg pneumo primers, add to separate tube and water according to schedule below.
Leg pneumo Primers
| Component | Desired | Units | PROBE RXN | "Sample" Tube |  | ROW | F | R
| mastermix |  |  | 18 |  |  | row A | 50
| Leg pneumo primer | 0.05 |  | 0.1 | 2 |  | row B | 250
| Leg pneumo primer | 0.25 |  |  |  |  | row C | 800
| Leg pneumo primer | 0.8 |  |  |  |  | row D |  | 50
| water |  |  | 1.9 | 38 |  | row E |  | 250
| TOTAL |  |  | 20 | 40 |  | row F |  | 800
| Component | Desired | Units | PROBE RXN | "Sample" Tube
| mastermix |  |  | 18
| Leg pneumo primer | 0.05
| Leg pneumo primer | 0.25 |  | 0.5 | 10
| Leg pneumo primer | 0.8
| water |  |  | 1.5 | 30
| TOTAL |  |  | 20 | 40
| Component | Desired | Units | PROBE RXN | "Sample" Tube
| mastermix |  |  | 18
| Leg pneumo primer | 0.05
| Leg pneumo primer | 0.25
| Leg pneumo primer | 0.8 |  | 1.6 | 32
| water |  |  | 0.4 | 8
| TOTAL |  |  | 20 | 40
20210511
not much difference observed between concentrations. Inhibition more pronounced at lower concentrations?
mastermix
Component | Desired | Units | PROBE RXN | 115.2
mastermix |  |  | 11.2622 | 1297.408
IAC F_primer@10uM |  | 0.05 | 0.1 | 11.52
IAC R_primer@10uM |  | 0.05 | 0.1 | 11.52
IAC probe@10uM |  | 0.4 | 0.8 | 92.16
IAC pos control (std_2) |  |  |  | 0
water |  |  | 5.7378 | 660.992
TOTAL |  |  | 18 | 2073.6
Conclusion
20210512
Had Raquel make me a new mix: this one doesn't contain dextran.
I’m worried that dextran is attenuating my fluorescence signal.
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 800
Stock Concentrations |  |  | Overage = |  |  | 5 | %
dNTP Mix | 25 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL | [13 empty] | 20210512 Asked Raquel to send me down all the IAC tubes and she sent me these.
10X Taq Buffer (No Mg2+) | 10 | x | [13 empty] | Sequences on tube label are different. Wrote Laara Moran to ask for IDT_DNA spec sheets.
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent | [5 empty] | Component | Desired | Units | PROBE RXN | "Sample" Tube |  |  |  | Order # | Name | Date | Ordered By:
[8 empty] | mastermix |  |  | 18 |  |  |  |  | 217544881 | Leg_spp_UltraIAC | 43677 | D. Acevedo
[8 empty] | Leg pneumo primer | 0.05 |  | 0.1 | 2 |  |  |  | 217544882 | Leg_pne_UltraIAC | 43677 | D. Acevedo
Component | Desired |  | Per 20 uL |  | Multiple Reactions |  |  | Leg pneumo primer | 0.25 | [6 empty] | 263398090 | Leg_spp_UltraIAC V1 | 43833 | D. Acevedo
dNTP Mix | 0.25 | mM | 0.32 | uL | 257.92 | uL |  | Leg pneumo primer | 0.8 | [6 empty] | 263398091 | Leg_spp_UltraIAC V2 | 43833 | D. Acevedo
AptaTaq | 2.5 | Units | 0.05 | uL | 40.3 | uL |  | water |  |  | 1.9 | 38 |  |  |  | 263398092 | Leg_spp_UltraIAC V3 | 43833 | D. Acevedo
Mg2+ | 3.5 | mM | 0.07 | uL | 56.42 | uL |  | TOTAL |  |  | 20 | 40
Trehalose | 10 | Percent | 4 | uL | 3224 | uL | [9 empty] | 20210513
BSA | 0.5 | mg/mL | 0.2 | uL | 161.2 | uL |  | Component | Desired | Units | PROBE RXN | "Sample" Tube |  |  |  | ﻿Hi, @Neil Sharma﻿, in designing new IAC Forward and Reverse primers, I ﻿discover that the IAC_F﻿ shares 15bp identify with an upstream region of the IAC ultramer. It may explain some IAC weirdness.
PCR Buffer 10X | 1 | x | 2 | uL | 1612 | uL |  | mastermix |  |  | 18
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 1791.1111 | uL |  | Leg pneumo primer | 0.05
Water |  |  | 0 | uL | 0 | uL |  | Leg pneumo primer | 0.25 |  | 0.5 | 10
|  | Total: | 8.8622 | uL | 7142.9511 | uL |  | Leg pneumo primer | 0.8
[8 empty] | water |  |  | 1.5 | 30
mastermix | [7 empty] | TOTAL |  |  | 20 | 40
Component | Desired | Units | PROBE RXN | 115.2
mastermix |  |  | 8.8622 | 1020.928 |  |  |  | Component | Desired | Units | PROBE RXN | "Sample" Tube
IAC F_primer@10uM |  | 0.05 | 0.1 | 11.52 |  |  |  | mastermix |  |  | 18
IAC R_primer@10uM |  | 0.05 | 0.1 | 11.52 |  |  |  | Leg pneumo primer | 0.05
IAC probe@10uM |  | 0.4 | 0.8 | 92.16 |  |  |  | Leg pneumo primer | 0.25
IAC pos control (std_2) |  |  |  | 0 |  |  |  | Leg pneumo primer | 0.8 |  | 1.6 | 32
water |  |  | 8.1378 | 937.472 |  |  |  | water |  |  | 0.4 | 8
TOTAL |  |  | 18 | 2073.6 |  |  |  | TOTAL |  |  | 20 | 40
Made a video for Neil:
https://web.microsoftstream.com/video/ec6aaaf9-1292-4fee-b2a0-a8c2ff0e71a0
Making 3 New F and R primers
[16 empty] | name | seq (5'-3') | Length | Tm
[16 empty] | IAC_Rev_2 | CAGCACCCTAGCTTGGTAGAA | 21 | 59.5
[16 empty] | IAC Rev_3 | CAGCACCCTAGCTTGGTAG | 19 | 55.1
[16 empty] | IAC_Rev_4 | ATCGATCAGCTACGTGAGGTC | 21 | 56.2
[16 empty] | IAC Fwd_2 | ACCTAACGTACGACTCGTTC | 20 | 54.5
[16 empty] | IAC Fwd_3 | TTCGTAGTCGACCTAACCTT | 20 | 54.1
[16 empty] | IAC Fwd_4 | CTCGTTCGTAGTCGACCTAA | 20 | 55.7
Conclusion:
Leg pneumo forward primer interferes with IAC reaction.
Increasing concentration of Leg pneumo forward primer by 5x overcomes interaction.

## F vs F2
IAC std curve with old primers
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 800
Stock Concentrations |  |  | Overage = |  |  | 5 | %
dNTP Mix | 25 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Component | Desired |  | Per 20 uL |  | Multiple Reactions
dNTP Mix | 0.25 | mM | 0.32 | uL | 256.32 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 40.05 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 56.07 | uL
Trehalose | 10 | Percent | 4 | uL | 3204 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 160.2 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 1602 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 1780 | uL
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 8.8622 | uL | 7098.64 | uL
Component | Desired | Units | PROBE RXN | 74
mmix |  |  | 8.8622 | 655.8044
IAC_F | 0.25uM |  | 0.5 | 37
IAC_R | 0.25uM |  | 0.5 | 37
IAC_P | 0.5uM |  | 1 | 74
water |  |  | 7.1378 | 528.1956
TOTAL |  |  | 18 | 1332
IAC_Fwd and IAC_Rev_std | [14 empty] | IAC Ultramer
Well | Target | Property | Dye | Ct | Ct_mean | Ct_StDEV | Well | Target | Property | Dye | Ct | Ct_mean | Ct_StDEV |  | TubeName | Conc (uM) | pM | Copy Number | Copy Number/rxn | Log(copy#/rxn) | F, R | Ct_StDEV | 2*STDEV | F2, R | Ct_StDEV | 2*STDEV | [5 empty] | TubeName | Conc (uM) | pM | Copy Number | Copy Number/rxn | Log(copy#/rxn)
A01 | IAC | Unknown | CY5 |  | 7.1567 | 0.1159 | A01 | IAC | Unknown | CY5 |  | 8.1167 | 0.42063 |  | IDT Source | 100 | 100000000 | [14 empty] | IDT Source | 100 | 100000000
A02 | IAC | Unknown | CY5 | 7.28 |  |  | A02 | IAC | Unknown | CY5 | 8.55 |  |  |  | std_0 | 1 | 1000000 | 6.022e+14 | 1.2044e+12 | [12 empty] | std_0 | 1 | 1000000 | 6.022e+14 | 1.2044e+12
A03 | IAC | Unknown | CY5 | 7.14 |  |  | A03 | IAC | Unknown | CY5 | 8.09 |  |  |  | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808 | 7.1567 | 0.1159 | 0.2318 | 8.1167 | 0.42063 | 0.84127 | [5 empty] | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808
A04 | IAC | Unknown | CY5 | 7.05 |  |  | A04 | IAC | Unknown | CY5 | 7.71 |  |  |  | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | 10.3725 | 0.44984 | 0.89969 | 11.295 | 0.46943 | 0.93886 | [5 empty] | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808
A05 | IAC | Unknown | CY5 |  |  |  | A05 | IAC | Unknown | CY5 |  |  |  |  | std_3 | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 9.0808 | 14.1075 | 0.92478 | 1.8496 | 15.2775 | 0.94008 | 1.8802 | [5 empty] | std_3 | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 9.0808
A06 | IAC | Unknown | CY5 | 31.88 | 30.375 | 0.22398 | A06 | IAC | Unknown | CY5 | 36.06 | 35.4025 | 0.3348 |  | std_4 | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808 | 17.2725 | 0.75975 | 1.5195 | 18.12 | 0.56077 | 1.1215 | [5 empty] | std_4 | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808
A07 | IAC | Unknown | CY5 | 30.59 |  |  | A07 | IAC | Unknown | CY5 | 35.47 |  |  |  | std_5 | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808 | 20.3925 | 0.54126 | 1.0825 | 22.0325 | 0.75066 | 1.5013 | [5 empty] | std_5 | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808
A08 | IAC | Unknown | CY5 | 30.43 |  |  | A08 | IAC | Unknown | CY5 | 34.92 |  |  |  | std_6 | 1e-06 | 1 | 602200000 | 1204400 | 6.0808 | 24.05 | 0.86306 | 1.7261 | 24.8425 | 0.52532 | 1.0506 | [5 empty] | std_6 | 1e-06 | 1 | 602200000 | 1204400 | 6.0808
A09 | IAC | Unknown | CY5 | 30.42 |  |  | A09 | IAC | Unknown | CY5 | 35.53 |  |  |  | std_7 | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808 | 26.8375 | 0.40558 | 0.81115 | 28.375 | 0.46508 | 0.93016 | [5 empty] | std_7 | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808
A10 | IAC | Unknown | CY5 | 30.06 |  |  | A10 | IAC | Unknown | CY5 | 35.69 |  |  |  | std_8 | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808 | 29.525 | 0.36519 | 0.73039 | 31.745 | 0.66184 | 1.3237 | [5 empty] | std_8 | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808
A11 | IAC | Unknown | CY5 |  |  |  | A11 | IAC | Unknown | CY5 |  |  |  |  | std_9 | 1e-09 | 0.001 | 602200 | 1204.4 | 3.0808 | 30.375 | 0.22398 | 0.44796 | 35.4025 | 0.3348 | 0.6696 | [5 empty] | std_9 | 1e-09 | 0.001 | 602200 | 1204.4 | 3.0808
A12 | IAC | Unknown | CY5 |  |  |  | A12 | IAC | Unknown | CY5 |  |  |  |  | std_10 | 1e-10 | 0.0001 | 60220 | 120.44 | 2.0808 | 33.02 | 0.26357 | 0.52713 | 39.815 | 0.99702 | 1.994 | [5 empty] | std_10 | 1e-10 | 0.0001 | 60220 | 120.44 | 2.0808
B01 | IAC | Unknown | CY5 | 10.16 | 10.3725 | 0.44984 | B01 | IAC | Unknown | CY5 | 12.11 | 11.295 | 0.46943 |  | std_11 | 1e-11 | 1e-05 | 6022 | 12.044 | 1.0808 | 33.3075 | 0.3318 | 0.6636 | 33.28 | 0.31654 | 0.63309 | [5 empty] | std_11 | 1e-11 | 1e-05 | 6022 | 12.044 | 1.0808
B02 | IAC | Unknown | CY5 | 10.64 |  |  | B02 | IAC | Unknown | CY5 | 11.58 |  |  |  | std_12 | 1e-12 | 1e-06 | 602.2 | 1.2044 | 0.080771 | 31.2975 | 1.3124 | 2.6247 | 36.7675 | 0.84984 | 1.6997 | [5 empty] | std_12 | 1e-12 | 1e-06 | 602.2 | 1.2044 | 0.080771
B03 | IAC | Unknown | CY5 | 10.6 |  |  | B03 | IAC | Unknown | CY5 | 11.58 |  |  |  | std_13 | 1e-13 | 1e-07 | 60.22 | 0.12044 | -0.91923 | 30.8975 | 0.26069 | 0.52138 | 38.435 | 1.0051 | 2.0103 | [5 empty] | std_13 | 1e-13 | 1e-07 | 60.22 | 0.12044 | -0.91923
B04 | IAC | Unknown | CY5 | 10.55 |  |  | B04 | IAC | Unknown | CY5 | 11.42 |  |  |  | std_14 | 1e-14 | 1e-08 | 6.022 | 0.012044 | -1.9192 | 30.05 | 0.26583 | 0.53166 | 37.3667 | 1.0152 | 2.0304 | [5 empty] | std_14 | 1e-14 | 1e-08 | 6.022 | 0.012044 | -1.9192
B05 | IAC | Unknown | CY5 | 9.7 |  |  | B05 | IAC | Unknown | CY5 | 10.6 |  |  |  | std_15 | 1e-15 | 1e-09 | 0.6022 | 0.0012044 | -2.9192 | 37.5725 | 1.1898 | 2.3797 | 40.265 | 3.2032 | 6.4064 | [5 empty] | std_15 | 1e-15 | 1e-09 | 0.6022 | 0.0012044 | -2.9192
B06 | IAC | Unknown | CY5 | 33.46 | 33.02 | 0.26357 | B06 | IAC | Unknown | CY5 | 39.29 | 39.815 | 0.99702 | [15 empty] | mx | b | PCR Efficiency
B07 | IAC | Unknown | CY5 | 33.39 |  |  | B07 | IAC | Unknown | CY5 | [16 empty] | F, R | stds_2..8 | -3.2034 | 43.0478 | 105.2%
B08 | IAC | Unknown | CY5 | 32.93 |  |  | B08 | IAC | Unknown | CY5 | [16 empty] | F2, R | stds_2..8 | -3.3667 | 45.5084 | 98.17%
B09 | IAC | Unknown | CY5 | 32.99 |  |  | B09 | IAC | Unknown | CY5 | 39.11
B10 | IAC | Unknown | CY5 | 32.77 |  |  | B10 | IAC | Unknown | CY5 | 40.52
B11 | IAC | Unknown | CY5 |  |  |  | B11 | IAC | Unknown | CY5
B12 | IAC | Unknown | CY5 |  |  |  | B12 | IAC | Unknown | CY5
C01 | IAC | Unknown | CY5 | 13.81 | 14.1075 | 0.92478 | C01 | IAC | Unknown | CY5 | 15.29 | 15.2775 | 0.94008
C02 | IAC | Unknown | CY5 | 14.79 |  |  | C02 | IAC | Unknown | CY5 | 16
C03 | IAC | Unknown | CY5 | 14.71 |  |  | C03 | IAC | Unknown | CY5 | 15.86
C04 | IAC | Unknown | CY5 | 14.14 |  |  | C04 | IAC | Unknown | CY5 | 15.31
C05 | IAC | Unknown | CY5 | 12.79 |  |  | C05 | IAC | Unknown | CY5 | 13.94
C06 | IAC | Unknown | CY5 | 34.12 | 33.3075 | 0.3318 | C06 | IAC | Unknown | CY5 | 34.45 | 33.28 | 0.31654
C07 | IAC | Unknown | CY5 | 33.76 |  |  | C07 | IAC | Unknown | CY5 | 33.61
C08 | IAC | Unknown | CY5 | 33.3 |  |  | C08 | IAC | Unknown | CY5 | 32.98
C09 | IAC | Unknown | CY5 | 32.97 |  |  | C09 | IAC | Unknown | CY5 | 33.04
C10 | IAC | Unknown | CY5 | 33.2 |  |  | C10 | IAC | Unknown | CY5 | 33.49
C11 | IAC | Unknown | CY5 |  |  |  | C11 | IAC | Unknown | CY5
C12 | IAC | Unknown | CY5 |  |  |  | C12 | IAC | Unknown | CY5
D01 | IAC | Unknown | CY5 | 17.14 | 17.2725 | 0.75975 | D01 | IAC | Unknown | CY5 | 19.46 | 18.12 | 0.56077
D02 | IAC | Unknown | CY5 | 17.71 |  |  | D02 | IAC | Unknown | CY5 | 18.56
D03 | IAC | Unknown | CY5 | 17.53 |  |  | D03 | IAC | Unknown | CY5 | 18.15
D04 | IAC | Unknown | CY5 | 17.71 |  |  | D04 | IAC | Unknown | CY5 | 18.45
D05 | IAC | Unknown | CY5 | 16.14 |  |  | D05 | IAC | Unknown | CY5 | 17.32
D06 | IAC | Unknown | CY5 | 31.85 | 31.2975 | 1.3124 | D06 | IAC | Unknown | CY5 | 36.86 | 36.7675 | 0.84984
D07 | IAC | Unknown | CY5 | 33.23 |  |  | D07 | IAC | Unknown | CY5 | 36.68
D08 | IAC | Unknown | CY5 | 30.84 |  |  | D08 | IAC | Unknown | CY5 | 37.88
D09 | IAC | Unknown | CY5 | 30.3 |  |  | D09 | IAC | Unknown | CY5 | 36.7
D10 | IAC | Unknown | CY5 | 30.82 |  |  | D10 | IAC | Unknown | CY5 | 35.81
D11 | IAC | Unknown | CY5 |  |  |  | D11 | IAC | Unknown | CY5
D12 | IAC | Unknown | CY5 |  |  |  | D12 | IAC | Unknown | CY5
E01 | IAC | Unknown | CY5 | 19.95 | 20.3925 | 0.54126 | E01 | IAC | Unknown | CY5 | 22.05 | 22.0325 | 0.75066
E02 | IAC | Unknown | CY5 | 20.95 |  |  | E02 | IAC | Unknown | CY5 | 22.43
E03 | IAC | Unknown | CY5 | 20.57 |  |  | E03 | IAC | Unknown | CY5 | 22.53
E04 | IAC | Unknown | CY5 | 20.39 |  |  | E04 | IAC | Unknown | CY5 | 22.25
E05 | IAC | Unknown | CY5 | 19.66 |  |  | E05 | IAC | Unknown | CY5 | 20.92
E06 | IAC | Unknown | CY5 | 32.01 | 30.8975 | 0.26069 | E06 | IAC | Unknown | CY5 | 39.6 | 38.435 | 1.0051
E07 | IAC | Unknown | CY5 | 31.27 |  |  | E07 | IAC | Unknown | CY5 | 38.64
E08 | IAC | Unknown | CY5 | 30.75 |  |  | E08 | IAC | Unknown | CY5 | 37.41
E09 | IAC | Unknown | CY5 | 30.69 |  |  | E09 | IAC | Unknown | CY5 | 39.74
E10 | IAC | Unknown | CY5 | 30.88 |  |  | E10 | IAC | Unknown | CY5 | 37.95
E11 | IAC | Unknown | CY5 |  |  |  | E11 | IAC | Unknown | CY5
E12 | IAC | Unknown | CY5 |  |  |  | E12 | IAC | Unknown | CY5
F01 | IAC | Unknown | CY5 | 24.49 | 24.05 | 0.86306 | F01 | IAC | Unknown | CY5 | 25.67 | 24.8425 | 0.52532
F02 | IAC | Unknown | CY5 | 24.71 |  |  | F02 | IAC | Unknown | CY5 | 25.19
F03 | IAC | Unknown | CY5 | 24.36 |  |  | F03 | IAC | Unknown | CY5 | 25.05
F04 | IAC | Unknown | CY5 | 24.35 |  |  | F04 | IAC | Unknown | CY5 | 25.07
F05 | IAC | Unknown | CY5 | 22.78 |  |  | F05 | IAC | Unknown | CY5 | 24.06
F06 | IAC | Unknown | CY5 | 30.85 | 30.05 | 0.26583 | F06 | IAC | Unknown | CY5 | 38.49 | 37.3667 | 1.0152
F07 | IAC | Unknown | CY5 | 30.29 |  |  | F07 | IAC | Unknown | CY5 | 37.78
F08 | IAC | Unknown | CY5 | 29.97 |  |  | F08 | IAC | Unknown | CY5
F09 | IAC | Unknown | CY5 | 30.23 |  |  | F09 | IAC | Unknown | CY5 | 38.11
F10 | IAC | Unknown | CY5 | 29.71 |  |  | F10 | IAC | Unknown | CY5 | 36.21
F11 | IAC | Unknown | CY5 |  |  |  | F11 | IAC | Unknown | CY5
F12 | IAC | Unknown | CY5 |  |  |  | F12 | IAC | Unknown | CY5
G01 | IAC | Unknown | CY5 | 21.99 | 26.8375 | 0.40558 | G01 | IAC | Unknown | CY5 | 28.83 | 28.375 | 0.46508
G02 | IAC | Unknown | CY5 | 27.06 |  |  | G02 | IAC | Unknown | CY5 | 28.68
G03 | IAC | Unknown | CY5 | 27.05 |  |  | G03 | IAC | Unknown | CY5 | 28.65
G04 | IAC | Unknown | CY5 | 27.01 |  |  | G04 | IAC | Unknown | CY5 | 28.48
G05 | IAC | Unknown | CY5 | 26.23 |  |  | G05 | IAC | Unknown | CY5 | 27.69
G06 | IAC | Unknown | CY5 | 43.92 | 37.5725 | 1.1898 | G06 | IAC | Unknown | CY5 |  | 40.265 | 3.2032
G07 | IAC | Unknown | CY5 | 38.36 |  |  | G07 | IAC | Unknown | CY5 | 42.53
G08 | IAC | Unknown | CY5 | 36.55 |  |  | G08 | IAC | Unknown | CY5 | 38
G09 | IAC | Unknown | CY5 | 38.82 |  |  | G09 | IAC | Unknown | CY5
G10 | IAC | Unknown | CY5 | 36.56 |  |  | G10 | IAC | Unknown | CY5
G11 | IAC | Unknown | CY5 |  |  |  | G11 | IAC | Unknown | CY5
G12 | IAC | Unknown | CY5 |  |  |  | G12 | IAC | Unknown | CY5
H01 | IAC | Unknown | CY5 | 29.29 | 29.525 | 0.36519 | H01 | IAC | Unknown | CY5 | 31.45 | 31.745 | 0.66184
H02 | IAC | Unknown | CY5 | 29.66 |  |  | H02 | IAC | Unknown | CY5 | 32.25
H03 | IAC | Unknown | CY5 | 29.71 |  |  | H03 | IAC | Unknown | CY5 | 31.87
H04 | IAC | Unknown | CY5 | 29.75 |  |  | H04 | IAC | Unknown | CY5 | 32.08
H05 | IAC | Unknown | CY5 | 28.98 |  |  | H05 | IAC | Unknown | CY5 | 30.78
H06 | IAC | Unknown | CY5 |  | 38.05 | #DIV/0! | H06 | IAC | Unknown | CY5 |  | #DIV/0! | #DIV/0!
H07 | IAC | Unknown | CY5 |  |  |  | H07 | IAC | Unknown | CY5
H08 | IAC | Unknown | CY5 |  |  |  | H08 | IAC | Unknown | CY5
H09 | IAC | Unknown | CY5 |  |  |  | H09 | IAC | Unknown | CY5
H10 | IAC | Unknown | CY5 | 38.05 |  |  | H10 | IAC | Unknown | CY5
H11 | IAC | Unknown | CY5 |  |  |  | H11 | IAC | Unknown | CY5
H12 | IAC | Unknown | CY5 |  |  |  | H12 | IAC | Unknown | CY5
using EvaGreen and stds to make a melting curve to see if any diff. between primer sets.
Component | Desired | Units | PROBE RXN | 20 |  | Component | Desired | Units | PROBE RXN | 20
mmix |  |  | 8.8622 | 177.2444 |  | mmix |  |  | 8.8622 | 177.2444
IAC_F | 0.25uM |  | 0.5 | 10 |  | IAC_F2 | 0.25uM |  | 0.5 | 10
IAC_R | 0.25uM |  | 0.5 | 10 |  | IAC_R | 0.25uM |  | 0.5 | 10 | [11 empty] | conclusions
IAC_P | 0.5uM |  | 1 | 20 |  | IAC_P | 0.5uM |  | 1 | 20 | [11 empty] | Small difference in melting when comparing IAC amplification with F vs F2 primer consistent with longer amplicon being generated by F2.
EvaGreen |  |  | 1 | 20 |  | EvaGreen |  |  | 1 | 20
water |  |  | 6.1378 | 122.7556 |  | water |  |  | 6.1378 | 122.7556
TOTAL |  |  | 18 | 360 |  | TOTAL |  |  | 18 | 360
also want to test for IAC amplification in the presence of Leg pneumo F primer.
Make it with "sample qPCR with std curve" but without std curve.
Component | Desired | Units | PROBE RXN | 42
mmix |  |  | 8.8622 | 372.2133 |  | What is Leg pneumo conc of int tubes? |  |  |  |  | How much Leg pneumo F @ 10uM to get desired conc
IAC_F | 0.25uM |  | 0.5 | 21 |  | C1 | V1 | C2 | V2 |  | C1 | V1 | C2 | V2 | WATER
IAC_R | 0.25uM |  | 0.5 | 21 |  | 0.5 | 26.4 | 0.05 | 264 |  | 10 | 3.5 | 0.5 | 70 | 66.5
IAC_P | 0.5uM |  | 1 | 42 |  | 5 | 26.4 | 0.5 | 264 |  | 10 | 35 | 5 | 70 | 35
EvaGreen |  |  | 1 | 42 |  | 8 | 26.4 | 0.8 | 264 |  | 10 | 56 | 8 | 70 | 14
IAC std_5 |  |  | 0.2 | 8.4
water |  |  | 5.9378 | 249.3867 |  | Make 70 of each of these sample tubes. From each tube, 26.4ul will be drawn and added to the wells.
TOTAL |  |  | 18 | 756
Component | Desired | Units | PROBE RXN | 42 |  | IAC amplification with F and F2 with doping in Leg pneumo at 50, 800nM. Using program qPCR sample with no std curve. Also testing melting.
mmix |  |  | 8.8622 | 372.2133 |  | ROW | IAC__F | IAC_F2 | IAC_R | Leg pneumo F Conc (nM) | EvaG_Mean | EvaG_STD | IAC_Mean | IAC_STD | [6 empty] | Results
IAC_F2 | 0.25uM |  | 0.5 | 21 |  | row A | x |  | x | 50 | 24.0482 | 0.15211 | 26.0764 | 0.11457 | [7 empty] | Normal IAC amplification in row A
IAC_R | 0.25uM |  | 0.5 | 21 |  | row B | x |  | x | 800 | 17.5436 | 0.27793 | #DIV/0! | #DIV/0! | [7 empty] | No IAC amplification in row B. Leg pneumo F inhibits?
IAC_P | 0.5uM |  | 1 | 42 |  | row C | x |  | x | NTC | 24.3382 | 0.12828 | 25.9027 | 0.096963 | [7 empty] | Melt curve analysis suggests 800nM Leg pneumo F generates alternate amplicon with Tm ~49C.
EvaGreen |  |  | 1 | 42 |  | row D |  | x | x | 50 | 25.6336 | 0.16549 | 26.9045 | 0.17558 | [7 empty] | 50nM
IAC std_5 |  |  | 0.2 | 8.4 |  | row E |  | x | x | 800 | 25.6136 | 0.16567 | 26.8555 | 0.21379 | [7 empty] | This explains previous results showing that decreased IAC F from 250 to 50nM had better fluorescence.
water |  |  | 5.9378 | 249.3867 |  | row F |  | x | x | NTC | 25.5773 | 0.14934 | 26.8682 | 0.14428 | [7 empty] | Also explains why IAC ultramer needed to be at a high concentration.
TOTAL |  |  | 18 | 756 | [17 empty] | Also shows that F2 primer acceptable replacement.
Well | Target | Property | Dye | Ct | Ct_Mean | Ct_STDEV
A01 | EvaGreen | Unknown | FAM | 22.92 | 24.0482 | 0.15211
A02 | EvaGreen | Unknown | FAM | 23.94
A03 | EvaGreen | Unknown | FAM | 24
A04 | EvaGreen | Unknown | FAM | 24.1
A05 | EvaGreen | Unknown | FAM | 24.09
A06 | EvaGreen | Unknown | FAM | 23.95
A07 | EvaGreen | Unknown | FAM | 24.19
A08 | EvaGreen | Unknown | FAM | 23.93
A09 | EvaGreen | Unknown | FAM | 23.98
A10 | EvaGreen | Unknown | FAM | 24
A11 | EvaGreen | Unknown | FAM | 23.92
A12 | EvaGreen | Unknown | FAM | 24.43
B01 | EvaGreen | Unknown | FAM | 16.05 | 17.5436 | 0.27793
B02 | EvaGreen | Unknown | FAM | 17.99
B03 | EvaGreen | Unknown | FAM | 17.81
B04 | EvaGreen | Unknown | FAM | 17.69
B05 | EvaGreen | Unknown | FAM | 17.71
B06 | EvaGreen | Unknown | FAM | 17.59
B07 | EvaGreen | Unknown | FAM | 17.76
B08 | EvaGreen | Unknown | FAM | 17.47
B09 | EvaGreen | Unknown | FAM | 17.29
B10 | EvaGreen | Unknown | FAM | 17.29
B11 | EvaGreen | Unknown | FAM | 17.29
B12 | EvaGreen | Unknown | FAM | 17.09
C01 | EvaGreen | Unknown | FAM | 23.35 | 24.3382 | 0.12828
C02 | EvaGreen | Unknown | FAM | 24.34
C03 | EvaGreen | Unknown | FAM | 24.43
C04 | EvaGreen | Unknown | FAM | 24.4
C05 | EvaGreen | Unknown | FAM | 24.48
C06 | EvaGreen | Unknown | FAM | 24.17
C07 | EvaGreen | Unknown | FAM | 24.36
C08 | EvaGreen | Unknown | FAM | 24.23
C09 | EvaGreen | Unknown | FAM | 24.51
C10 | EvaGreen | Unknown | FAM | 24.35
C11 | EvaGreen | Unknown | FAM | 24.09
C12 | EvaGreen | Unknown | FAM | 24.36
D01 | EvaGreen | Unknown | FAM | 24.89 | 25.6336 | 0.16549
D02 | EvaGreen | Unknown | FAM | 25.83
D03 | EvaGreen | Unknown | FAM | 25.49
D04 | EvaGreen | Unknown | FAM | 25.81
D05 | EvaGreen | Unknown | FAM | 25.48
D06 | EvaGreen | Unknown | FAM | 25.43
D07 | EvaGreen | Unknown | FAM | 25.49
D08 | EvaGreen | Unknown | FAM | 25.58
D09 | EvaGreen | Unknown | FAM | 25.53
D10 | EvaGreen | Unknown | FAM | 25.66
D11 | EvaGreen | Unknown | FAM | 25.87
D12 | EvaGreen | Unknown | FAM | 25.8
E01 | EvaGreen | Unknown | FAM | 24.68 | 25.6136 | 0.16567
E02 | EvaGreen | Unknown | FAM | 25.56
E03 | EvaGreen | Unknown | FAM | 25.7
E04 | EvaGreen | Unknown | FAM | 25.66
E05 | EvaGreen | Unknown | FAM | 25.69
E06 | EvaGreen | Unknown | FAM | 25.49
E07 | EvaGreen | Unknown | FAM | 26.04
E08 | EvaGreen | Unknown | FAM | 25.59
E09 | EvaGreen | Unknown | FAM | 25.58
E10 | EvaGreen | Unknown | FAM | 25.43
E11 | EvaGreen | Unknown | FAM | 25.5
E12 | EvaGreen | Unknown | FAM | 25.51
F01 | EvaGreen | Unknown | FAM | 24.98 | 25.5773 | 0.14934
F02 | EvaGreen | Unknown | FAM | 25.76
F03 | EvaGreen | Unknown | FAM | 25.57
F04 | EvaGreen | Unknown | FAM | 25.72
F05 | EvaGreen | Unknown | FAM | 25.49
F06 | EvaGreen | Unknown | FAM | 25.39
F07 | EvaGreen | Unknown | FAM | 25.69
F08 | EvaGreen | Unknown | FAM | 25.53
F09 | EvaGreen | Unknown | FAM | 25.83
F10 | EvaGreen | Unknown | FAM | 25.46
F11 | EvaGreen | Unknown | FAM | 25.5
F12 | EvaGreen | Unknown | FAM | 25.41
G01 | EvaGreen | Unknown | FAM
G02 | EvaGreen | Unknown | FAM
G03 | EvaGreen | Unknown | FAM
G04 | EvaGreen | Unknown | FAM
G05 | EvaGreen | Unknown | FAM
G06 | EvaGreen | Unknown | FAM
G07 | EvaGreen | Unknown | FAM
G08 | EvaGreen | Unknown | FAM
G09 | EvaGreen | Unknown | FAM
G10 | EvaGreen | Unknown | FAM
G11 | EvaGreen | Unknown | FAM
G12 | EvaGreen | Unknown | FAM
H01 | EvaGreen | Unknown | FAM
H02 | EvaGreen | Unknown | FAM
H03 | EvaGreen | Unknown | FAM
H04 | EvaGreen | Unknown | FAM
H05 | EvaGreen | Unknown | FAM
H06 | EvaGreen | Unknown | FAM
H07 | EvaGreen | Unknown | FAM
H08 | EvaGreen | Unknown | FAM
H09 | EvaGreen | Unknown | FAM
H10 | EvaGreen | Unknown | FAM
H11 | EvaGreen | Unknown | FAM
H12 | EvaGreen | Unknown | FAM
A01 | IAC | Unknown | CY5 | 25.17 | 26.0764 | 0.11457
A02 | IAC | Unknown | CY5 | 25.98
A03 | IAC | Unknown | CY5 | 26.04
A04 | IAC | Unknown | CY5 | 26.11
A05 | IAC | Unknown | CY5 | 26.1
A06 | IAC | Unknown | CY5 | 26.03
A07 | IAC | Unknown | CY5 | 26.33
A08 | IAC | Unknown | CY5 | 25.97
A09 | IAC | Unknown | CY5 | 26.09
A10 | IAC | Unknown | CY5 | 26.13
A11 | IAC | Unknown | CY5 | 26.16
A12 | IAC | Unknown | CY5 | 25.9
B01 | IAC | Unknown | CY5 |  | #DIV/0! | #DIV/0!
B02 | IAC | Unknown | CY5
B03 | IAC | Unknown | CY5
B04 | IAC | Unknown | CY5
B05 | IAC | Unknown | CY5
B06 | IAC | Unknown | CY5
B07 | IAC | Unknown | CY5
B08 | IAC | Unknown | CY5
B09 | IAC | Unknown | CY5
B10 | IAC | Unknown | CY5
B11 | IAC | Unknown | CY5
B12 | IAC | Unknown | CY5
C01 | IAC | Unknown | CY5 | 24.97 | 25.9027 | 0.096963
C02 | IAC | Unknown | CY5 | 25.97
C03 | IAC | Unknown | CY5 | 26.02
C04 | IAC | Unknown | CY5 | 25.94
C05 | IAC | Unknown | CY5 | 25.95
C06 | IAC | Unknown | CY5 | 25.79
C07 | IAC | Unknown | CY5 | 25.95
C08 | IAC | Unknown | CY5 | 25.78
C09 | IAC | Unknown | CY5 | 25.96
C10 | IAC | Unknown | CY5 | 25.92
C11 | IAC | Unknown | CY5 | 25.71
C12 | IAC | Unknown | CY5 | 25.94
D01 | IAC | Unknown | CY5 | 26.3 | 26.9045 | 0.17558
D02 | IAC | Unknown | CY5 | 27.14
D03 | IAC | Unknown | CY5 | 26.73
D04 | IAC | Unknown | CY5 | 27.01
D05 | IAC | Unknown | CY5 | 26.69
D06 | IAC | Unknown | CY5 | 26.79
D07 | IAC | Unknown | CY5 | 26.8
D08 | IAC | Unknown | CY5 | 26.83
D09 | IAC | Unknown | CY5 | 26.79
D10 | IAC | Unknown | CY5 | 26.87
D11 | IAC | Unknown | CY5 | 27.17
D12 | IAC | Unknown | CY5 | 27.13
E01 | IAC | Unknown | CY5 | 25.97 | 26.8555 | 0.21379
E02 | IAC | Unknown | CY5 | 27.01
E03 | IAC | Unknown | CY5 | 27.04
E04 | IAC | Unknown | CY5 | 26.9
E05 | IAC | Unknown | CY5 | 26.9
E06 | IAC | Unknown | CY5 | 26.66
E07 | IAC | Unknown | CY5 | 27.3
E08 | IAC | Unknown | CY5 | 26.81
E09 | IAC | Unknown | CY5 | 26.85
E10 | IAC | Unknown | CY5 | 26.65
E11 | IAC | Unknown | CY5 | 26.53
E12 | IAC | Unknown | CY5 | 26.76
F01 | IAC | Unknown | CY5 | 26.27 | 26.8682 | 0.14428
F02 | IAC | Unknown | CY5 | 27.1
F03 | IAC | Unknown | CY5 | 26.89
F04 | IAC | Unknown | CY5 | 26.92
F05 | IAC | Unknown | CY5 | 26.75
F06 | IAC | Unknown | CY5 | 26.71
F07 | IAC | Unknown | CY5 | 27.07
F08 | IAC | Unknown | CY5 | 26.74
F09 | IAC | Unknown | CY5 | 27.03
F10 | IAC | Unknown | CY5 | 26.76
F11 | IAC | Unknown | CY5 | 26.84
F12 | IAC | Unknown | CY5 | 26.74
G01 | IAC | Unknown | CY5
G02 | IAC | Unknown | CY5
G03 | IAC | Unknown | CY5
G04 | IAC | Unknown | CY5
G05 | IAC | Unknown | CY5
G06 | IAC | Unknown | CY5
G07 | IAC | Unknown | CY5
G08 | IAC | Unknown | CY5
G09 | IAC | Unknown | CY5
G10 | IAC | Unknown | CY5
G11 | IAC | Unknown | CY5
G12 | IAC | Unknown | CY5
H01 | IAC | Unknown | CY5
H02 | IAC | Unknown | CY5
H03 | IAC | Unknown | CY5
H04 | IAC | Unknown | CY5
H05 | IAC | Unknown | CY5
H06 | IAC | Unknown | CY5
H07 | IAC | Unknown | CY5
H08 | IAC | Unknown | CY5
H09 | IAC | Unknown | CY5
H10 | IAC | Unknown | CY5
H11 | IAC | Unknown | CY5
H12 | IAC | Unknown | CY5
Note on P34: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Raquel provided me 100uM IAC ultramer from a tube upstairs. I diluted 10ul in 990ul water. This is std_0. Another 10ul from std_0 was used and put into 990ul std_1. Every dilution after this is 1/10, 100 into 900ul.
Note on Q36: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: confirmed that beginning concentration = 100uM.
Note on AJ36: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: confirmed that beginning concentration = 100uM.
Note on E154: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: From this tube, fill wells A1, B1 and C1 with 237.6ul.

## 50..800_F
Sum of F1 | Well
Sampling Count | A02 | A03 | A04 | A05 | A06 | A07 | A08 | A09 | A10 | A11 | B02 | B03 | B04 | B05 | B06 | B07 | B08 | B09 | B10 | B11 | B12 | C02 | C03 | C04 | C05 | C06 | C07 | C08 | C09 | C10 | C11 | C12 | Grand Total
1 | 224.8 | 225.31 | 183.89 | 206.62 | 210.79 | 141.45 | 245.55 | 159.49 | 160.71 | 152.55 | 113.99 | 163.56 | 159.64 | 136.74 | 171.61 | 168.2 | 222.93 | 198.37 | 173.25 | 153.5 | 61.07 | 234.77 | 234.68 | 148.42 | 219.54 | 300.91 | 269.09 | 307.91 | 235.35 | 250.62 | 287 | 170.47 | 6292.78
2 | 226.99 | 227.5 | 186.02 | 208.68 | 212.91 | 143.32 | 247.85 | 161.5 | 162.75 | 154.51 | 114.58 | 164.56 | 160.92 | 137.61 | 173.03 | 169.52 | 224.56 | 199.9 | 174.66 | 154.68 | 61.99 | 237.26 | 237.02 | 150.22 | 221.83 | 303.72 | 271.68 | 310.69 | 237.56 | 253.14 | 289.82 | 172.55 | 6353.53
3 | 229.81 | 230.33 | 188.88 | 211.34 | 215.67 | 145.89 | 250.79 | 164.24 | 165.51 | 157.18 | 115.24 | 165.7 | 162.54 | 138.64 | 174.82 | 171.18 | 226.55 | 201.8 | 176.45 | 156.16 | 63.35 | 240.52 | 240.05 | 152.65 | 224.82 | 307.29 | 275 | 314.19 | 240.39 | 256.4 | 293.45 | 175.36 | 6432.19
4 | 232.88 | 233.39 | 192 | 214.24 | 218.66 | 148.72 | 253.96 | 167.24 | 168.54 | 160.11 | 115.99 | 166.95 | 164.37 | 139.79 | 176.84 | 173.04 | 228.73 | 203.92 | 178.46 | 157.82 | 64.99 | 244.07 | 243.33 | 155.32 | 228.07 | 311.13 | 278.59 | 317.96 | 243.44 | 259.93 | 297.37 | 178.44 | 6518.29
5 | 236.09 | 236.6 | 195.3 | 217.28 | 221.81 | 151.72 | 257.28 | 170.42 | 171.75 | 163.22 | 116.84 | 168.3 | 166.37 | 141.04 | 179.05 | 175.07 | 231.08 | 206.23 | 180.67 | 159.65 | 66.89 | 247.81 | 246.77 | 158.15 | 231.48 | 315.15 | 282.34 | 321.9 | 246.63 | 263.63 | 301.49 | 181.7 | 6609.71
6 | 239.47 | 239.97 | 198.78 | 220.49 | 225.12 | 154.9 | 260.76 | 173.79 | 175.14 | 166.52 | 117.83 | 169.76 | 168.56 | 142.42 | 181.48 | 177.3 | 233.61 | 208.74 | 183.1 | 161.68 | 69.1 | 251.74 | 250.39 | 161.16 | 235.06 | 319.34 | 286.27 | 326 | 249.96 | 267.5 | 305.8 | 185.15 | 6706.89
7 | 243.02 | 243.51 | 202.45 | 223.85 | 228.6 | 158.28 | 264.41 | 177.35 | 178.73 | 170.01 | 118.96 | 171.35 | 170.98 | 143.94 | 184.15 | 179.74 | 236.35 | 211.49 | 185.78 | 163.91 | 71.65 | 255.87 | 254.17 | 164.34 | 238.82 | 323.71 | 290.38 | 330.29 | 253.45 | 271.55 | 310.32 | 188.81 | 6810.22
8 | 246.74 | 247.23 | 206.33 | 227.39 | 232.27 | 161.86 | 268.23 | 181.12 | 182.52 | 173.71 | 120.27 | 173.07 | 173.64 | 145.61 | 187.1 | 182.43 | 239.3 | 214.51 | 188.73 | 166.39 | 74.6 | 260.22 | 258.15 | 167.72 | 242.78 | 328.28 | 294.68 | 334.75 | 257.11 | 275.8 | 315.04 | 192.67 | 6920.25
9 | 250.66 | 251.14 | 210.42 | 231.12 | 236.13 | 165.67 | 272.23 | 185.11 | 186.52 | 177.62 | 121.8 | 174.94 | 176.57 | 147.46 | 190.34 | 185.39 | 242.49 | 217.81 | 192.01 | 169.15 | 77.99 | 264.79 | 262.32 | 171.3 | 246.93 | 333.05 | 299.18 | 339.42 | 260.92 | 280.25 | 320 | 196.76 | 7037.49
10 | 254.78 | 255.24 | 214.74 | 235.04 | 240.19 | 169.7 | 276.43 | 189.33 | 190.76 | 181.76 | 123.58 | 176.98 | 179.82 | 149.51 | 193.93 | 188.67 | 245.96 | 221.45 | 195.63 | 172.21 | 81.9 | 269.6 | 266.69 | 175.08 | 251.28 | 338.02 | 303.89 | 344.28 | 264.92 | 284.91 | 325.19 | 201.09 | 7162.56
11 | 259.11 | 259.55 | 219.29 | 239.16 | 244.46 | 173.97 | 280.83 | 193.8 | 195.23 | 186.15 | 125.68 | 179.22 | 183.41 | 151.79 | 197.9 | 192.29 | 249.73 | 225.45 | 199.66 | 175.63 | 86.4 | 274.66 | 271.28 | 179.1 | 255.86 | 343.21 | 308.82 | 349.36 | 269.1 | 289.79 | 330.63 | 205.66 | 7296.18
12 | 263.66 | 264.08 | 224.1 | 243.5 | 248.96 | 178.5 | 285.44 | 198.53 | 199.95 | 190.79 | 128.15 | 181.67 | 187.4 | 154.32 | 202.31 | 196.3 | 253.83 | 229.88 | 204.14 | 179.45 | 91.58 | 279.98 | 276.09 | 183.35 | 260.67 | 348.63 | 313.98 | 354.65 | 273.47 | 294.9 | 336.32 | 210.5 | 7439.08
13 | 268.45 | 268.83 | 229.16 | 248.07 | 253.69 | 183.31 | 290.28 | 203.53 | 204.94 | 195.71 | 131.09 | 184.39 | 191.83 | 157.17 | 207.22 | 200.76 | 258.3 | 234.77 | 209.13 | 183.74 | 97.56 | 285.57 | 281.13 | 187.85 | 265.71 | 354.28 | 319.37 | 360.18 | 278.03 | 300.24 | 342.28 | 215.61 | 7592.18
14 | 273.48 | 273.83 | 234.51 | 252.87 | 258.68 | 188.4 | 295.35 | 208.81 | 210.2 | 200.91 | 134.6 | 187.39 | 196.77 | 160.37 | 212.68 | 205.74 | 263.19 | 240.2 | 214.71 | 188.55 | 104.44 | 291.45 | 286.42 | 192.61 | 271 | 360.18 | 325.01 | 365.94 | 282.81 | 305.84 | 348.52 | 221.01 | 7756.47
15 | 278.77 | 279.08 | 240.14 | 257.93 | 263.92 | 193.79 | 300.67 | 214.4 | 215.76 | 206.41 | 138.81 | 190.75 | 202.29 | 164 | 218.78 | 211.28 | 268.54 | 246.23 | 220.96 | 193.98 | 112.37 | 297.63 | 291.96 | 197.66 | 276.55 | 366.32 | 330.9 | 371.94 | 287.8 | 311.68 | 355.05 | 226.72 | 7933.07
16 | 284.33 | 284.6 | 246.07 | 263.25 | 269.44 | 199.5 | 306.23 | 220.3 | 221.61 | 212.22 | 143.88 | 194.52 | 208.46 | 168.13 | 225.59 | 217.48 | 274.41 | 252.94 | 227.95 | 200.1 | 121.48 | 304.13 | 297.76 | 202.99 | 282.37 | 372.71 | 337.06 | 378.2 | 293 | 317.79 | 361.87 | 232.75 | 8123.12
17 | 290.17 | 290.38 | 252.32 | 268.85 | 275.24 | 205.55 | 312.07 | 226.53 | 227.78 | 218.37 | 150 | 198.79 | 215.38 | 172.87 | 233.23 | 224.42 | 280.86 | 260.41 | 235.79 | 207.02 | 131.96 | 310.95 | 303.84 | 208.63 | 288.46 | 379.37 | 343.49 | 384.71 | 298.44 | 324.18 | 369.01 | 239.11 | 8328.18
18 | 296.3 | 296.45 | 258.9 | 274.73 | 281.34 | 211.94 | 318.17 | 233.1 | 234.27 | 224.85 | 157.38 | 203.66 | 223.15 | 178.33 | 241.78 | 232.2 | 287.96 | 268.74 | 244.58 | 214.85 | 143.97 | 318.11 | 310.2 | 214.6 | 294.84 | 386.29 | 350.19 | 391.49 | 304.11 | 330.84 | 376.45 | 245.82 | 8549.59
19 | 302.74 | 302.82 | 265.81 | 280.9 | 287.75 | 218.69 | 324.56 | 240.02 | 241.1 | 231.7 | 166.3 | 209.29 | 231.87 | 184.68 | 251.38 | 240.93 | 295.8 | 278.03 | 254.43 | 223.71 | 157.7 | 325.61 | 316.84 | 220.9 | 301.51 | 393.48 | 357.19 | 398.53 | 310.01 | 337.78 | 384.22 | 252.89 | 8789.17
20 | 309.49 | 309.49 | 273.06 | 287.38 | 294.48 | 225.82 | 331.23 | 247.31 | 248.28 | 238.91 | 177.04 | 215.83 | 241.68 | 192.11 | 262.14 | 250.72 | 304.46 | 288.39 | 265.47 | 233.74 | 173.33 | 333.48 | 323.78 | 227.55 | 308.48 | 400.94 | 364.47 | 405.84 | 316.16 | 345.01 | 392.32 | 260.34 | 9048.73
21 | 316.56 | 316.46 | 280.68 | 294.18 | 301.53 | 233.34 | 338.2 | 254.97 | 255.81 | 246.5 | 189.9 | 223.53 | 252.71 | 200.83 | 274.21 | 261.7 | 314.06 | 299.94 | 277.81 | 245.06 | 191.03 | 341.71 | 331.02 | 234.58 | 315.76 | 408.67 | 372.04 | 413.43 | 322.55 | 352.53 | 400.74 | 268.17 | 9330.21
22 | 323.95 | 323.76 | 288.65 | 301.3 | 308.92 | 241.25 | 345.47 | 263.02 | 263.69 | 254.49 | 205.18 | 232.67 | 265.09 | 211.15 | 287.73 | 274.01 | 324.71 | 312.78 | 291.57 | 257.83 | 210.92 | 350.32 | 338.57 | 241.98 | 323.34 | 416.67 | 379.91 | 421.28 | 329.18 | 360.33 | 409.5 | 276.39 | 9635.61
23 | 331.68 | 331.37 | 296.98 | 308.74 | 316.65 | 249.57 | 353.04 | 271.45 | 271.93 | 262.88 | 223.17 | 243.61 | 278.99 | 223.38 | 302.83 | 287.78 | 336.55 | 327.04 | 306.87 | 272.17 | 233.1 | 359.3 | 346.43 | 249.79 | 331.23 | 424.92 | 388.07 | 429.39 | 336.05 | 368.43 | 418.59 | 285.02 | 9967
24 | 339.75 | 339.3 | 305.67 | 316.51 | 324.72 | 258.31 | 360.9 | 280.27 | 280.53 | 271.67 | 244.06 | 256.79 | 294.53 | 237.93 | 319.66 | 303.15 | 349.74 | 342.83 | 323.8 | 288.2 | 257.6 | 368.67 | 354.59 | 258 | 339.42 | 433.44 | 396.52 | 437.77 | 343.16 | 376.8 | 427.99 | 294.05 | 10326.33
25 | 348.14 | 347.55 | 314.73 | 324.62 | 333.13 | 267.46 | 369.06 | 289.48 | 289.48 | 280.87 | 267.98 | 272.74 | 311.87 | 255.23 | 338.37 | 320.26 | 364.45 | 360.25 | 342.45 | 306.04 | 284.38 | 378.42 | 363.05 | 266.64 | 347.92 | 442.19 | 405.25 | 446.39 | 350.5 | 385.45 | 437.71 | 303.5 | 10715.56
26 | 356.87 | 356.11 | 324.12 | 333.04 | 341.88 | 277.02 | 377.51 | 299.06 | 298.77 | 290.47 | 294.9 | 292.05 | 331.12 | 275.77 | 359.1 | 339.26 | 380.88 | 379.42 | 362.9 | 325.78 | 313.31 | 388.54 | 371.82 | 275.72 | 356.7 | 451.17 | 414.26 | 455.25 | 358.06 | 394.35 | 447.73 | 313.34 | 11136.28
27 | 365.91 | 364.97 | 333.85 | 341.78 | 350.96 | 286.98 | 386.22 | 309.01 | 308.39 | 300.46 | 324.66 | 315.38 | 352.4 | 300.06 | 381.99 | 360.27 | 399.22 | 400.42 | 385.22 | 347.52 | 344.22 | 399.03 | 380.87 | 285.24 | 365.76 | 460.35 | 423.52 | 464.32 | 365.83 | 403.5 | 458.03 | 323.58 | 11589.92
28 | 375.25 | 374.11 | 343.89 | 350.82 | 360.35 | 297.32 | 395.19 | 319.3 | 318.3 | 310.84 | 356.99 | 343.4 | 375.79 | 328.6 | 407.18 | 383.45 | 419.73 | 423.38 | 409.47 | 371.34 | 376.87 | 409.86 | 390.2 | 295.21 | 375.08 | 469.71 | 433.01 | 473.59 | 373.78 | 412.86 | 468.58 | 334.2 | 12077.65
29 | 384.87 | 383.5 | 354.2 | 360.14 | 370.04 | 308.03 | 404.39 | 329.91 | 328.49 | 321.57 | 391.55 | 376.68 | 401.38 | 361.85 | 434.84 | 408.93 | 442.66 | 448.42 | 435.72 | 397.36 | 411.02 | 421.02 | 399.79 | 305.63 | 384.63 | 479.24 | 442.72 | 483.02 | 381.91 | 422.43 | 479.35 | 345.18 | 12600.47
30 | 394.74 | 393.12 | 364.76 | 369.7 | 379.99 | 319.08 | 413.78 | 340.79 | 338.9 | 332.63 | 428.05 | 415.65 | 429.24 | 400.11 | 465.17 | 436.9 | 468.33 | 475.73 | 464.1 | 425.73 | 446.48 | 432.48 | 409.62 | 316.51 | 394.39 | 488.88 | 452.62 | 492.59 | 390.18 | 432.15 | 490.32 | 356.49 | 13159.21
31 | 404.83 | 402.93 | 375.51 | 379.49 | 390.17 | 330.42 | 423.33 | 351.91 | 349.5 | 343.99 | 466.41 | 460.42 | 459.49 | 443.5 | 498.45 | 467.58 | 497.09 | 505.55 | 494.76 | 456.71 | 483.17 | 444.2 | 419.66 | 327.85 | 404.33 | 498.61 | 462.66 | 502.26 | 398.56 | 442 | 501.43 | 368.1 | 13754.87
32 | 415.09 | 412.89 | 386.39 | 389.45 | 400.54 | 342 | 433 | 363.22 | 360.23 | 355.59 | 506.81 | 510.73 | 492.37 | 491.89 | 535.05 | 501.29 | 529.42 | 538.26 | 527.98 | 490.66 | 521.23 | 456.14 | 429.88 | 339.63 | 414.4 | 508.39 | 472.83 | 511.99 | 407.02 | 451.94 | 512.64 | 379.95 | 14388.9
33 | 425.46 | 422.94 | 397.35 | 399.53 | 411.04 | 353.78 | 442.73 | 374.64 | 371.03 | 367.39 | 549.84 | 565.95 | 528.35 | 544.9 | 575.52 | 538.48 | 565.89 | 574.37 | 564.2 | 528.11 | 561.05 | 468.26 | 440.25 | 351.86 | 424.57 | 518.18 | 483.08 | 521.74 | 415.53 | 461.93 | 523.9 | 392.01 | 15063.86
34 | 435.9 | 433.03 | 408.32 | 409.68 | 421.63 | 365.67 | 452.47 | 386.11 | 381.84 | 379.31 | 596.42 | 625.32 | 568.26 | 602.07 | 620.6 | 579.83 | 607.27 | 614.61 | 604.04 | 569.78 | 603.38 | 480.49 | 450.74 | 364.52 | 434.78 | 527.93 | 493.36 | 531.46 | 424.04 | 471.91 | 535.16 | 404.2 | 15784.13
35 | 446.34 | 443.09 | 419.22 | 419.83 | 432.24 | 377.61 | 462.16 | 397.57 | 392.57 | 391.3 | 647.75 | 688.37 | 613.62 | 663.14 | 671.33 | 626.21 | 654.59 | 659.99 | 648.4 | 616.63 | 649.38 | 492.77 | 461.3 | 377.6 | 445 | 537.62 | 503.64 | 541.12 | 432.52 | 481.85 | 546.35 | 416.46 | 16557.57
36 | 456.7 | 453.05 | 429.96 | 429.92 | 442.8 | 389.53 | 471.73 | 408.92 | 403.17 | 403.28 | 705.24 | 755.48 | 666.87 | 728.46 | 729.09 | 678.82 | 709.16 | 711.8 | 698.53 | 669.93 | 700.64 | 505.05 | 471.91 | 391.07 | 455.18 | 547.19 | 513.87 | 550.67 | 440.94 | 491.7 | 557.44 | 428.72 | 17396.82
37 | 466.92 | 462.85 | 440.47 | 439.88 | 453.25 | 401.33 | 481.12 | 420.1 | 413.55 | 415.17 | 770.48 | 828.37 | 731.65 | 799.29 | 795.6 | 739.13 | 772.65 | 771.71 | 756.04 | 731.23 | 759.19 | 517.24 | 482.52 | 404.93 | 465.27 | 556.62 | 524.02 | 560.07 | 449.24 | 501.41 | 568.35 | 440.91 | 18320.56
38 | 476.93 | 472.41 | 450.68 | 449.64 | 463.51 | 412.94 | 490.26 | 431.04 | 423.66 | 426.9 | 845.21 | 909.86 | 812.77 | 877.81 | 872.85 | 808.86 | 846.98 | 841.73 | 822.99 | 802.44 | 827.48 | 529.29 | 493.1 | 419.14 | 475.22 | 565.89 | 534.04 | 569.29 | 457.4 | 510.96 | 579.06 | 452.96 | 19353.3
39 | 486.66 | 481.68 | 460.51 | 459.13 | 473.53 | 424.29 | 499.11 | 441.67 | 433.44 | 438.39 | 931.23 | 1003.14 | 915.3 | 966.6 | 962.92 | 889.77 | 934.17 | 924.15 | 901.72 | 885.75 | 908.33 | 541.13 | 503.63 | 433.7 | 485.01 | 574.99 | 543.91 | 578.31 | 465.39 | 520.31 | 589.52 | 464.81 | 20522.2
40 | 496.05 | 490.61 | 469.91 | 468.3 | 483.25 | 435.31 | 507.61 | 451.94 | 442.85 | 449.6 | 1030 | 1110.46 | 1042.48 | 1068.01 | 1067.58 | 983.39 | 1035.95 | 1021.33 | 994.67 | 983.48 | 1004.76 | 552.69 | 514.08 | 448.58 | 494.61 | 583.9 | 553.63 | 587.12 | 473.19 | 529.45 | 599.71 | 476.41 | 21850.91
41 | 505.06 | 499.15 | 478.85 | 477.09 | 492.63 | 445.95 | 515.74 | 461.85 | 451.88 | 460.48 | 1142.04 | 1232.08 | 1192.52 | 1183.49 | 1187.72 | 1090.51 | 1153.31 | 1135.27 | 1103.9 | 1097.78 | 1119.68 | 563.92 | 524.44 | 463.76 | 504.01 | 592.64 | 563.17 | 595.71 | 480.8 | 538.37 | 609.63 | 487.74 | 23351.17
42 | 513.66 | 507.29 | 487.31 | 485.5 | 501.67 | 456.2 | 523.48 | 471.38 | 460.54 | 471 | 1266.48 | 1365.84 | 1357.07 | 1313.16 | 1322.52 | 1210.74 | 1285.95 | 1267.1 | 1230.46 | 1229.95 | 1255.18 | 574.79 | 534.72 | 479.25 | 513.21 | 601.25 | 572.55 | 604.11 | 488.22 | 547.1 | 619.29 | 498.77 | 25015.74
43 | 521.85 | 515.04 | 495.31 | 493.53 | 510.37 | 466.05 | 530.84 | 480.58 | 468.87 | 481.19 | 1400.52 | 1507.56 | 1524.42 | 1455.39 | 1468.56 | 1342.11 | 1431.95 | 1416.11 | 1373.75 | 1379.45 | 1411.48 | 585.28 | 544.93 | 495.03 | 522.24 | 609.78 | 581.8 | 612.35 | 495.48 | 555.67 | 628.73 | 509.56 | 26815.78
44 | 529.64 | 522.43 | 502.92 | 501.19 | 518.77 | 475.54 | 537.86 | 489.52 | 476.95 | 491.07 | 1537.08 | 1650.51 | 1684.44 | 1604.95 | 1618.59 | 1479.84 | 1586.1 | 1577.94 | 1530.43 | 1542 | 1585.49 | 595.39 | 555.11 | 511.11 | 531.14 | 618.29 | 590.97 | 620.5 | 502.61 | 564.15 | 638.01 | 520.17 | 28690.71
45 | 537.1 | 529.56 | 510.23 | 508.57 | 526.95 | 484.74 | 544.61 | 498.32 | 484.89 | 500.73 | 1660.06 | 1781.09 | 1824.74 | 1746.1 | 1758.02 | 1611.21 | 1733.43 | 1739.22 | 1691.1 | 1704.63 | 1766.83 | 605.17 | 565.32 | 527.51 | 539.97 | 626.89 | 600.14 | 628.63 | 509.68 | 572.6 | 647.23 | 530.72 | 30495.99
46 | 544.31 | 536.51 | 517.39 | 515.76 | 535.05 | 493.77 | 551.19 | 507.12 | 492.83 | 510.29 | 1754.27 | 1884 | 1933.19 | 1858.06 | 1870.51 | 1719.59 | 1853.6 | 1877.53 | 1835.37 | 1844.86 | 1931.75 | 614.67 | 575.63 | 544.26 | 548.84 | 635.7 | 609.38 | 636.86 | 516.78 | 581.16 | 656.5 | 541.38 | 32028.11
47 | 551.42 | 543.46 | 524.57 | 522.93 | 543.21 | 502.77 | 557.74 | 516.12 | 500.95 | 519.9 | 1809.95 | 1946.4 | 1999.66 | 1929.59 | 1946.2 | 1794.62 | 1934.18 | 1978.8 | 1944.59 | 1947.06 | 2058.39 | 624 | 586.16 | 561.41 | 557.85 | 644.87 | 618.83 | 645.33 | 524.02 | 589.95 | 665.98 | 552.38 | 33143.29
48 | 558.59 | 550.59 | 532 | 530.24 | 551.65 | 511.94 | 564.44 | 525.53 | 509.46 | 529.76 | 1817.48 | 1956.17 | 2013.42 | 1951.83 | 1974.14 | 1825.86 | 1964.27 | 2032.94 | 2006.69 | 1999.63 | 2133.84 | 633.33 | 597.03 | 579.04 | 567.18 | 654.59 | 628.64 | 654.19 | 531.56 | 599.16 | 675.86 | 564.02 | 33725.07
49 | 566.03 | 558.14 | 539.96 | 537.95 | 560.62 | 521.5 | 571.48 | 535.59 | 518.58 | 540.11 | 1770.21 | 1905.79 | 1966.81 | 1917.7 | 1944.55 | 1803.4 | 1935.11 | 2029.91 | 2010.55 | 1992.29 | 2146.72 | 642.85 | 608.43 | 597.26 | 577 | 665.08 | 638.98 | 663.65 | 539.57 | 609 | 686.36 | 576.61 | 33677.79
50 | 573.99 | 566.38 | 548.74 | 546.31 | 570.39 | 531.68 | 579.1 | 546.56 | 528.55 | 551.19 | 1668.15 | 1795.85 | 1858.85 | 1825.74 | 1854.59 | 1723.15 | 1844.79 | 1964.66 | 1950.86 | 1921.09 | 2091.2 | 652.83 | 620.56 | 616.21 | 587.54 | 676.59 | 650.07 | 673.95 | 548.28 | 619.7 | 697.73 | 590.55 | 32975.83
51 | 582.75 | 575.6 | 558.65 | 555.61 | 581.26 | 542.75 | 587.57 | 558.7 | 539.6 | 563.28 | 1518.64 | 1635.06 | 1695.99 | 1681.99 | 1710.41 | 1590.06 | 1700.13 | 1840.52 | 1831.68 | 1791.42 | 1971.16 | 663.59 | 633.65 | 636.1 | 599.09 | 689.44 | 662.14 | 685.35 | 557.93 | 631.55 | 710.27 | 606.24 | 31688.18
52 | 592.63 | 586.1 | 570.05 | 566.17 | 593.54 | 554.97 | 597.17 | 572.25 | 551.97 | 576.65 | 1340.06 | 1443.6 | 1497.83 | 1500.62 | 1526.68 | 1418.9 | 1516.21 | 1667.42 | 1664.54 | 1615.81 | 1798.76 | 675.49 | 647.98 | 657.2 | 611.97 | 703.97 | 675.47 | 698.16 | 568.83 | 644.86 | 724.3 | 624.11 | 29984.27
53 | 603.93 | 598.21 | 583.25 | 578.3 | 607.56 | 568.63 | 608.19 | 587.43 | 565.88 | 591.59 | 1160.24 | 1251.24 | 1296.75 | 1308.72 | 1331.38 | 1237.85 | 1320.99 | 1468.46 | 1469.43 | 1418.53 | 1594.85 | 689 | 663.88 | 679.87 | 626.55 | 720.56 | 690.38 | 712.73 | 581.32 | 659.99 | 740.2 | 644.56 | 28160.45
54 | 617.01 | 612.25 | 598.61 | 592.33 | 623.62 | 583.99 | 620.95 | 604.47 | 581.54 | 608.37 | 993.61 | 1073.15 | 1109.82 | 1126.02 | 1145.48 | 1066.39 | 1135.09 | 1270.53 | 1273.26 | 1225 | 1387.27 | 704.6 | 681.71 | 704.55 | 643.28 | 739.66 | 707.2 | 729.46 | 595.79 | 677.33 | 758.38 | 667.98 | 26458.7
55 | 632.2 | 628.54 | 616.43 | 608.57 | 642 | 601.32 | 635.77 | 623.54 | 599.16 | 627.23 | 844.84 | 913.87 | 942.38 | 959.53 | 976.7 | 910.46 | 966.04 | 1086.17 | 1090.08 | 1046.01 | 1191.07 | 722.86 | 701.87 | 731.8 | 662.67 | 761.76 | 726.33 | 748.77 | 612.69 | 697.31 | 779.27 | 694.7 | 24981.94
56 | 649.84 | 647.39 | 637 | 627.32 | 662.99 | 620.86 | 652.99 | 644.85 | 618.93 | 648.43 | 715.52 | 774.72 | 796.21 | 812.27 | 827.98 | 772.19 | 816.93 | 920.99 | 925.21 | 886.02 | 1012.74 | 744.4 | 724.83 | 762.26 | 685.26 | 787.38 | 748.2 | 771.16 | 632.52 | 720.4 | 803.38 | 724.98 | 23776.15
57 | 670.31 | 669.1 | 660.62 | 648.87 | 686.83 | 642.91 | 672.97 | 668.58 | 641.09 | 672.22 | 605.44 | 655.49 | 671.19 | 685.18 | 699.87 | 652.22 | 688.52 | 777.27 | 780.93 | 746.78 | 855.58 | 769.89 | 751.09 | 796.66 | 711.69 | 817.09 | 773.32 | 797.17 | 655.83 | 747.1 | 831.25 | 759 | 22862.06
58 | 693.95 | 693.99 | 687.55 | 673.53 | 713.79 | 667.75 | 696.1 | 694.97 | 665.87 | 698.9 | 513.24 | 554.99 | 565.97 | 577.61 | 591.41 | 550.14 | 580.01 | 654.9 | 657.51 | 628.11 | 720.4 | 800.07 | 781.25 | 835.79 | 742.64 | 851.5 | 802.23 | 827.41 | 683.24 | 777.96 | 863.49 | 796.9 | 22243.17
59 | 721.17 | 722.41 | 718.07 | 701.61 | 744.18 | 695.71 | 722.8 | 724.33 | 693.59 | 728.78 | 436.89 | 471.39 | 478.51 | 487.9 | 500.78 | 464.71 | 489.6 | 552.3 | 553.75 | 528.54 | 606.22 | 835.68 | 815.96 | 880.41 | 778.82 | 891.24 | 835.58 | 862.54 | 715.45 | 813.57 | 900.75 | 838.73 | 21911.97
60 | 752.38 | 754.73 | 752.52 | 733.47 | 778.32 | 727.2 | 753.53 | 757.02 | 724.64 | 762.28 | 374.16 | 402.59 | 406.44 | 413.87 | 425.79 | 394.12 | 415.03 | 467.18 | 467.63 | 445.98 | 510.99 | 877.51 | 855.97 | 931.25 | 820.96 | 936.94 | 874.09 | 903.31 | 753.2 | 854.55 | 943.79 | 884.6 | 21856.04
61 | 788.09 | 791.43 | 791.28 | 769.53 | 816.66 | 762.7 | 788.84 | 793.56 | 759.49 | 799.89 | 322.9 | 346.4 | 347.45 | 353.24 | 364.19 | 336.32 | 353.92 | 397.1 | 396.78 | 378.07 | 432.24 | 926.35 | 902.16 | 988.89 | 869.81 | 989.22 | 918.59 | 950.53 | 797.35 | 901.57 | 993.46 | 934.67 | 22062.68
62 | 828.86 | 833.07 | 834.86 | 810.32 | 859.76 | 802.79 | 829.32 | 834.53 | 798.72 | 842.23 | 281.18 | 300.77 | 299.38 | 303.85 | 313.87 | 289.28 | 304.07 | 339.68 | 338.83 | 322.51 | 367.48 | 982.98 | 955.49 | 1053.81 | 926.05 | 1048.7 | 970.07 | 1005.12 | 848.77 | 955.38 | 1050.74 | 989.27 | 22521.74
63 | 875.38 | 880.38 | 883.9 | 856.46 | 908.31 | 848.17 | 875.69 | 880.66 | 843.05 | 890.07 | 247.3 | 263.86 | 260.33 | 263.76 | 272.9 | 251.16 | 263.49 | 292.79 | 291.57 | 277.18 | 314.38 | 1048.12 | 1017.07 | 1126.38 | 990.33 | 1115.98 | 1029.63 | 1068.08 | 908.45 | 1016.73 | 1116.73 | 1048.97 | 23227.26
64 | 928.44 | 934.24 | 939.24 | 908.73 | 963.16 | 899.69 | 928.76 | 932.79 | 893.28 | 944.29 | 219.84 | 234.06 | 228.68 | 231.29 | 239.58 | 220.3 | 230.51 | 254.59 | 253.07 | 240.23 | 270.87 | 1122.44 | 1088.13 | 1206.99 | 1063.23 | 1191.66 | 1098.58 | 1140.49 | 977.41 | 1086.47 | 1192.76 | 1114.72 | 24178.52
65 | 988.98 | 995.74 | 1001.88 | 968.03 | 1025.34 | 958.33 | 989.48 | 991.9 | 950.36 | 1005.91 | 197.59 | 210 | 203.07 | 205.03 | 212.47 | 195.32 | 203.71 | 223.5 | 221.69 | 210.06 | 235.18 | 1206.58 | 1169.98 | 1296.15 | 1145.25 | 1276.4 | 1178.39 | 1223.51 | 1056.67 | 1165.46 | 1280.36 | 1187.8 | 25380.12
66 | 1058.09 | 1066.15 | 1073.01 | 1035.44 | 1096 | 1025.21 | 1058.96 | 1059.1 | 1015.33 | 1076.11 | 179.55 | 190.56 | 182.33 | 183.81 | 190.36 | 175.06 | 181.92 | 198.21 | 196.09 | 185.36 | 205.85 | 1301.19 | 1263.95 | 1394.75 | 1236.93 | 1371.02 | 1270.67 | 1318.36 | 1147.3 | 1254.68 | 1381.29 | 1269.87 | 26842.51
67 | 1136.96 | 1146.84 | 1153.94 | 1112.14 | 1176.45 | 1101.53 | 1138.41 | 1135.65 | 1089.35 | 1156.12 | 164.86 | 174.78 | 165.52 | 166.62 | 172.23 | 158.52 | 164.16 | 177.59 | 175.13 | 165.06 | 181.66 | 1407.09 | 1371.31 | 1504.22 | 1338.89 | 1476.67 | 1377.11 | 1426.29 | 1250.27 | 1355.17 | 1497.55 | 1362.87 | 28580.96
68 | 1226.86 | 1239.18 | 1246.02 | 1199.43 | 1268.06 | 1188.5 | 1229.15 | 1222.95 | 1173.62 | 1247.27 | 152.82 | 161.86 | 151.81 | 152.64 | 157.27 | 144.93 | 149.62 | 160.72 | 157.9 | 148.27 | 161.6 | 1525.47 | 1493.21 | 1626.66 | 1452.12 | 1594.95 | 1499.32 | 1548.53 | 1366.52 | 1468.24 | 1631.35 | 1468.99 | 30615.84
69 | 1329 | 1344.36 | 1350.56 | 1298.6 | 1372.26 | 1287.25 | 1332.39 | 1322.4 | 1269.33 | 1350.77 | 142.81 | 151.16 | 140.51 | 141.17 | 144.78 | 133.61 | 137.61 | 146.82 | 143.63 | 134.28 | 144.86 | 1658.13 | 1630.68 | 1764.79 | 1578.13 | 1728.06 | 1638.76 | 1686.41 | 1496.93 | 1595.49 | 1785.04 | 1590.6 | 32971.18
70 | 1444.44 | 1463.23 | 1468.7 | 1410.77 | 1490.36 | 1398.76 | 1449.15 | 1435.24 | 1377.5 | 1467.59 | 134.35 | 142.13 | 131.07 | 131.6 | 134.23 | 124.05 | 127.57 | 135.25 | 131.7 | 122.51 | 130.78 | 1807.62 | 1784.9 | 1921.8 | 1719.11 | 1878.78 | 1796.69 | 1841.44 | 1642.45 | 1739.09 | 1961.08 | 1730.22 | 35674.16
71 | 1573.88 | 1596.23 | 1601.26 | 1536.65 | 1623.3 | 1523.83 | 1579.91 | 1562.2 | 1498.81 | 1598.23 | 127.03 | 134.35 | 123.01 | 123.48 | 125.15 | 115.82 | 119.03 | 125.47 | 121.6 | 112.5 | 118.82 | 1977.08 | 1957.5 | 2101.14 | 1877.83 | 2050.22 | 1974.32 | 2015.68 | 1804.31 | 1901.71 | 2161.97 | 1890.39 | 38752.71
72 | 1717.62 | 1743.33 | 1748.66 | 1676.22 | 1771.3 | 1662.88 | 1724.36 | 1703.11 | 1633.35 | 1742.46 | 120.56 | 127.47 | 115.97 | 116.41 | 117.2 | 108.61 | 111.63 | 117.07 | 112.91 | 103.87 | 108.55 | 2170.05 | 2150.84 | 2306.36 | 2057.34 | 2245.48 | 2173.01 | 2211.89 | 1984.25 | 2086.48 | 2390.3 | 2073.54 | 42233.08
73 | 1875.87 | 1904.22 | 1911.14 | 1828.67 | 1933.65 | 1815.71 | 1881.24 | 1856.76 | 1780.51 | 1899.5 | 114.7 | 121.25 | 109.68 | 110.11 | 110.13 | 102.16 | 105.08 | 109.71 | 105.32 | 96.34 | 99.65 | 2390.64 | 2368.33 | 2541.59 | 2260.89 | 2467.76 | 2394.76 | 2433.91 | 2184.89 | 2296.63 | 2649.24 | 2282.21 | 46142.25
74 | 2050.51 | 2079.77 | 2090.49 | 1993.84 | 2109.8 | 1982.33 | 2049.65 | 2021.7 | 1939.73 | 2069.23 | 109.3 | 115.5 | 103.94 | 104.37 | 103.73 | 96.3 | 99.18 | 103.14 | 98.59 | 89.68 | 91.85 | 2645.49 | 2616.14 | 2813.63 | 2493.31 | 2722.07 | 2644.06 | 2688.31 | 2410.62 | 2536.12 | 2944.89 | 2520.82 | 50538.09
75 | 2249.74 | 2276.95 | 2294.94 | 2176.84 | 2303.97 | 2167.22 | 2233.9 | 2200.01 | 2113.75 | 2256.19 | 104.22 | 110.11 | 98.6 | 99.04 | 97.86 | 90.9 | 93.78 | 97.19 | 92.52 | 83.74 | 84.94 | 2948.9 | 2908.64 | 3137.28 | 2765.38 | 3020.61 | 2933.21 | 2989.25 | 2670.82 | 2813.13 | 3292.2 | 2800.83 | 55606.66
76 | 2496.4 | 2517.9 | 2547.81 | 2397.23 | 2535.26 | 2388.65 | 2453.95 | 2405.56 | 2316.27 | 2477.6 | 99.42 | 104.99 | 93.59 | 94.03 | 92.43 | 85.88 | 88.76 | 91.72 | 87 | 78.37 | 78.76 | 3331.3 | 3276.9 | 3543.35 | 3101.41 | 3393.24 | 3291.47 | 3367.38 | 2987.49 | 3148.46 | 3724.33 | 3150.11 | 61847.02
77 | 2836.37 | 2846.21 | 2891.37 | 2699 | 2850.11 | 2687.84 | 2757.92 | 2676.58 | 2583.9 | 2772.18 | 94.83 | 100.09 | 88.83 | 89.26 | 87.36 | 81.18 | 84.05 | 86.64 | 81.93 | 73.49 | 73.19 | 3835.22 | 3760.32 | 4068.95 | 3540.75 | 3888.12 | 3761.28 | 3865.39 | 3403.05 | 3582.68 | 4285.21 | 3612.34 | 70145.64
78 | 3284.65 | 3270.48 | 3326.65 | 3101.26 | 3271.81 | 3081.04 | 3170.55 | 3056.17 | 2957.35 | 3169.88 | 90.45 | 95.39 | 84.3 | 84.72 | 82.62 | 76.76 | 79.62 | 81.91 | 77.23 | 69.02 | 68.14 | 4421.1 | 4311.86 | 4661.07 | 4071.82 | 4476.19 | 4306.2 | 4443.01 | 3923.79 | 4118.51 | 4923.68 | 4158.17 | 80395.4
79 | 3733.65 | 3679.51 | 3733.41 | 3507.29 | 3701.67 | 3474.99 | 3594.37 | 3485.7 | 3378.97 | 3589.06 | 86.25 | 90.88 | 79.99 | 80.38 | 78.16 | 72.59 | 75.44 | 77.47 | 72.87 | 64.93 | 63.5 | 4923.01 | 4763.44 | 5142.22 | 4551.65 | 4996.87 | 4768.05 | 4933.78 | 4428.54 | 4632.85 | 5449.16 | 4636.74 | 89947.39
80 | 4016.3 | 3908.14 | 3942.78 | 3761.12 | 3976.54 | 3720.55 | 3868.96 | 3816.38 | 3704.94 | 3866.06 | 82.24 | 86.57 | 75.89 | 76.24 | 73.98 | 68.67 | 71.48 | 73.3 | 68.81 | 61.15 | 59.23 | 5138.27 | 4919.21 | 5304.93 | 4785.42 | 5245.22 | 4954.04 | 5135.55 | 4734.33 | 4941.67 | 5636.04 | 4856.9 | 95030.91
81 | 3977.66 | 3812.2 | 3814.94 | 3717.1 | 3941.27 | 3677.04 | 3841.99 | 3889.56 | 3782.1 | 3846.64 | 78.42 | 82.45 | 71.99 | 72.3 | 70.06 | 64.99 | 67.74 | 69.39 | 65.02 | 57.67 | 55.26 | 4906.92 | 4636.72 | 4997.44 | 4614.21 | 5050.09 | 4713.43 | 4890.51 | 4671.93 | 4873.25 | 5318.6 | 4661.01 | 92389.9
82 | 3552.2 | 3345.36 | 3315.02 | 3313.81 | 3526.6 | 3282.9 | 3444.12 | 3603.96 | 3511.33 | 3458.22 | 74.79 | 78.52 | 68.31 | 68.57 | 66.39 | 61.53 | 64.21 | 65.72 | 61.49 | 54.46 | 51.54 | 4209.88 | 3918.55 | 4222.57 | 4003.6 | 4376.12 | 4030.74 | 4183.05 | 4170.77 | 4353.95 | 4497.99 | 4020.55 | 81056.82
83 | 2822.74 | 2608.34 | 2556.47 | 2629.1 | 2809.55 | 2611.04 | 2748.4 | 2985.2 | 2915.03 | 2767.15 | 71.35 | 74.81 | 64.83 | 65.04 | 62.95 | 58.3 | 60.9 | 62.29 | 58.21 | 51.5 | 48.04 | 3217.46 | 2947.83 | 3176.93 | 3091.36 | 3381.24 | 3071.37 | 3186.84 | 3324.28 | 3476.85 | 3384.25 | 3088.41 | 63478.06
84 | 2016.12 | 1832.17 | 1777.63 | 1875.13 | 2011.3 | 1867.13 | 1969.71 | 2205.42 | 2158.2 | 1975.86 | 68.11 | 71.3 | 61.58 | 61.72 | 59.76 | 55.28 | 57.79 | 59.09 | 55.16 | 48.77 | 44.73 | 2246.44 | 2032.98 | 2192.06 | 2158.34 | 2382.89 | 2139.61 | 2219.46 | 2387.31 | 2502.37 | 2333.17 | 2163.57 | 45090.16
85 | 1368.9 | 1232.3 | 1187.87 | 1271.78 | 1367.22 | 1268.51 | 1339.37 | 1513.28 | 1482.99 | 1325.53 | 65.07 | 68.01 | 58.54 | 58.61 | 56.79 | 52.47 | 54.89 | 56.11 | 52.33 | 46.27 | 41.6 | 1524.48 | 1370.62 | 1479.21 | 1449.85 | 1626.96 | 1450.27 | 1504.2 | 1629.77 | 1710.71 | 1572.33 | 1469.96 | 30756.8
86 | 931.62 | 835.12 | 801.98 | 864.81 | 930.81 | 863.27 | 911.84 | 1025.3 | 1005.59 | 889.06 | 62.23 | 64.94 | 55.71 | 55.72 | 54.03 | 49.88 | 52.19 | 53.35 | 49.72 | 43.97 | 38.62 | 1046.53 | 937.67 | 1013.4 | 983.65 | 1121.23 | 994.89 | 1031.99 | 1113.69 | 1169.33 | 1075.38 | 1008.69 | 21136.21
87 | 648.33 | 580.18 | 555.88 | 601.49 | 647.68 | 600.52 | 634.5 | 706.68 | 693.39 | 610.75 | 59.59 | 62.07 | 53.09 | 53.04 | 51.49 | 47.48 | 49.69 | 50.8 | 47.32 | 41.86 | 35.79 | 736.72 | 658.65 | 713.39 | 684.89 | 791.05 | 699.8 | 726.23 | 777.19 | 815.48 | 755.35 | 708.64 | 14899.01
88 | 464.35 | 415.34 | 397.4 | 430.67 | 463.59 | 429.88 | 454.26 | 500.38 | 490.95 | 432.81 | 57.14 | 59.41 | 50.68 | 50.56 | 49.16 | 45.27 | 47.39 | 48.44 | 45.11 | 39.93 | 33.11 | 534.25 | 476.77 | 517.98 | 491.84 | 574.15 | 506.84 | 526.45 | 557.63 | 584.32 | 546.9 | 511.95 | 10834.91
89 | 342.83 | 306.64 | 293.19 | 317.93 | 341.83 | 317.19 | 335.1 | 365.14 | 358.07 | 316.96 | 54.88 | 56.96 | 48.46 | 48.28 | 47.01 | 43.24 | 45.26 | 46.28 | 43.1 | 38.18 | 30.58 | 399.29 | 355.7 | 387.93 | 364.54 | 429.06 | 378.19 | 393.33 | 412.12 | 431.05 | 408.25 | 380.43 | 8137
90 | 260.71 | 233.26 | 222.98 | 241.83 | 259.45 | 241.05 | 254.53 | 274.56 | 269.03 | 239.74 | 52.8 | 54.7 | 46.43 | 46.2 | 45.05 | 41.39 | 43.31 | 44.3 | 41.25 | 36.59 | 28.18 | 307.12 | 273.1 | 299.19 | 278.52 | 329.75 | 290.39 | 302.52 | 313.55 | 327.24 | 313.73 | 290 | 6302.45
91 | 203.89 | 182.53 | 174.53 | 189.24 | 202.38 | 188.32 | 198.75 | 212.3 | 207.93 | 186.92 | 50.9 | 52.63 | 44.58 | 44.29 | 43.26 | 39.7 | 41.53 | 42.5 | 39.58 | 35.14 | 25.92 | 242.56 | 215.32 | 237.06 | 218.94 | 260.12 | 228.99 | 239.01 | 245.16 | 255.25 | 247.64 | 225.47 | 5022.34
92 | 163.67 | 146.66 | 140.3 | 152.07 | 161.91 | 150.9 | 159.28 | 168.22 | 164.88 | 149.75 | 49.15 | 50.74 | 42.9 | 42.55 | 41.63 | 38.17 | 39.91 | 40.85 | 38.06 | 33.84 | 23.81 | 196.19 | 173.89 | 192.45 | 176.65 | 210.12 | 184.98 | 193.48 | 196.56 | 204.11 | 200.25 | 177.34 | 4105.27
93 | 134.63 | 120.78 | 115.57 | 125.28 | 132.56 | 123.75 | 130.78 | 135.82 | 133.63 | 122.73 | 47.57 | 49.02 | 41.37 | 40.98 | 40.15 | 36.77 | 38.44 | 39.36 | 36.68 | 32.67 | 21.83 | 162.09 | 143.48 | 159.66 | 145.94 | 173.43 | 152.66 | 160.04 | 161.19 | 166.92 | 165.45 | 139.88 | 3431.11
94 | 113.24 | 101.74 | 97.3 | 105.61 | 110.82 | 103.62 | 109.81 | 110.86 | 110.07 | 102.29 | 46.12 | 47.44 | 39.99 | 39.56 | 38.81 | 35.51 | 37.1 | 38 | 35.44 | 31.61 | 19.98 | 136.45 | 120.64 | 135.01 | 123.13 | 145.93 | 128.38 | 134.93 | 134.88 | 139.26 | 139.3 | 109.83 | 2922.66
95 | 97.21 | 87.43 | 83.49 | 90.92 | 94.37 | 88.38 | 94.13 | 90.56 | 91.45 | 86.09 | 44.82 | 46.02 | 38.74 | 38.27 | 37.59 | 34.37 | 35.89 | 36.78 | 34.32 | 30.66 | 18.27 | 116.79 | 103.14 | 116.11 | 105.81 | 124.93 | 109.72 | 115.68 | 114.87 | 118.23 | 119.25 | 85.42 | 2529.71
96 | 84.94 | 76.46 | 72.79 | 79.75 | 81.67 | 76.61 | 82.18 | 73.24 | 75.99 | 72.63 | 43.63 | 44.73 | 37.61 | 37.11 | 36.5 | 33.35 | 34.8 | 35.67 | 33.32 | 29.81 | 16.68 | 101.44 | 89.46 | 101.35 | 92.39 | 108.6 | 95.12 | 100.66 | 99.36 | 101.92 | 103.58 | 65.66 | 2219.01
97 | 75.34 | 67.85 | 64.31 | 71.09 | 71.67 | 67.34 | 72.9 | 58.03 | 62.63 | 61 | 42.56 | 43.56 | 36.6 | 36.07 | 35.51 | 32.42 | 33.82 | 34.68 | 32.42 | 29.06 | 15.21 | 89.27 | 78.59 | 89.65 | 81.81 | 95.68 | 83.49 | 88.73 | 87.12 | 89.04 | 91.12 | 49.89 | 1968.46
98 | 67.67 | 60.94 | 57.45 | 64.23 | 63.65 | 59.9 | 65.55 | 44.58 | 50.83 | 50.72 | 41.6 | 42.51 | 35.69 | 35.13 | 34.62 | 31.59 | 32.94 | 33.78 | 31.61 | 28.39 | 13.86 | 79.49 | 69.85 | 80.25 | 73.33 | 85.33 | 74.11 | 79.15 | 77.32 | 78.7 | 81.07 | 37.56 | 1763.4
99 | 61.4 | 55.28 | 51.81 | 58.69 | 57.14 | 53.84 | 59.6 | 32.82 | 40.37 | 41.57 | 40.73 | 41.56 | 34.88 | 34.29 | 33.83 | 30.85 | 32.15 | 32.98 | 30.9 | 27.79 | 12.61 | 71.56 | 62.74 | 72.61 | 66.44 | 76.94 | 66.46 | 71.36 | 69.38 | 70.31 | 72.88 | 28.13 | 1593.9
100 | 56.21 | 50.57 | 47.12 | 54.12 | 51.79 | 48.86 | 54.69 | 22.78 | 31.18 | 33.51 | 39.96 | 40.71 | 34.15 | 33.55 | 33.12 | 30.18 | 31.44 | 32.26 | 30.26 | 27.26 | 11.47 | 65.07 | 56.94 | 66.36 | 60.81 | 70.08 | 60.19 | 64.99 | 62.92 | 63.44 | 66.17 | 21.08 | 1453.24
101 | 51.85 | 46.62 | 43.21 | 50.31 | 47.37 | 44.75 | 50.6 | 14.43 | 23.28 | 26.5 | 39.26 | 39.95 | 33.5 | 32.88 | 32.48 | 29.59 | 30.81 | 31.62 | 29.69 | 26.8 | 10.42 | 59.74 | 52.2 | 61.23 | 56.18 | 64.45 | 55.04 | 59.76 | 57.63 | 57.79 | 60.65 | 15.94 | 1336.53
102 | 48.17 | 43.28 | 39.94 | 47.11 | 43.72 | 41.34 | 47.14 | 7.69 | 16.62 | 20.53 | 38.64 | 39.27 | 32.93 | 32.29 | 31.91 | 29.06 | 30.25 | 31.05 | 29.2 | 26.39 | 9.47 | 55.36 | 48.32 | 57 | 52.36 | 59.83 | 50.8 | 55.45 | 53.3 | 53.15 | 56.12 | 12.29 | 1239.98
103 | 45.07 | 40.46 | 37.22 | 44.4 | 40.71 | 38.53 | 44.22 | 2.39 | 11.13 | 15.54 | 38.09 | 38.66 | 32.42 | 31.76 | 31.41 | 28.59 | 29.75 | 30.55 | 28.76 | 26.04 | 8.59 | 51.75 | 45.16 | 53.52 | 49.23 | 56.03 | 47.33 | 51.92 | 49.76 | 49.35 | 52.4 | 9.77 | 1160.51
104 | 42.47 | 38.09 | 34.97 | 42.12 | 38.23 | 36.22 | 41.75 | -1.66 | 6.71 | 11.45 | 37.61 | 38.12 | 31.96 | 31.3 | 30.96 | 28.17 | 29.31 | 30.1 | 28.37 | 25.73 | 7.79 | 48.8 | 42.59 | 50.66 | 46.66 | 52.94 | 44.51 | 49.03 | 46.88 | 46.26 | 49.37 | 8.09 | 1095.56
105 | 40.3 | 36.11 | 33.12 | 40.2 | 36.21 | 34.35 | 39.66 | -4.67 | 3.2 | 8.14 | 37.17 | 37.64 | 31.57 | 30.9 | 30.57 | 27.81 | 28.92 | 29.71 | 28.04 | 25.46 | 7.07 | 46.39 | 40.53 | 48.33 | 44.57 | 50.42 | 42.22 | 46.67 | 44.56 | 43.76 | 46.91 | 7.03 | 1042.87
106 | 38.49 | 34.47 | 31.62 | 38.6 | 34.57 | 32.85 | 37.92 | -6.85 | 0.49 | 5.52 | 36.8 | 37.22 | 31.22 | 30.54 | 30.22 | 27.49 | 28.59 | 29.36 | 27.75 | 25.24 | 6.41 | 44.44 | 38.88 | 46.43 | 42.89 | 48.4 | 40.38 | 44.77 | 42.7 | 41.75 | 44.95 | 6.43 | 1000.54
107 | 37.02 | 33.13 | 30.42 | 37.27 | 33.27 | 31.66 | 36.47 | -8.37 | -1.58 | 3.48 | 36.47 | 36.85 | 30.92 | 30.23 | 29.92 | 27.21 | 28.29 | 29.06 | 27.51 | 25.05 | 5.82 | 42.87 | 37.59 | 44.91 | 41.55 | 46.78 | 38.93 | 43.26 | 41.22 | 40.17 | 43.4 | 6.17 | 966.95
108 | 35.82 | 32.04 | 29.46 | 36.19 | 32.24 | 30.74 | 35.28 | -9.38 | -3.12 | 1.91 | 36.18 | 36.53 | 30.66 | 29.97 | 29.67 | 26.97 | 28.04 | 28.81 | 27.3 | 24.9 | 5.28 | 41.64 | 36.59 | 43.71 | 40.5 | 45.52 | 37.8 | 42.07 | 40.07 | 38.94 | 42.19 | 6.15 | 940.67
109 | 34.87 | 31.18 | 28.73 | 35.32 | 31.46 | 30.05 | 34.31 | -9.99 | -4.24 | 0.73 | 35.94 | 36.25 | 30.44 | 29.75 | 29.45 | 26.77 | 27.83 | 28.59 | 27.13 | 24.78 | 4.79 | 40.68 | 35.85 | 42.78 | 39.7 | 44.56 | 36.95 | 41.15 | 39.2 | 38.01 | 41.29 | 6.3 | 920.61
110 | 34.14 | 30.51 | 28.19 | 34.63 | 30.88 | 29.56 | 33.55 | -10.29 | -5.01 | -0.12 | 35.73 | 36.01 | 30.26 | 29.56 | 29.26 | 26.6 | 27.65 | 28.41 | 26.99 | 24.68 | 4.36 | 39.97 | 35.32 | 42.09 | 39.12 | 43.85 | 36.34 | 40.48 | 38.57 | 37.34 | 40.63 | 6.6 | 905.86
111 | 33.59 | 30.02 | 27.8 | 34.12 | 30.47 | 29.23 | 32.97 | -10.36 | -5.5 | -0.71 | 35.56 | 35.82 | 30.11 | 29.41 | 29.11 | 26.46 | 27.5 | 28.26 | 26.88 | 24.61 | 3.97 | 39.47 | 34.98 | 41.6 | 38.73 | 43.37 | 35.92 | 40.01 | 38.14 | 36.88 | 40.2 | 7 | 895.62
112 | 33.21 | 29.67 | 27.57 | 33.75 | 30.22 | 29.05 | 32.55 | -10.25 | -5.77 | -1.09 | 35.42 | 35.65 | 29.99 | 29.28 | 28.99 | 26.34 | 27.39 | 28.14 | 26.8 | 24.57 | 3.63 | 39.15 | 34.79 | 41.3 | 38.49 | 43.07 | 35.68 | 39.72 | 37.88 | 36.62 | 39.96 | 7.48 | 889.25
113 | 32.97 | 29.46 | 27.45 | 33.51 | 30.11 | 28.99 | 32.27 | -10.01 | -5.85 | -1.28 | 35.31 | 35.52 | 29.9 | 29.19 | 28.89 | 26.26 | 27.3 | 28.05 | 26.75 | 24.55 | 3.32 | 39 | 34.75 | 41.15 | 38.41 | 42.95 | 35.6 | 39.6 | 37.78 | 36.51 | 39.88 | 8.03 | 886.32
114 | 32.87 | 29.37 | 27.45 | 33.39 | 30.11 | 29.04 | 32.12 | -9.64 | -5.79 | -1.32 | 35.23 | 35.42 | 29.83 | 29.12 | 28.82 | 26.19 | 27.24 | 27.99 | 26.72 | 24.55 | 3.05 | 38.98 | 34.84 | 41.15 | 38.45 | 42.97 | 35.65 | 39.61 | 37.82 | 36.56 | 39.95 | 8.64 | 886.38
115 | 32.88 | 29.39 | 27.54 | 33.38 | 30.22 | 29.19 | 32.08 | -9.19 | -5.61 | -1.23 | 35.18 | 35.35 | 29.79 | 29.08 | 28.78 | 26.16 | 27.2 | 27.95 | 26.71 | 24.57 | 2.82 | 39.1 | 35.04 | 41.28 | 38.6 | 43.12 | 35.82 | 39.75 | 37.98 | 36.73 | 40.16 | 9.3 | 889.12
116 | 32.99 | 29.51 | 27.73 | 33.47 | 30.42 | 29.43 | 32.14 | -8.67 | -5.32 | -1.04 | 35.15 | 35.3 | 29.78 | 29.06 | 28.76 | 26.14 | 27.18 | 27.93 | 26.73 | 24.6 | 2.62 | 39.34 | 35.34 | 41.52 | 38.86 | 43.4 | 36.1 | 40 | 38.24 | 37.02 | 40.48 | 10 | 894.21
117 | 33.19 | 29.7 | 27.98 | 33.64 | 30.7 | 29.73 | 32.29 | -8.12 | -4.97 | -0.77 | 35.14 | 35.29 | 29.78 | 29.07 | 28.77 | 26.14 | 27.19 | 27.94 | 26.76 | 24.66 | 2.46 | 39.66 | 35.71 | 41.85 | 39.2 | 43.77 | 36.46 | 40.34 | 38.59 | 37.39 | 40.89 | 10.71 | 901.14
118 | 33.4 | 29.89 | 28.21 | 33.82 | 30.95 | 29.99 | 32.47 | -7.75 | -4.71 | -0.55 | 35.19 | 35.33 | 29.83 | 29.12 | 28.81 | 26.18 | 27.24 | 27.99 | 26.82 | 24.73 | 2.37 | 39.96 | 36.04 | 42.16 | 39.5 | 44.1 | 36.78 | 40.65 | 38.9 | 37.72 | 41.25 | 11.2 | 907.59
Grand Total | 79735.21 | 78313.34 | 76606.59 | 75740.67 | 79618.53 | 72671.17 | 79630.36 | 75679.96 | 73661.16 | 75216.22 | 41754.41 | 45327.12 | 45311.11 | 44042.23 | 46292.3 | 43068.09 | 46827.08 | 48002.2 | 46691.9 | 44812.57 | 45124.32 | 95883.48 | 91903.64 | 93776.75 | 89365.27 | 101456.24 | 95100.16 | 99350.39 | 88201.92 | 94465.75 | 104815.37 | 88172.19 | 2306617.7

## 50..800_F2
Sum of F1 | Well
Sampling Count | D02 | D03 | D04 | D05 | D06 | D07 | D08 | D09 | D10 | D11 | D12 | E02 | E03 | E04 | E05 | E06 | E07 | E08 | E09 | E10 | E11 | E12 | F02 | F03 | F04 | F05 | F06 | F07 | F08 | F09 | F10 | F11 | F12 | Grand Total
1 | 287.04 | 337.78 | 278.97 | 384.21 | 350.59 | 359 | 435.68 | 412.3 | 393.94 | 316.28 | 206.48 | 296.34 | 250.68 | 244.16 | 288.69 | 350.13 | 273.22 | 375.75 | 326.93 | 353.86 | 331.98 | 240.51 | 303.82 | 299.11 | 281.76 | 387.13 | 360.94 | 341.62 | 443.98 | 341.28 | 295.07 | 391.96 | 238.66 | 10779.85
2 | 289.47 | 340.45 | 281.16 | 387 | 353.09 | 361.56 | 438.64 | 415.24 | 396.89 | 318.81 | 208.8 | 298.89 | 252.92 | 246.32 | 290.94 | 352.89 | 275.33 | 378.62 | 329.52 | 356.64 | 334.89 | 242.71 | 306.42 | 301.45 | 284.05 | 390.08 | 363.81 | 344.26 | 447.12 | 343.63 | 298.26 | 395.35 | 241.77 | 10866.98
3 | 292.49 | 343.7 | 283.82 | 390.32 | 356.04 | 364.58 | 442.06 | 418.69 | 400.42 | 321.91 | 211.9 | 302.08 | 255.76 | 249.05 | 293.68 | 356.24 | 277.87 | 382.07 | 332.69 | 360.01 | 338.53 | 245.5 | 309.67 | 304.3 | 286.88 | 393.64 | 367.32 | 347.47 | 450.81 | 346.36 | 302.45 | 399.58 | 246 | 10973.89
4 | 295.74 | 347.17 | 286.67 | 393.83 | 359.16 | 367.77 | 445.65 | 422.33 | 404.16 | 325.23 | 215.28 | 305.51 | 258.84 | 252 | 296.6 | 359.81 | 280.58 | 385.75 | 336.06 | 363.6 | 342.46 | 248.53 | 313.17 | 307.35 | 289.9 | 397.42 | 371.07 | 350.88 | 454.7 | 349.24 | 307.02 | 404.13 | 250.67 | 11088.28
5 | 299.14 | 350.8 | 289.64 | 397.47 | 362.41 | 371.08 | 449.34 | 426.08 | 408.03 | 328.68 | 218.86 | 309.1 | 262.07 | 255.1 | 299.65 | 363.52 | 283.41 | 389.56 | 339.58 | 367.33 | 346.55 | 251.73 | 316.83 | 310.55 | 293.07 | 401.35 | 374.98 | 354.43 | 458.72 | 352.2 | 311.83 | 408.88 | 255.63 | 11207.6
6 | 302.69 | 354.6 | 292.73 | 401.23 | 365.79 | 374.5 | 453.14 | 429.96 | 412.04 | 332.28 | 222.65 | 312.86 | 265.46 | 258.36 | 302.84 | 367.37 | 286.35 | 393.51 | 343.24 | 371.22 | 350.84 | 255.11 | 320.66 | 313.89 | 296.38 | 405.42 | 379.06 | 358.12 | 462.87 | 355.25 | 316.89 | 413.85 | 260.88 | 11332.04
7 | 306.4 | 358.56 | 295.95 | 405.13 | 369.29 | 378.04 | 457.05 | 433.96 | 416.19 | 336.03 | 226.65 | 316.8 | 269.04 | 261.79 | 306.16 | 371.39 | 289.41 | 397.62 | 347.04 | 375.25 | 355.31 | 258.67 | 324.68 | 317.38 | 299.84 | 409.66 | 383.31 | 361.96 | 467.16 | 358.4 | 322.2 | 419.05 | 266.44 | 11461.81
8 | 310.29 | 362.7 | 299.31 | 409.18 | 372.95 | 381.72 | 461.08 | 438.09 | 420.49 | 339.94 | 230.88 | 320.93 | 272.8 | 265.41 | 309.64 | 375.56 | 292.6 | 401.88 | 351.01 | 379.46 | 359.98 | 262.44 | 328.88 | 321.04 | 303.46 | 414.07 | 387.75 | 365.95 | 471.6 | 361.64 | 327.8 | 424.48 | 272.35 | 11597.36
9 | 314.36 | 367.04 | 302.81 | 413.37 | 376.75 | 385.53 | 465.23 | 442.36 | 424.95 | 344.02 | 235.35 | 325.26 | 276.77 | 269.21 | 313.27 | 379.9 | 295.93 | 406.31 | 355.14 | 383.83 | 364.87 | 266.41 | 333.28 | 324.87 | 307.26 | 418.65 | 392.38 | 370.11 | 476.19 | 364.98 | 333.68 | 430.17 | 278.6 | 11738.84
10 | 318.61 | 371.58 | 306.47 | 417.71 | 380.71 | 389.49 | 469.5 | 446.77 | 429.58 | 348.28 | 240.08 | 329.8 | 280.94 | 273.22 | 317.06 | 384.42 | 299.4 | 410.92 | 359.45 | 388.38 | 369.99 | 270.61 | 337.9 | 328.88 | 311.23 | 423.42 | 397.22 | 374.44 | 480.95 | 368.43 | 339.87 | 436.13 | 285.24 | 11886.68
11 | 323.07 | 376.33 | 310.29 | 422.22 | 384.83 | 393.59 | 473.9 | 451.32 | 434.37 | 352.72 | 245.08 | 334.56 | 285.34 | 277.45 | 321.03 | 389.12 | 303.02 | 415.7 | 363.93 | 393.11 | 375.34 | 275.04 | 342.74 | 333.09 | 315.39 | 428.38 | 402.28 | 378.95 | 485.87 | 371.98 | 346.39 | 442.37 | 292.27 | 12041.07
12 | 327.75 | 381.31 | 314.28 | 426.89 | 389.14 | 397.85 | 478.44 | 456.03 | 439.34 | 357.36 | 250.37 | 339.56 | 289.98 | 281.9 | 325.19 | 394.01 | 306.8 | 420.68 | 368.61 | 398.05 | 380.93 | 279.73 | 347.82 | 337.51 | 319.76 | 433.55 | 407.57 | 383.65 | 490.96 | 375.65 | 353.24 | 448.91 | 299.72 | 12202.54
13 | 332.64 | 386.53 | 318.45 | 431.74 | 393.63 | 402.27 | 483.11 | 460.9 | 444.49 | 362.21 | 255.96 | 344.8 | 294.87 | 286.59 | 329.54 | 399.11 | 310.75 | 425.85 | 373.48 | 403.18 | 386.78 | 284.69 | 353.15 | 342.14 | 324.34 | 438.93 | 413.09 | 388.54 | 496.24 | 379.42 | 360.44 | 455.76 | 307.62 | 12371.24
14 | 337.77 | 392.01 | 322.8 | 436.77 | 398.32 | 406.86 | 487.93 | 465.93 | 449.84 | 367.27 | 261.88 | 350.31 | 300.03 | 291.54 | 334.09 | 404.42 | 314.87 | 431.23 | 378.56 | 408.53 | 392.91 | 289.94 | 358.75 | 347.01 | 329.14 | 444.53 | 418.87 | 393.64 | 501.71 | 383.31 | 368.01 | 462.93 | 315.98 | 12547.69
15 | 343.15 | 397.76 | 327.35 | 442 | 403.22 | 411.62 | 492.89 | 471.12 | 455.38 | 372.55 | 268.12 | 356.08 | 305.47 | 296.77 | 338.87 | 409.94 | 319.17 | 436.83 | 383.86 | 414.09 | 399.31 | 295.48 | 364.62 | 352.11 | 334.18 | 450.37 | 424.91 | 398.96 | 507.37 | 387.32 | 375.97 | 470.45 | 324.84 | 12732.13
16 | 348.79 | 403.79 | 332.1 | 447.41 | 408.34 | 416.57 | 498 | 476.49 | 461.12 | 378.07 | 274.72 | 362.15 | 311.22 | 302.28 | 343.86 | 415.7 | 323.67 | 442.65 | 389.37 | 419.88 | 406.01 | 301.35 | 370.78 | 357.48 | 339.46 | 456.45 | 431.23 | 404.5 | 513.25 | 391.45 | 384.33 | 478.33 | 334.22 | 12925.02
17 | 354.7 | 410.12 | 337.06 | 453.04 | 413.69 | 421.71 | 503.27 | 482.02 | 467.07 | 383.83 | 281.69 | 368.52 | 317.28 | 308.09 | 349.1 | 421.68 | 328.37 | 448.7 | 395.12 | 425.9 | 413.01 | 307.55 | 377.25 | 363.12 | 345.01 | 462.79 | 437.84 | 410.28 | 519.33 | 395.7 | 393.11 | 486.58 | 344.14 | 13126.67
18 | 360.89 | 416.77 | 342.25 | 458.86 | 419.29 | 427.05 | 508.68 | 487.74 | 473.23 | 389.85 | 289.04 | 375.2 | 323.68 | 314.23 | 354.59 | 427.91 | 333.28 | 454.99 | 401.11 | 432.17 | 420.32 | 314.1 | 384.04 | 369.05 | 350.84 | 469.39 | 444.75 | 416.3 | 525.63 | 400.06 | 402.32 | 495.22 | 354.63 | 13337.46
19 | 367.38 | 423.75 | 347.67 | 464.91 | 425.15 | 432.58 | 514.26 | 493.63 | 479.61 | 396.13 | 296.8 | 382.21 | 330.43 | 320.71 | 360.34 | 434.39 | 338.41 | 461.52 | 407.34 | 438.68 | 427.96 | 321.03 | 391.15 | 375.28 | 356.95 | 476.26 | 451.98 | 422.57 | 532.16 | 404.55 | 411.97 | 504.26 | 365.71 | 13557.73
20 | 374.17 | 431.09 | 353.32 | 471.17 | 431.28 | 438.33 | 519.98 | 499.7 | 486.2 | 402.68 | 304.97 | 389.56 | 337.55 | 327.54 | 366.36 | 441.13 | 343.76 | 468.3 | 413.82 | 445.44 | 435.93 | 328.35 | 398.62 | 381.83 | 363.37 | 483.42 | 459.53 | 429.09 | 538.92 | 409.15 | 422.08 | 513.72 | 377.4 | 13787.76
21 | 381.27 | 438.79 | 359.22 | 477.65 | 437.7 | 444.29 | 525.86 | 505.95 | 493.01 | 409.5 | 313.58 | 397.27 | 345.07 | 334.75 | 372.67 | 448.12 | 349.35 | 475.34 | 420.55 | 452.46 | 444.24 | 336.08 | 406.44 | 388.71 | 370.11 | 490.86 | 467.42 | 435.88 | 545.92 | 413.86 | 432.66 | 523.6 | 389.73 | 14027.91
22 | 388.7 | 446.87 | 365.38 | 484.35 | 444.41 | 450.47 | 531.89 | 512.37 | 500.05 | 416.6 | 322.62 | 405.34 | 352.99 | 342.34 | 379.27 | 455.39 | 355.18 | 482.63 | 427.55 | 459.73 | 452.89 | 344.22 | 414.63 | 395.95 | 377.17 | 498.6 | 475.65 | 442.94 | 553.15 | 418.69 | 443.7 | 533.93 | 402.71 | 14278.36
23 | 396.47 | 455.34 | 371.78 | 491.27 | 451.44 | 456.87 | 538.07 | 518.96 | 507.3 | 423.99 | 332.12 | 413.79 | 361.35 | 350.35 | 386.18 | 462.92 | 361.26 | 490.19 | 434.8 | 467.26 | 461.88 | 352.81 | 423.2 | 403.55 | 384.57 | 506.64 | 484.25 | 450.27 | 560.64 | 423.61 | 455.21 | 544.69 | 416.36 | 14539.39
24 | 404.57 | 464.24 | 378.45 | 498.42 | 458.8 | 463.5 | 544.39 | 525.71 | 514.76 | 431.66 | 342.08 | 422.63 | 370.15 | 358.78 | 393.41 | 470.72 | 367.59 | 498 | 442.32 | 475.05 | 471.22 | 361.84 | 432.16 | 411.53 | 392.33 | 514.99 | 493.2 | 457.87 | 568.37 | 428.62 | 467.19 | 555.9 | 430.67 | 14811.12
25 | 413.01 | 473.55 | 385.39 | 505.78 | 466.5 | 470.35 | 550.84 | 532.61 | 522.42 | 439.62 | 352.5 | 431.85 | 379.42 | 367.66 | 400.96 | 478.78 | 374.18 | 506.07 | 450.09 | 483.08 | 480.89 | 371.33 | 441.51 | 419.9 | 400.45 | 523.65 | 502.53 | 465.75 | 576.35 | 433.72 | 479.63 | 567.55 | 445.65 | 15093.57
26 | 421.79 | 483.3 | 392.58 | 513.35 | 474.54 | 477.43 | 557.41 | 539.66 | 530.28 | 447.85 | 363.37 | 441.47 | 389.16 | 376.98 | 408.84 | 487.12 | 381.02 | 514.39 | 458.11 | 491.34 | 490.89 | 381.27 | 451.25 | 428.69 | 408.93 | 532.61 | 512.22 | 473.89 | 584.58 | 438.88 | 492.51 | 579.63 | 461.3 | 15386.64
27 | 430.9 | 493.49 | 400.03 | 521.11 | 482.95 | 484.72 | 564.09 | 546.83 | 538.32 | 456.35 | 374.68 | 451.48 | 399.39 | 386.77 | 417.06 | 495.71 | 388.12 | 522.95 | 466.37 | 499.84 | 501.2 | 391.68 | 461.37 | 437.89 | 417.8 | 541.88 | 522.29 | 482.3 | 593.06 | 444.1 | 505.81 | 592.14 | 477.59 | 15690.27
28 | 440.35 | 504.13 | 407.74 | 529.06 | 491.73 | 492.24 | 570.86 | 554.1 | 546.51 | 465.11 | 386.43 | 461.88 | 410.11 | 397.04 | 425.62 | 504.55 | 395.46 | 531.74 | 474.86 | 508.54 | 511.81 | 402.54 | 471.88 | 447.51 | 427.04 | 551.44 | 532.72 | 490.96 | 601.77 | 449.34 | 519.5 | 605.04 | 494.5 | 16004.11
29 | 450.11 | 515.21 | 415.68 | 537.18 | 500.89 | 499.96 | 577.71 | 561.46 | 554.85 | 474.1 | 398.58 | 472.65 | 421.34 | 407.77 | 434.51 | 513.62 | 403.05 | 540.75 | 483.55 | 517.43 | 522.68 | 413.83 | 482.75 | 457.57 | 436.66 | 561.29 | 543.5 | 499.85 | 610.72 | 454.59 | 533.52 | 618.31 | 511.98 | 16327.65
30 | 460.16 | 526.73 | 423.85 | 545.43 | 510.42 | 507.89 | 584.6 | 568.87 | 563.29 | 483.3 | 411.09 | 483.78 | 433.07 | 418.98 | 443.73 | 522.91 | 410.87 | 549.94 | 492.43 | 526.49 | 533.77 | 425.54 | 493.97 | 468.06 | 446.64 | 571.4 | 554.61 | 508.96 | 619.89 | 459.82 | 547.84 | 631.9 | 529.97 | 16660.2
31 | 470.48 | 538.67 | 432.23 | 553.8 | 520.34 | 516 | 591.52 | 576.3 | 571.82 | 492.68 | 423.93 | 495.24 | 445.29 | 430.66 | 453.28 | 532.39 | 418.91 | 559.3 | 501.47 | 535.66 | 545.05 | 437.63 | 505.5 | 478.98 | 456.99 | 581.77 | 566.04 | 518.26 | 629.27 | 465 | 562.37 | 645.78 | 548.39 | 17001
32 | 481.03 | 551.01 | 440.78 | 562.24 | 530.62 | 524.29 | 598.44 | 583.72 | 580.38 | 502.2 | 437.02 | 507 | 457.98 | 442.78 | 463.13 | 542.03 | 427.13 | 568.79 | 510.63 | 544.93 | 556.46 | 450.05 | 517.3 | 490.32 | 467.67 | 592.35 | 577.75 | 527.71 | 638.85 | 470.11 | 577.05 | 659.86 | 567.15 | 17348.76
33 | 491.77 | 563.72 | 449.47 | 570.73 | 541.26 | 532.72 | 605.32 | 591.09 | 588.95 | 511.82 | 450.31 | 519.01 | 471.13 | 455.33 | 473.26 | 551.81 | 435.52 | 578.37 | 519.87 | 554.25 | 567.95 | 462.76 | 529.33 | 502.07 | 478.68 | 603.13 | 589.71 | 537.27 | 648.59 | 475.09 | 591.77 | 674.08 | 586.13 | 17702.27
34 | 502.65 | 576.75 | 458.28 | 579.21 | 552.24 | 541.29 | 612.14 | 598.37 | 597.48 | 521.48 | 463.7 | 531.22 | 484.7 | 468.27 | 483.65 | 561.69 | 444.03 | 588.02 | 529.14 | 563.56 | 579.43 | 475.68 | 541.53 | 514.2 | 489.96 | 614.06 | 601.86 | 546.9 | 658.49 | 479.92 | 606.43 | 688.34 | 605.2 | 18059.87
35 | 513.61 | 590.04 | 467.14 | 587.63 | 563.52 | 549.96 | 618.85 | 605.52 | 605.91 | 531.12 | 477.1 | 543.56 | 498.64 | 481.55 | 494.27 | 571.62 | 452.63 | 597.67 | 538.4 | 572.83 | 590.85 | 488.73 | 553.81 | 526.67 | 501.48 | 625.09 | 614.16 | 556.55 | 668.51 | 484.56 | 620.91 | 702.55 | 624.19 | 18419.63
36 | 524.58 | 603.54 | 476 | 595.93 | 575.08 | 558.71 | 625.43 | 612.47 | 614.2 | 540.67 | 490.4 | 555.97 | 512.89 | 495.13 | 505.07 | 581.56 | 461.27 | 607.29 | 547.57 | 581.98 | 602.11 | 501.82 | 566.11 | 539.46 | 513.19 | 636.19 | 626.54 | 566.14 | 678.62 | 488.97 | 635.09 | 716.6 | 642.95 | 18779.53
37 | 535.49 | 617.15 | 484.81 | 604.06 | 586.85 | 567.49 | 631.83 | 619.2 | 622.28 | 550.07 | 503.5 | 568.35 | 527.39 | 508.93 | 516 | 591.45 | 469.89 | 616.81 | 556.62 | 590.97 | 613.12 | 514.86 | 578.33 | 552.51 | 525.03 | 647.29 | 638.92 | 575.62 | 688.79 | 493.1 | 648.84 | 730.35 | 661.27 | 19137.17
38 | 546.24 | 630.79 | 493.49 | 611.95 | 598.78 | 576.29 | 638.03 | 625.65 | 630.11 | 559.25 | 516.28 | 580.61 | 542.06 | 522.88 | 527.01 | 601.25 | 478.43 | 626.19 | 565.45 | 599.74 | 623.81 | 527.74 | 590.36 | 565.77 | 536.93 | 658.33 | 651.24 | 584.91 | 698.98 | 496.93 | 662.01 | 743.7 | 678.97 | 19490.16
39 | 556.76 | 644.35 | 501.98 | 619.54 | 610.8 | 585.06 | 643.98 | 631.78 | 637.63 | 568.12 | 528.6 | 592.64 | 556.79 | 536.89 | 538.04 | 610.88 | 486.83 | 635.36 | 574.03 | 608.24 | 634.07 | 540.36 | 602.1 | 579.17 | 548.81 | 669.25 | 663.39 | 593.93 | 709.15 | 500.43 | 674.47 | 756.5 | 695.84 | 19835.77
40 | 566.96 | 657.71 | 510.2 | 626.74 | 622.82 | 593.77 | 649.66 | 637.54 | 644.79 | 576.61 | 540.37 | 604.32 | 571.49 | 550.86 | 549.02 | 620.29 | 495.01 | 644.28 | 582.27 | 616.43 | 643.82 | 552.61 | 613.43 | 592.63 | 560.58 | 679.98 | 675.29 | 602.59 | 719.25 | 503.57 | 686.1 | 768.62 | 711.69 | 20171.3
41 | 576.73 | 670.75 | 518.06 | 633.51 | 634.75 | 602.38 | 655.04 | 642.91 | 651.53 | 584.65 | 551.47 | 615.55 | 586.03 | 564.68 | 559.88 | 629.43 | 502.91 | 652.9 | 590.11 | 624.27 | 652.98 | 564.4 | 624.21 | 606.07 | 572.14 | 690.45 | 686.85 | 610.82 | 729.23 | 506.34 | 696.78 | 779.94 | 726.34 | 20494.09
42 | 585.99 | 683.34 | 525.51 | 639.77 | 646.49 | 610.86 | 660.1 | 647.86 | 657.83 | 592.17 | 561.83 | 626.18 | 600.31 | 578.25 | 570.55 | 638.22 | 510.44 | 661.18 | 597.52 | 631.74 | 661.49 | 575.68 | 634.35 | 619.41 | 583.39 | 700.59 | 697.97 | 618.53 | 739.03 | 508.75 | 706.41 | 790.34 | 739.65 | 20801.73
43 | 594.65 | 695.37 | 532.45 | 645.47 | 657.92 | 619.17 | 664.82 | 652.36 | 663.65 | 599.12 | 571.4 | 636.12 | 614.18 | 591.44 | 580.95 | 646.61 | 517.53 | 669.06 | 604.44 | 638.86 | 669.28 | 586.39 | 643.71 | 632.56 | 594.22 | 710.33 | 708.56 | 625.63 | 748.61 | 510.81 | 714.94 | 799.73 | 751.48 | 21091.82
44 | 602.66 | 706.71 | 538.84 | 650.57 | 668.96 | 627.29 | 669.21 | 656.42 | 668.97 | 605.45 | 580.13 | 645.23 | 627.53 | 604.17 | 591.02 | 654.56 | 524.14 | 676.55 | 610.85 | 645.64 | 676.31 | 596.52 | 652.18 | 645.44 | 604.54 | 719.61 | 718.55 | 632.06 | 757.9 | 512.56 | 722.34 | 808.02 | 761.77 | 21362.7
45 | 609.93 | 717.26 | 544.62 | 655.05 | 679.49 | 635.21 | 673.26 | 660.05 | 673.78 | 611.15 | 588.07 | 653.41 | 640.24 | 616.32 | 600.69 | 662.01 | 530.19 | 683.61 | 616.76 | 652.16 | 682.57 | 606.12 | 659.68 | 657.96 | 614.24 | 728.35 | 727.87 | 637.75 | 766.85 | 514.07 | 728.62 | 815.18 | 770.51 | 21613.03
46 | 616.45 | 726.96 | 549.76 | 658.9 | 689.45 | 642.92 | 676.99 | 663.26 | 678.09 | 616.19 | 595.28 | 660.58 | 652.2 | 627.83 | 609.94 | 668.94 | 535.66 | 690.27 | 622.2 | 658.49 | 688.05 | 615.27 | 666.14 | 670.06 | 623.23 | 736.51 | 736.47 | 642.67 | 775.4 | 515.43 | 733.84 | 821.18 | 777.74 | 21842.35
47 | 622.18 | 735.76 | 554.26 | 662.15 | 698.78 | 650.43 | 680.42 | 666.11 | 681.93 | 620.59 | 601.87 | 666.66 | 663.33 | 638.65 | 618.72 | 675.35 | 540.53 | 696.54 | 627.22 | 664.77 | 692.79 | 624.12 | 671.5 | 681.68 | 631.43 | 744.04 | 744.35 | 646.78 | 783.52 | 516.78 | 738.1 | 826.03 | 783.58 | 22050.95
48 | 627.16 | 743.67 | 558.16 | 664.87 | 707.48 | 657.78 | 683.6 | 668.65 | 685.34 | 624.39 | 608.02 | 671.61 | 673.57 | 648.76 | 627.04 | 681.23 | 544.8 | 702.49 | 631.95 | 671.18 | 696.83 | 632.86 | 675.74 | 692.8 | 638.78 | 750.91 | 751.52 | 650.08 | 791.18 | 518.27 | 741.57 | 829.8 | 788.22 | 22240.31
49 | 631.41 | 750.72 | 561.52 | 667.18 | 715.58 | 665.04 | 686.59 | 670.99 | 688.41 | 627.64 | 613.95 | 675.44 | 682.9 | 658.18 | 634.94 | 686.65 | 548.51 | 708.18 | 636.54 | 677.92 | 700.26 | 641.74 | 678.9 | 703.42 | 645.26 | 757.11 | 758.05 | 652.61 | 798.38 | 520.12 | 744.47 | 832.58 | 791.91 | 22413.1
50 | 635.03 | 757.03 | 564.46 | 669.25 | 723.18 | 672.31 | 689.45 | 673.22 | 691.22 | 630.41 | 619.93 | 678.19 | 691.32 | 666.99 | 642.49 | 691.67 | 551.73 | 713.73 | 641.2 | 685.26 | 703.18 | 651.05 | 681.04 | 713.57 | 650.86 | 762.64 | 764.06 | 654.43 | 805.15 | 522.56 | 747.06 | 834.49 | 794.98 | 22573.14
51 | 638.14 | 762.76 | 567.16 | 671.31 | 730.45 | 679.72 | 692.27 | 675.48 | 693.9 | 632.82 | 626.28 | 679.96 | 698.92 | 675.34 | 649.81 | 696.43 | 554.57 | 719.27 | 646.24 | 693.46 | 705.71 | 661.12 | 682.27 | 723.33 | 655.62 | 767.54 | 769.7 | 655.66 | 811.54 | 525.88 | 749.66 | 835.69 | 797.79 | 22725.8
52 | 640.91 | 768.12 | 569.85 | 673.66 | 737.64 | 687.48 | 695.18 | 677.95 | 696.59 | 634.96 | 633.32 | 680.91 | 705.79 | 683.41 | 657.07 | 701.1 | 557.17 | 724.96 | 651.98 | 702.83 | 708.02 | 672.26 | 682.78 | 732.82 | 659.64 | 771.86 | 775.23 | 656.44 | 817.68 | 530.39 | 752.61 | 836.38 | 800.76 | 22877.75
53 | 643.54 | 773.38 | 572.83 | 676.7 | 745.05 | 695.83 | 698.31 | 680.83 | 699.49 | 636.99 | 641.4 | 681.27 | 712.13 | 691.45 | 664.5 | 705.92 | 559.72 | 730.99 | 658.85 | 713.68 | 710.29 | 684.79 | 682.79 | 742.22 | 663.05 | 775.69 | 780.91 | 657 | 823.73 | 536.41 | 756.29 | 836.79 | 804.33 | 23037.15
54 | 646.3 | 778.86 | 576.45 | 680.89 | 753.08 | 705.08 | 701.81 | 684.34 | 702.79 | 639.04 | 650.84 | 681.37 | 718.15 | 699.8 | 672.36 | 711.17 | 562.46 | 737.57 | 667.32 | 726.3 | 712.71 | 698.97 | 682.59 | 751.73 | 666.02 | 779.17 | 787.12 | 657.57 | 829.92 | 544.29 | 761.1 | 837.14 | 808.92 | 23213.23
55 | 649.49 | 784.92 | 581.12 | 686.76 | 762.16 | 715.59 | 705.89 | 688.77 | 706.75 | 641.27 | 661.96 | 681.6 | 724.14 | 708.8 | 680.99 | 717.2 | 565.65 | 744.94 | 677.9 | 740.94 | 715.51 | 715.02 | 682.55 | 761.63 | 668.81 | 782.44 | 794.26 | 658.5 | 836.52 | 554.34 | 767.45 | 837.69 | 814.99 | 23416.55
56 | 653.46 | 791.98 | 587.33 | 694.91 | 772.79 | 727.78 | 710.76 | 694.43 | 711.64 | 643.86 | 675 | 682.47 | 730.42 | 718.89 | 690.73 | 724.41 | 569.62 | 753.32 | 691.12 | 757.79 | 718.95 | 733.03 | 683.1 | 772.22 | 671.7 | 785.71 | 802.78 | 660.14 | 843.87 | 566.88 | 775.72 | 838.73 | 822.93 | 23658.47
57 | 658.59 | 800.45 | 595.58 | 705.97 | 785.45 | 742.08 | 716.68 | 701.68 | 717.8 | 646.99 | 690.14 | 684.61 | 737.35 | 730.47 | 701.97 | 733.22 | 574.72 | 762.93 | 707.47 | 776.96 | 723.29 | 753.01 | 684.76 | 783.85 | 675.02 | 789.2 | 813.15 | 662.95 | 852.32 | 582.13 | 786.29 | 840.54 | 833.11 | 23950.73
58 | 665.32 | 810.79 | 606.43 | 720.54 | 800.66 | 758.91 | 723.97 | 710.93 | 725.59 | 650.88 | 707.5 | 688.75 | 745.32 | 744 | 715.08 | 744.1 | 581.37 | 774 | 727.38 | 798.45 | 728.87 | 774.84 | 688.11 | 796.86 | 679.15 | 793.18 | 825.88 | 667.46 | 862.3 | 600.27 | 799.48 | 843.45 | 845.85 | 24305.67
59 | 674.1 | 823.45 | 620.41 | 739.2 | 818.88 | 778.67 | 732.97 | 722.62 | 735.43 | 655.74 | 727.09 | 695.76 | 754.73 | 759.88 | 730.43 | 757.51 | 590.02 | 786.7 | 751.14 | 822.17 | 736.03 | 798.31 | 693.82 | 811.61 | 684.51 | 797.95 | 841.42 | 674.25 | 874.21 | 621.41 | 815.56 | 847.84 | 861.41 | 24735.23
60 | 685.41 | 838.9 | 638.05 | 762.36 | 840.5 | 801.67 | 744.08 | 737.24 | 747.79 | 661.86 | 748.82 | 706.62 | 765.98 | 778.47 | 748.32 | 773.89 | 601.19 | 801.17 | 778.85 | 847.93 | 745.19 | 823.12 | 702.63 | 828.44 | 691.51 | 803.84 | 860.22 | 684.01 | 888.5 | 645.59 | 834.74 | 854.14 | 879.96 | 25250.99
61 | 699.73 | 857.57 | 659.84 | 790.23 | 865.83 | 828.08 | 757.79 | 755.31 | 763.21 | 669.56 | 772.56 | 722.4 | 779.46 | 800.07 | 768.96 | 793.62 | 615.43 | 817.52 | 810.36 | 875.51 | 756.86 | 848.98 | 715.34 | 847.68 | 700.62 | 811.22 | 882.62 | 697.52 | 905.58 | 672.79 | 857.12 | 862.87 | 901.62 | 25863.86
62 | 717.55 | 879.9 | 686.16 | 822.76 | 895.03 | 857.94 | 774.61 | 777.37 | 782.29 | 679.26 | 798.11 | 744.24 | 795.55 | 824.9 | 792.51 | 817.04 | 633.4 | 835.83 | 845.28 | 904.71 | 771.63 | 875.65 | 732.79 | 869.63 | 712.29 | 820.52 | 908.9 | 715.64 | 925.88 | 703.01 | 882.73 | 874.7 | 926.44 | 26584.25
63 | 739.3 | 906.3 | 717.28 | 859.59 | 928.18 | 891.15 | 795.17 | 803.94 | 805.69 | 691.5 | 825.34 | 773.24 | 814.67 | 853.12 | 819.04 | 844.37 | 655.79 | 856.17 | 883 | 935.41 | 790.25 | 903.11 | 755.79 | 894.67 | 726.99 | 832.26 | 939.24 | 739.35 | 949.85 | 736.24 | 911.53 | 890.48 | 954.45 | 27422.46
64 | 765.44 | 937.18 | 753.37 | 900.14 | 965.23 | 927.58 | 820.17 | 835.54 | 834.14 | 707.01 | 854.27 | 810.38 | 837.24 | 884.89 | 848.58 | 875.77 | 683.36 | 878.68 | 922.88 | 967.69 | 813.65 | 931.6 | 785.08 | 923.27 | 745.24 | 847.12 | 973.73 | 769.68 | 977.94 | 772.56 | 943.46 | 911.3 | 985.74 | 28389.91
65 | 796.37 | 973.01 | 794.47 | 943.78 | 1006.2 | 967.15 | 850.46 | 872.65 | 868.42 | 726.77 | 885.23 | 856.34 | 863.86 | 920.48 | 881.28 | 911.39 | 716.88 | 903.63 | 964.45 | 1001.92 | 843.03 | 961.8 | 821.25 | 956.21 | 767.63 | 866.04 | 1012.49 | 807.72 | 1010.7 | 812.14 | 978.54 | 938.57 | 1020.57 | 29501.43
66 | 832.55 | 1014.45 | 840.65 | 990.15 | 1051.29 | 1010.12 | 887.06 | 915.78 | 909.41 | 752.17 | 919.03 | 911.42 | 895.37 | 960.44 | 917.5 | 951.48 | 757.16 | 931.6 | 1007.75 | 1038.87 | 879.85 | 994.94 | 864.73 | 994.63 | 794.96 | 890.37 | 1055.79 | 854.55 | 1048.87 | 855.27 | 1017.08 | 974.11 | 1059.56 | 30778.96
67 | 874.61 | 1062.51 | 892.18 | 1039.69 | 1101.23 | 1057.41 | 931.26 | 965.56 | 958.11 | 785.12 | 957.12 | 975.49 | 933.09 | 1005.85 | 958.15 | 996.64 | 804.98 | 963.64 | 1053.7 | 1079.86 | 926.01 | 1032.84 | 915.83 | 1040.27 | 828.48 | 922.12 | 1104.28 | 911.23 | 1093.56 | 902.47 | 1059.86 | 1020.27 | 1103.95 | 32257.37
68 | 923.63 | 1118.89 | 949.96 | 1094.04 | 1157.62 | 1110.91 | 984.78 | 1022.98 | 1015.82 | 828.35 | 1001.82 | 1048.2 | 979.14 | 1058.64 | 1004.94 | 1048.14 | 861.18 | 1001.5 | 1104.37 | 1126.99 | 983.92 | 1078.12 | 975.04 | 1095.55 | 870.22 | 964.3 | 1159.35 | 978.77 | 1146.53 | 954.69 | 1108.49 | 1080.06 | 1155.88 | 33992.82
69 | 981.47 | 1186.42 | 1016.04 | 1156.59 | 1223.47 | 1173.94 | 1050.07 | 1089.8 | 1084.41 | 885.72 | 1056.52 | 1129.67 | 1036.93 | 1121.92 | 1060.78 | 1108.39 | 926.88 | 1047.95 | 1163.23 | 1183.42 | 1056.76 | 1134.33 | 1043.48 | 1163.78 | 923.54 | 1021.35 | 1223.57 | 1058.44 | 1210.63 | 1013.71 | 1165.7 | 1157.31 | 1218.76 | 36074.98
70 | 1051.44 | 1269.7 | 1094.42 | 1232.79 | 1303.71 | 1251.7 | 1130.82 | 1169.21 | 1166.87 | 962.74 | 1126.02 | 1221.6 | 1111.86 | 1200.46 | 1130.23 | 1181.52 | 1004.14 | 1107.22 | 1235.34 | 1253.74 | 1148.92 | 1206.32 | 1123.66 | 1249.45 | 993.78 | 1099.55 | 1301.2 | 1152.26 | 1290.34 | 1082.75 | 1235.82 | 1256.95 | 1297.72 | 38644.25
71 | 1139.05 | 1375.99 | 1191.83 | 1330.48 | 1405.86 | 1351.9 | 1232.77 | 1266.74 | 1268.2 | 1067.33 | 1217.17 | 1328.82 | 1212.3 | 1301.21 | 1219.92 | 1274.03 | 1097.05 | 1185.64 | 1327.71 | 1344.57 | 1266.78 | 1300.92 | 1220.4 | 1358.91 | 1088.85 | 1207.4 | 1398.81 | 1264.3 | 1392.56 | 1167.14 | 1325.22 | 1385.52 | 1400.17 | 41915.55
72 | 1252.97 | 1516.3 | 1318.53 | 1460.47 | 1540.76 | 1485.8 | 1364.94 | 1391.37 | 1396.64 | 1210.74 | 1339.78 | 1460.8 | 1350.59 | 1434.06 | 1339.28 | 1395.7 | 1213.16 | 1292.77 | 1449.97 | 1465.27 | 1420.04 | 1427.78 | 1341.78 | 1501.52 | 1219.65 | 1355.61 | 1526.1 | 1402.32 | 1527.55 | 1275.07 | 1443.13 | 1552.12 | 1536.74 | 46209.31
73 | 1406.05 | 1706.65 | 1489.11 | 1637.57 | 1723.78 | 1669.56 | 1541.19 | 1556.79 | 1565.23 | 1408.8 | 1508.07 | 1632.75 | 1543.84 | 1612.96 | 1501.5 | 1560.7 | 1364.88 | 1443.06 | 1615.42 | 1629.02 | 1623.67 | 1600.66 | 1500.21 | 1691.07 | 1399.88 | 1556.91 | 1696.99 | 1579.74 | 1710.2 | 1418.35 | 1602.75 | 1770.2 | 1722.61 | 51990.17
74 | 1616.77 | 1969.68 | 1723.62 | 1882.21 | 1976.51 | 1926.09 | 1782.01 | 1782.98 | 1793.59 | 1683.45 | 1742.46 | 1866.49 | 1814.01 | 1857.45 | 1724.99 | 1789.22 | 1570.5 | 1658.5 | 1842.66 | 1854.26 | 1900.67 | 1839.09 | 1713.69 | 1947.64 | 1646.11 | 1826.47 | 1931.45 | 1817.2 | 1961.68 | 1613.48 | 1822.93 | 2060.5 | 1979.42 | 59917.78
75 | 1911.2 | 2337.11 | 2049.65 | 2222.94 | 2329.29 | 2287.34 | 2116.83 | 2098.28 | 2109.96 | 2064.77 | 2072.16 | 2192.03 | 2187.7 | 2194.89 | 2035.57 | 2110.02 | 1855.25 | 1972.53 | 2157.88 | 2166.89 | 2285.37 | 2170.7 | 2007.8 | 2299.74 | 1979.33 | 2184.89 | 2258.35 | 2143.94 | 2311.89 | 1883.32 | 2130.45 | 2454.97 | 2337.95 | 70920.99
76 | 2325.85 | 2853.54 | 2505.38 | 2699.61 | 2824.49 | 2797.4 | 2586.98 | 2542.23 | 2553.81 | 2593.56 | 2538.33 | 2650.53 | 2697.82 | 2663.11 | 2469.31 | 2564.08 | 2253.23 | 2434.86 | 2597.93 | 2603.34 | 2826.93 | 2634.18 | 2418.5 | 2787.47 | 2429.85 | 2664.76 | 2719.97 | 2599.87 | 2802.74 | 2259.57 | 2562.82 | 3001.09 | 2841.44 | 86304.58
77 | 2910.45 | 3580.13 | 3142.24 | 3366.38 | 3519.29 | 3514.79 | 3248.52 | 3168.43 | 3178.4 | 3323.81 | 3196.9 | 3297.84 | 3389.23 | 3312.55 | 3074.9 | 3208.44 | 2809.72 | 3114.09 | 3213.6 | 3214.06 | 3591.15 | 3282.19 | 2995.46 | 3465.78 | 3044.37 | 3317.75 | 3377.81 | 3238.21 | 3491.15 | 2785.7 | 3171.34 | 3764.3 | 3548.93 | 107857.91
78 | 3727.94 | 4591.64 | 4021.84 | 4289.76 | 4483.03 | 4508.81 | 4169.25 | 4044.04 | 4050.35 | 4318.23 | 4115.34 | 4204.81 | 4323.34 | 4204.44 | 3912.8 | 4116.35 | 3579.61 | 4090.02 | 4069.36 | 4064.22 | 4652.11 | 4180.31 | 3803.04 | 4404.18 | 3887.55 | 4213.78 | 4315.68 | 4124.23 | 4446.13 | 3517.66 | 4022.27 | 4824.77 | 4533.99 | 137810.88
79 | 4817.67 | 5917.44 | 5166.4 | 5515.35 | 5760 | 5806.82 | 5376.73 | 5224.5 | 5226.54 | 5563.87 | 5319.76 | 5434.55 | 5545.29 | 5385.32 | 5036.43 | 5351.71 | 4597.75 | 5393.83 | 5224.53 | 5218.98 | 5992.33 | 5378.16 | 4908.67 | 5661.44 | 5015.54 | 5415.62 | 5617.26 | 5306.93 | 5712.17 | 4511.54 | 5187.74 | 6225.07 | 5851.42 | 177667.36
80 | 6031.41 | 7344.35 | 6393.91 | 6913.54 | 7208.06 | 7212.28 | 6682.38 | 6606 | 6611.78 | 6804.7 | 6618.49 | 6891.87 | 6921.63 | 6746.55 | 6362.84 | 6811.85 | 5747.82 | 6825.05 | 6620.26 | 6638.39 | 7317.98 | 6758.73 | 6284.04 | 7131.47 | 6333.21 | 6856.39 | 7218.23 | 6672.04 | 7152.67 | 5731.91 | 6645.72 | 7767.11 | 7369.64 | 223232.3
81 | 7035.93 | 8451.32 | 7346.68 | 8112.83 | 8439.04 | 8331.19 | 7710.54 | 7837.02 | 7865.61 | 7666.45 | 7632.8 | 8217.26 | 8078.07 | 7938.23 | 7571.72 | 8125.89 | 6714.96 | 7987.96 | 7937.89 | 8019.19 | 8207.13 | 7960.61 | 7638.03 | 8444.01 | 7508.44 | 8201.64 | 8732.89 | 7863.77 | 8388.53 | 6914.08 | 8096.38 | 9027.21 | 8689.22 | 262692.52
82 | 7499.84 | 8847.93 | 7690.06 | 8717.3 | 9046.31 | 8786.25 | 8108.24 | 8525.79 | 8599.91 | 7833.74 | 8009.46 | 8996.69 | 8622.78 | 8564.7 | 8282.49 | 8876.3 | 7171.88 | 8506.32 | 8764.62 | 8945.53 | 8316 | 8589.83 | 8558.41 | 9179.22 | 8167.65 | 9027.59 | 9677.81 | 8492.05 | 9017.29 | 7696.74 | 9105.94 | 9590.32 | 9379.47 | 283194.46
83 | 7204.57 | 8317.63 | 7234.75 | 8455.44 | 8755.38 | 8346.49 | 7669.15 | 8383.61 | 8515.51 | 7173.5 | 7548.23 | 8910.93 | 8299.47 | 8338.47 | 8193.85 | 8752.88 | 6905.01 | 8143.31 | 8773.54 | 9067.82 | 7521.62 | 8368.41 | 8696.34 | 9023.26 | 8032.83 | 9004.91 | 9689.64 | 8279.79 | 8759.43 | 7780.11 | 9299.16 | 9187.91 | 9139.62 | 275772.57
84 | 6166.89 | 6948.15 | 6051.7 | 7316.34 | 7561.37 | 7062.72 | 6455 | 7362.27 | 7541.68 | 5833.39 | 6321.87 | 7886.05 | 7119.93 | 7231.91 | 7234.03 | 7698.64 | 5924.42 | 6935.88 | 7861.09 | 8242.56 | 6020.61 | 7272.42 | 7909.05 | 7921.15 | 7055.28 | 8038.18 | 8667.23 | 7201.51 | 7597.67 | 7047.48 | 8506.23 | 7849.21 | 7946.29 | 239788.2
85 | 4684.39 | 5147.5 | 4490 | 5607.85 | 5786.43 | 5305.6 | 4821.24 | 5730.37 | 5925.48 | 4232.53 | 4694.53 | 6186.99 | 5408.86 | 5546.95 | 5652.93 | 5990.34 | 4500.96 | 5238.42 | 6244.89 | 6651.02 | 4296.14 | 5596.45 | 6376.01 | 6161.96 | 5490.94 | 6360.89 | 6867.48 | 5548.13 | 5841.39 | 5663.74 | 6893.29 | 5955.17 | 6121.11 | 185019.98
86 | 3253.73 | 3508.99 | 3063.5 | 3884.55 | 4004.32 | 3655.91 | 3306.9 | 4021.85 | 4194.19 | 2878.62 | 3203.46 | 4373.17 | 3720.43 | 3841.04 | 3976.23 | 4183.16 | 3096.22 | 3620.83 | 4430.16 | 4788.97 | 2888.98 | 3887.24 | 4583.13 | 4322.53 | 3851.86 | 4495.9 | 4858.76 | 3860.76 | 4058.57 | 4061.11 | 4962.17 | 4134.48 | 4258.33 | 129230.05
87 | 2206.12 | 2358.25 | 2058.66 | 2599.77 | 2679.57 | 2469.97 | 2228.3 | 2711.37 | 2842.09 | 1951.43 | 2150.9 | 2962.54 | 2484.43 | 2570.24 | 2681.86 | 2780.59 | 2067.47 | 2447.46 | 2969.92 | 3241.03 | 1949.06 | 2602.89 | 3097.18 | 2913.94 | 2594.19 | 3005.9 | 3250.98 | 2592.25 | 2721.86 | 2739.96 | 3331.35 | 2802.93 | 2856.22 | 86920.68
88 | 1517.88 | 1617.23 | 1410.71 | 1764.35 | 1818.86 | 1697.41 | 1528.79 | 1846.21 | 1940.15 | 1351.58 | 1471.6 | 2024.54 | 1689.11 | 1745.16 | 1825.63 | 1849.8 | 1404.56 | 1680.46 | 2001.53 | 2192.91 | 1348.73 | 1764.73 | 2094.22 | 1984.99 | 1764.03 | 2021.88 | 2187.77 | 1763.41 | 1849.56 | 1850.64 | 2224 | 1928.04 | 1939.21 | 59099.68
89 | 1071.89 | 1141.73 | 994.69 | 1231.07 | 1269.28 | 1198.74 | 1078.26 | 1289.74 | 1356.62 | 962.88 | 1035.28 | 1419.09 | 1183.91 | 1218.57 | 1274.7 | 1256.46 | 982.62 | 1184.92 | 1384.42 | 1518.27 | 962.51 | 1229.06 | 1450.3 | 1387.74 | 1230.5 | 1396.13 | 1510.88 | 1232.51 | 1291.35 | 1280.06 | 1518.65 | 1360.9 | 1351.85 | 41255.58
90 | 779.22 | 831.33 | 723.03 | 886.12 | 913.7 | 872.13 | 783.52 | 928.14 | 975.84 | 706.5 | 749.69 | 1024.8 | 857.71 | 877.68 | 916.6 | 879.64 | 709.81 | 860.41 | 987.53 | 1082.16 | 708.81 | 882.44 | 1034.8 | 999.62 | 884.27 | 994.4 | 1075.9 | 888.15 | 929.48 | 911.8 | 1070.09 | 988.54 | 970.99 | 29684.85
91 | 583.02 | 623.83 | 541.37 | 657.77 | 678.33 | 653.36 | 586.2 | 688.07 | 722.32 | 533.41 | 556.87 | 762.71 | 641.69 | 651.83 | 678.77 | 637.3 | 529.15 | 643.21 | 726.58 | 794.89 | 537.85 | 652.85 | 761.27 | 741.92 | 654.8 | 730.72 | 790.26 | 659.75 | 689.52 | 669.27 | 779.55 | 738.73 | 718.42 | 22015.59
92 | 448.29 | 481.51 | 416.77 | 502.58 | 518.39 | 503.16 | 450.72 | 524.61 | 549.44 | 413.57 | 421.46 | 584.18 | 494.62 | 498.3 | 516.74 | 477.88 | 406.28 | 494.21 | 550.44 | 601.07 | 419.55 | 496.28 | 576.64 | 566.51 | 498.82 | 553.04 | 597.83 | 504.39 | 526.3 | 505.54 | 586.3 | 567.04 | 546.53 | 16798.99
93 | 353.48 | 381.35 | 329.12 | 394.29 | 406.77 | 397.43 | 355.3 | 410.36 | 428.6 | 328.5 | 322.5 | 459.48 | 391.75 | 391.19 | 403.46 | 370.05 | 320.44 | 389.4 | 428.33 | 466.92 | 335.53 | 385.7 | 448.72 | 444.04 | 390 | 430.08 | 464.78 | 396.02 | 412.41 | 392.13 | 454 | 446.14 | 426.43 | 13154.7
94 | 285.15 | 309.09 | 265.98 | 316.77 | 326.87 | 321.2 | 286.42 | 328.45 | 342.05 | 266.63 | 247.97 | 370.2 | 317.92 | 314.56 | 322.28 | 294.95 | 258.93 | 313.84 | 341.48 | 371.69 | 274.35 | 304.44 | 357.84 | 356.41 | 312.14 | 342.8 | 370.43 | 318.55 | 330.96 | 311.61 | 360.82 | 358.98 | 340.18 | 10541.94
95 | 234.77 | 255.73 | 219.45 | 259.94 | 268.28 | 264.95 | 235.55 | 268.33 | 278.62 | 220.62 | 191.12 | 304.77 | 263.64 | 258.43 | 262.73 | 241.1 | 213.79 | 258.1 | 278.22 | 302.46 | 228.76 | 242.34 | 291.75 | 292.22 | 255.1 | 279.35 | 301.9 | 261.88 | 271.35 | 253.09 | 293.41 | 294.75 | 276.41 | 8622.91
96 | 196.84 | 215.46 | 184.45 | 217.36 | 224.39 | 222.55 | 197.2 | 223.22 | 231.12 | 185.7 | 147.84 | 255.76 | 222.82 | 216.42 | 218.13 | 201.45 | 179.96 | 216.1 | 231.09 | 250.99 | 194.06 | 193.41 | 242.63 | 244.15 | 212.4 | 232.18 | 250.99 | 219.49 | 226.76 | 209.64 | 243.43 | 246.42 | 227.82 | 7182.23
97 | 167.72 | 184.48 | 157.6 | 184.83 | 190.85 | 189.96 | 167.75 | 188.73 | 194.86 | 158.7 | 115.27 | 218.31 | 191.49 | 184.34 | 184.08 | 171.54 | 154.09 | 183.83 | 195.27 | 211.95 | 167.15 | 154.2 | 205.38 | 207.42 | 179.78 | 196.38 | 212.38 | 187.11 | 192.73 | 176.73 | 205.54 | 209.36 | 189.71 | 6079.52
98 | 144.99 | 160.24 | 136.67 | 159.53 | 164.78 | 164.47 | 144.78 | 161.92 | 166.7 | 137.49 | 91.1 | 189.17 | 167 | 159.4 | 157.63 | 148.48 | 133.97 | 158.63 | 167.54 | 181.81 | 145.93 | 122.68 | 176.63 | 178.84 | 154.43 | 168.68 | 182.53 | 161.91 | 166.3 | 151.38 | 176.25 | 180.46 | 159.15 | 5221.47
99 | 126.98 | 141 | 120.12 | 139.59 | 144.21 | 144.25 | 126.61 | 140.78 | 144.52 | 120.6 | 73.37 | 166.15 | 147.56 | 139.73 | 136.79 | 130.38 | 118.08 | 138.69 | 145.73 | 158.18 | 128.99 | 97.51 | 154.08 | 156.27 | 134.43 | 146.91 | 159.07 | 141.98 | 145.45 | 131.54 | 153.22 | 157.59 | 134.29 | 4544.65
100 | 112.56 | 125.58 | 106.89 | 123.68 | 127.79 | 128.04 | 112.1 | 123.93 | 126.86 | 107 | 60.47 | 147.73 | 131.95 | 124.03 | 120.17 | 115.99 | 105.37 | 122.74 | 128.38 | 139.43 | 115.33 | 77.66 | 136.15 | 138.22 | 118.46 | 129.56 | 140.4 | 126.01 | 128.8 | 115.82 | 134.89 | 139.29 | 113.98 | 4005.26
101 | 100.92 | 113.13 | 96.23 | 110.89 | 114.58 | 114.95 | 100.42 | 110.4 | 112.68 | 95.99 | 51.1 | 132.88 | 119.3 | 111.4 | 106.81 | 104.43 | 95.12 | 109.91 | 114.46 | 124.4 | 104.24 | 62.24 | 121.75 | 123.68 | 105.6 | 115.61 | 125.4 | 113.13 | 115.39 | 103.28 | 120.16 | 124.53 | 97.41 | 3572.42
102 | 91.49 | 103.05 | 87.61 | 100.58 | 103.89 | 104.36 | 90.99 | 99.48 | 101.26 | 87.05 | 44.28 | 120.83 | 109.02 | 101.18 | 96.02 | 95.1 | 86.8 | 99.55 | 103.23 | 112.31 | 95.23 | 50.45 | 110.12 | 111.9 | 95.21 | 104.34 | 113.3 | 102.69 | 104.52 | 93.21 | 108.27 | 112.57 | 83.99 | 3223.88
103 | 83.85 | 94.88 | 80.63 | 92.26 | 95.22 | 95.78 | 83.37 | 90.67 | 92.05 | 79.77 | 39.29 | 111.05 | 100.64 | 92.92 | 87.29 | 87.55 | 80.05 | 91.18 | 94.16 | 102.54 | 87.9 | 41.58 | 100.69 | 102.35 | 86.79 | 95.2 | 103.51 | 94.22 | 95.7 | 85.11 | 98.65 | 102.87 | 73.23 | 2942.95
104 | 77.66 | 88.29 | 74.98 | 85.55 | 88.21 | 88.86 | 77.21 | 83.56 | 84.63 | 73.86 | 35.61 | 103.12 | 93.82 | 86.23 | 80.24 | 81.47 | 74.56 | 84.44 | 86.83 | 94.66 | 81.95 | 34.99 | 93.05 | 94.6 | 79.99 | 87.8 | 95.6 | 87.37 | 88.54 | 78.61 | 90.88 | 95.01 | 64.69 | 2716.87
105 | 72.66 | 82.99 | 70.43 | 80.16 | 82.55 | 83.29 | 72.26 | 77.84 | 78.67 | 69.09 | 32.89 | 96.7 | 88.28 | 80.85 | 74.56 | 76.59 | 70.13 | 79.02 | 80.94 | 88.31 | 77.15 | 30.19 | 86.86 | 88.34 | 74.51 | 81.84 | 89.24 | 81.86 | 82.74 | 73.41 | 84.63 | 88.67 | 57.99 | 2535.64
106 | 68.66 | 78.76 | 66.79 | 75.87 | 78.01 | 78.84 | 68.3 | 73.26 | 73.91 | 65.26 | 30.89 | 91.55 | 83.82 | 76.55 | 70.02 | 72.7 | 66.58 | 74.7 | 76.22 | 83.22 | 73.29 | 26.75 | 81.89 | 83.31 | 70.12 | 77.06 | 84.16 | 77.44 | 78.08 | 69.28 | 79.64 | 83.57 | 52.8 | 2391.3
107 | 65.47 | 75.41 | 63.9 | 72.48 | 74.4 | 75.32 | 65.16 | 69.63 | 70.14 | 62.21 | 29.43 | 87.44 | 80.25 | 73.13 | 66.41 | 69.63 | 63.75 | 71.27 | 72.47 | 79.17 | 70.23 | 24.35 | 77.92 | 79.3 | 66.63 | 73.26 | 80.13 | 73.94 | 74.36 | 66.03 | 75.67 | 79.51 | 48.83 | 2277.23
108 | 62.97 | 72.8 | 61.64 | 69.84 | 71.56 | 72.56 | 62.7 | 66.78 | 67.18 | 59.81 | 28.4 | 84.19 | 77.42 | 70.45 | 63.57 | 67.25 | 61.52 | 68.59 | 69.53 | 75.98 | 67.82 | 22.74 | 74.77 | 76.13 | 63.88 | 70.27 | 76.99 | 71.2 | 71.42 | 63.5 | 72.57 | 76.3 | 45.85 | 2188.18
109 | 61.03 | 70.81 | 59.9 | 67.83 | 69.37 | 70.45 | 60.81 | 64.58 | 64.9 | 57.96 | 27.71 | 81.68 | 75.23 | 68.39 | 61.39 | 65.45 | 59.81 | 66.53 | 67.26 | 73.5 | 65.96 | 21.74 | 72.33 | 73.67 | 61.77 | 67.97 | 74.57 | 69.1 | 69.14 | 61.57 | 70.18 | 73.82 | 43.67 | 2120.08
110 | 59.57 | 69.33 | 58.6 | 66.33 | 67.73 | 68.88 | 59.4 | 62.93 | 63.19 | 56.57 | 27.3 | 79.78 | 73.56 | 66.85 | 59.74 | 64.14 | 58.53 | 64.99 | 65.55 | 71.63 | 64.57 | 21.22 | 70.47 | 71.8 | 60.17 | 66.23 | 72.76 | 67.53 | 67.41 | 60.15 | 68.4 | 71.95 | 42.15 | 2069.41
111 | 58.52 | 68.29 | 57.67 | 65.28 | 66.54 | 67.76 | 58.4 | 61.74 | 61.96 | 55.58 | 27.13 | 78.39 | 72.34 | 65.75 | 58.56 | 63.23 | 57.62 | 63.9 | 64.32 | 70.26 | 63.58 | 21.06 | 69.11 | 70.43 | 59.02 | 64.99 | 71.48 | 66.41 | 66.16 | 59.14 | 67.12 | 70.59 | 41.15 | 2033.48
112 | 57.81 | 67.62 | 57.07 | 64.61 | 65.75 | 67.03 | 57.75 | 60.94 | 61.13 | 54.91 | 27.17 | 77.45 | 71.51 | 65.03 | 57.76 | 62.68 | 57.02 | 63.17 | 63.49 | 69.33 | 62.93 | 21.2 | 68.18 | 69.5 | 58.24 | 64.15 | 70.63 | 65.67 | 65.31 | 58.5 | 66.27 | 69.67 | 40.6 | 2010.08
113 | 57.4 | 67.28 | 56.75 | 64.26 | 65.31 | 66.64 | 57.39 | 60.48 | 60.65 | 54.54 | 27.38 | 76.88 | 71.02 | 64.63 | 57.3 | 62.44 | 56.69 | 62.77 | 63.02 | 68.77 | 62.57 | 21.59 | 67.61 | 68.94 | 57.79 | 63.67 | 70.17 | 65.27 | 64.81 | 58.16 | 65.79 | 69.12 | 40.41 | 1997.5
114 | 57.25 | 67.22 | 56.67 | 64.2 | 65.16 | 66.53 | 57.3 | 60.32 | 60.48 | 54.42 | 27.76 | 76.65 | 70.82 | 64.51 | 57.13 | 62.46 | 56.59 | 62.65 | 62.85 | 68.54 | 62.47 | 22.17 | 67.36 | 68.7 | 57.62 | 63.51 | 70.03 | 65.15 | 64.62 | 58.09 | 65.63 | 68.91 | 40.53 | 1994.3
115 | 57.32 | 67.4 | 56.8 | 64.39 | 65.28 | 66.69 | 57.44 | 60.41 | 60.57 | 54.53 | 28.28 | 76.71 | 70.88 | 64.65 | 57.22 | 62.72 | 56.71 | 62.77 | 62.94 | 68.6 | 62.61 | 22.93 | 67.39 | 68.75 | 57.7 | 63.62 | 70.19 | 65.29 | 64.69 | 58.26 | 65.76 | 68.99 | 40.92 | 1999.41
116 | 57.6 | 67.81 | 57.13 | 64.79 | 65.64 | 67.07 | 57.78 | 60.74 | 60.9 | 54.83 | 28.94 | 77.02 | 71.17 | 65 | 57.53 | 63.2 | 57.01 | 63.12 | 63.28 | 68.91 | 62.94 | 23.84 | 67.67 | 69.05 | 58 | 63.97 | 70.6 | 65.65 | 65.01 | 58.63 | 66.14 | 69.32 | 41.54 | 2011.83
117 | 58.04 | 68.38 | 57.6 | 65.37 | 66.17 | 67.62 | 58.28 | 61.24 | 61.41 | 55.29 | 29.67 | 77.54 | 71.64 | 65.53 | 58.02 | 63.83 | 57.46 | 63.63 | 63.8 | 69.42 | 63.43 | 24.82 | 68.15 | 69.55 | 58.47 | 64.5 | 71.21 | 66.18 | 65.52 | 59.16 | 66.71 | 69.86 | 42.31 | 2029.81
118 | 58.46 | 68.91 | 58.04 | 65.9 | 66.68 | 68.13 | 58.74 | 61.72 | 61.9 | 55.72 | 30.24 | 78.05 | 72.11 | 66.03 | 58.49 | 64.39 | 57.88 | 64.12 | 64.3 | 69.92 | 63.9 | 25.55 | 68.62 | 70.05 | 58.92 | 65.02 | 71.78 | 66.68 | 66.01 | 59.65 | 67.25 | 70.38 | 42.94 | 2046.48
Grand Total | 109323.28 | 128289.85 | 108327.95 | 127049.53 | 130669.93 | 126594.51 | 121740.21 | 125925.91 | 127122.85 | 111370.09 | 110511.86 | 127214.77 | 120336.97 | 120105.77 | 118152.35 | 126485.61 | 102274.83 | 123448 | 126405.96 | 131914.66 | 120298.12 | 119622.45 | 123563.12 | 130025.09 | 115175.28 | 132117.39 | 138801.58 | 121817.75 | 136398.09 | 112271.77 | 133057.66 | 140668.24 | 133224.48 | 4080305.91

## 50..800_amp
Sum of F4 | Well
Cycles | A02 | A03 | A04 | A05 | A06 | A07 | A08 | A09 | A10 | A11 | B02 | B03 | B04 | B05 | B06 | B07 | B08 | B09 | B10 | B11 | B12 | C02 | C03 | C04 | C05 | C06 | C07 | C08 | C09 | C10 | C11 | C12 | Grand Total
1 | 5887.3 | 5871 | 5944.5 | 6000.5 | 6146 | 5364.1 | 6442.1 | 6192.8 | 6030.2 | 5978.4 | 5534 | 5851.5 | 6056.7 | 6096.5 | 5909.4 | 5496.6 | 6239.5 | 6492.4 | 6084.4 | 5848.2 | 5695.4 | 5729.1 | 5571 | 5802.6 | 5825 | 6279.7 | 5918.8 | 6407.6 | 5999.8 | 5892.3 | 6478.6 | 5938.3 | 191004.3
2 | 5963.89 | 5978.1 | 6034.24 | 6064.68 | 6200.81 | 5483.42 | 6526.2 | 6327.55 | 6136.79 | 6027.82 | 5600.6 | 5960.67 | 6139.5 | 6164.74 | 5963.33 | 5570.51 | 6284.97 | 6542.35 | 6139.77 | 5870.3 | 5711.15 | 5755.12 | 5629.55 | 5891.66 | 5862.95 | 6346.5 | 5967.17 | 6471.25 | 6043.63 | 5952.8 | 6501.95 | 5906.05 | 193020.02
3 | 5967.72 | 5984.77 | 6047.49 | 6078.64 | 6209.5 | 5490.87 | 6538.11 | 6331.5 | 6146.75 | 6025.13 | 5613.15 | 5966.49 | 6146.81 | 6178.46 | 5963 | 5574.68 | 6284.59 | 6541.02 | 6143.9 | 5860.17 | 5706.12 | 5761.04 | 5634.87 | 5904.1 | 5868.13 | 6369.32 | 5972.13 | 6471.54 | 6047.11 | 5952.8 | 6492.34 | 5887.91 | 193160.16
4 | 5973 | 5993.56 | 6058.98 | 6091.49 | 6218.18 | 5502.11 | 6549.83 | 6343.04 | 6159.63 | 6025.65 | 5622.91 | 5973.58 | 6155.47 | 6191.42 | 5964.67 | 5577.86 | 6284.39 | 6537.88 | 6145.32 | 5851.95 | 5706.45 | 5767.53 | 5637.84 | 5917.4 | 5875.1 | 6373.5 | 5975.62 | 6470.62 | 6050.03 | 5954.91 | 6484.13 | 5866.21 | 193300.26
5 | 5975.24 | 6004.17 | 6074.12 | 6104.08 | 6223.9 | 5516.46 | 6562.95 | 6360.23 | 6172.1 | 6028.27 | 5636.11 | 5983.78 | 6164.31 | 6203.33 | 5966.7 | 5582.95 | 6284.22 | 6535.46 | 6147.95 | 5846.66 | 5706.73 | 5775.52 | 5649.02 | 5930.29 | 5880.14 | 6378.39 | 5982.23 | 6471.16 | 6050.12 | 5955.16 | 6477.47 | 5848.84 | 193478.06
6 | 5977.33 | 6011.48 | 6089.18 | 6116.39 | 6233.73 | 5531.18 | 6572.27 | 6372.05 | 6185.18 | 6032.38 | 5645.2 | 5991.28 | 6169.61 | 6210.48 | 5968.13 | 5587.01 | 6283.25 | 6533.28 | 6149.84 | 5841.3 | 5710.29 | 5780.56 | 5659.76 | 5942.96 | 5883.34 | 6382.95 | 5987.9 | 6473.12 | 6049.3 | 5954.68 | 6473.47 | 5835.58 | 193634.46
7 | 5982.79 | 6016.59 | 6102.64 | 6129.74 | 6244.37 | 5543.42 | 6584.45 | 6386.42 | 6198.99 | 6037.92 | 5654.46 | 5996.52 | 6174.66 | 6218.1 | 5969.95 | 5588.47 | 6280.89 | 6530.86 | 6150.3 | 5837.59 | 5714.81 | 5785.24 | 5665 | 5951.21 | 5887.44 | 6386.2 | 5990.23 | 6471.97 | 6048.77 | 5954.01 | 6469.25 | 5822.98 | 193776.24
8 | 5989.77 | 6023.18 | 6117.28 | 6143.77 | 6254.3 | 5557.69 | 6597.08 | 6401.75 | 6211.82 | 6043.02 | 5662.64 | 6001.66 | 6177.7 | 6225.13 | 5970.13 | 5590.02 | 6278.43 | 6528.55 | 6151.73 | 5833.7 | 5719.86 | 5787.72 | 5667.97 | 5959.97 | 5892.74 | 6391.66 | 5990.11 | 6467.32 | 6048.95 | 5953.14 | 6464.84 | 5810.02 | 193913.65
9 | 5996.61 | 6028.77 | 6128.09 | 6157.54 | 6266.95 | 5568.08 | 6609.11 | 6418.16 | 6223.5 | 6049.79 | 5667.77 | 6005.45 | 6180.94 | 6229.41 | 5970.55 | 5590.68 | 6277.26 | 6526.2 | 6153.41 | 5828.7 | 5726.33 | 5790.28 | 5671.68 | 5965.2 | 5896.88 | 6396.39 | 5991.73 | 6463.07 | 6048.79 | 5952.46 | 6460.16 | 5797.18 | 194037.12
10 | 6000.1 | 6033.87 | 6141.14 | 6170.82 | 6280.56 | 5578.93 | 6620.89 | 6428.82 | 6235.01 | 6056.14 | 5673.84 | 6008 | 6184.54 | 6233.15 | 5972.55 | 5592.5 | 6275.08 | 6522.79 | 6153.65 | 5824.63 | 5733.2 | 5791.25 | 5675.49 | 5970.8 | 5899 | 6398.8 | 5994.71 | 6459.78 | 6048.75 | 5952.52 | 6455.77 | 5784.81 | 194151.89
11 | 6004.01 | 6039.03 | 6154.72 | 6178.46 | 6288.96 | 5586.36 | 6631.79 | 6440.49 | 6241.91 | 6060.56 | 5680.79 | 6008.06 | 6187.91 | 6234.96 | 5975.96 | 5594.6 | 6273.35 | 6520.22 | 6154.09 | 5819.93 | 5738.63 | 5790.86 | 5677.79 | 5974.18 | 5901.66 | 6398.86 | 5995.26 | 6456.8 | 6048.76 | 5952.1 | 6450.99 | 5772.32 | 194234.37
12 | 6006.29 | 6042.89 | 6161.95 | 6188.83 | 6295.55 | 5596.3 | 6642.44 | 6448.49 | 6248.18 | 6064.49 | 5684.95 | 6009.04 | 6190.8 | 6235.92 | 5975.57 | 5594.54 | 6273.29 | 6517.11 | 6153.8 | 5817.06 | 5744.58 | 5790.58 | 5677.71 | 5976.58 | 5904.22 | 6398.76 | 5994.66 | 6454.76 | 6048.42 | 5950.5 | 6446.02 | 5764.51 | 194298.79
13 | 6008.39 | 6047.75 | 6170.77 | 6200.95 | 6302.48 | 5607.19 | 6649.28 | 6453.79 | 6255.01 | 6067.61 | 5689.5 | 6009.13 | 6195.9 | 6237.12 | 5975.19 | 5594.71 | 6272.81 | 6514.99 | 6153.64 | 5815.87 | 5748.85 | 5792.82 | 5678.91 | 5979.76 | 5905.14 | 6397.58 | 5994.15 | 6452.76 | 6047.37 | 5949.8 | 6440.46 | 5755.07 | 194364.75
14 | 6010.68 | 6052.17 | 6176.5 | 6212.18 | 6306.46 | 5614.45 | 6657.71 | 6457.63 | 6261.93 | 6068.39 | 5693.87 | 6011.47 | 6197.66 | 6237.18 | 5974.3 | 5594 | 6272.04 | 6512.9 | 6155.28 | 5814.75 | 5753.78 | 5792.52 | 5679.34 | 5983.74 | 5906.94 | 6396.45 | 5994.05 | 6450.99 | 6045.21 | 5948.7 | 6437.02 | 5748.4 | 194418.69
15 | 6012.81 | 6055.4 | 6185.06 | 6219.02 | 6311.64 | 5617.51 | 6663.63 | 6459.77 | 6262.16 | 6067.03 | 5696.21 | 6013.34 | 6199.8 | 6239.44 | 5974.31 | 5594.15 | 6271.92 | 6512.86 | 6157.85 | 5816.44 | 5759.07 | 5792.27 | 5681.49 | 5983.95 | 5906.8 | 6395.7 | 5993.1 | 6448.7 | 6041.89 | 5947.85 | 6435.11 | 5740.1 | 194456.38
16 | 6014.56 | 6057.16 | 6190.97 | 6223.09 | 6316.74 | 5621.04 | 6667.77 | 6461.85 | 6263.44 | 6063.8 | 5699.77 | 6017.81 | 6205.01 | 6242.58 | 5978.06 | 5597.88 | 6274.78 | 6514.79 | 6162.32 | 5821.1 | 5764.86 | 5794.1 | 5683.31 | 5986.01 | 5906.12 | 6394.38 | 5991.77 | 6446.76 | 6039.6 | 5945.47 | 6432.97 | 5734.77 | 194514.64
17 | 6015.35 | 6057.01 | 6197.68 | 6229.12 | 6319.78 | 5622.23 | 6673.58 | 6463.9 | 6263.73 | 6061.37 | 5707.11 | 6026.39 | 6212.52 | 6248.15 | 5986.25 | 5604.79 | 6281.32 | 6520.26 | 6171.76 | 5829.56 | 5772.42 | 5796.18 | 5683.3 | 5987.1 | 5905.85 | 6393.95 | 5991.4 | 6445.37 | 6038.17 | 5943.34 | 6431.03 | 5729.27 | 194609.24
18 | 6018.96 | 6056.48 | 6206.15 | 6235.95 | 6325.43 | 5625.77 | 6677.4 | 6465.07 | 6265.09 | 6061.54 | 5718.37 | 6037.08 | 6222.9 | 6256.83 | 5997.52 | 5612.82 | 6290.88 | 6530.51 | 6185.88 | 5838.95 | 5781.31 | 5798.76 | 5681.92 | 5988.74 | 5907.47 | 6393.92 | 5990.79 | 6445.26 | 6039.31 | 5944.9 | 6430.75 | 5725.86 | 194758.57
19 | 6026.9 | 6061.65 | 6217.47 | 6244.63 | 6335.14 | 5634.53 | 6684.38 | 6467.42 | 6270.78 | 6066.11 | 5733.17 | 6050.29 | 6236.7 | 6269.07 | 6009.41 | 5623.09 | 6303.31 | 6542.77 | 6199.84 | 5849.35 | 5791.02 | 5802.66 | 5686.73 | 5995.33 | 5912.84 | 6399.27 | 5994.61 | 6451.12 | 6044.19 | 5950.29 | 6434.42 | 5726.63 | 195015.12
20 | 6044.4 | 6076.86 | 6238.44 | 6263.74 | 6350.63 | 5650.78 | 6700.53 | 6479.38 | 6284.04 | 6076.29 | 5748.97 | 6065.67 | 6252.67 | 6282.7 | 6021.59 | 5636.32 | 6317.56 | 6554.58 | 6213.18 | 5859.9 | 5800.09 | 5817.16 | 5701.34 | 6008.95 | 5927.27 | 6412.65 | 6007.38 | 6466.16 | 6056.93 | 5964.82 | 6446.55 | 5737.36 | 195464.89
21 | 6082.19 | 6109.31 | 6276.32 | 6301.21 | 6384.57 | 5680.1 | 6737.55 | 6508.1 | 6315.32 | 6103.34 | 5764.62 | 6082.14 | 6268.14 | 6296.42 | 6033.62 | 5649.3 | 6331.58 | 6561.72 | 6225.29 | 5869.76 | 5805.4 | 5846.36 | 5733.62 | 6041.17 | 5958.95 | 6443.87 | 6035.11 | 6497.04 | 6085.33 | 5996.36 | 6479.31 | 5765.44 | 196268.56
22 | 6157.02 | 6176.42 | 6347.77 | 6369.43 | 6458.17 | 5745.04 | 6811.27 | 6572.28 | 6379.6 | 6167.56 | 5774.93 | 6099 | 6281.77 | 6311.27 | 6043.83 | 5660.96 | 6340.73 | 6570.65 | 6236.92 | 5876.06 | 5810.12 | 5907.12 | 5795.97 | 6103.67 | 6020.07 | 6510.29 | 6094.89 | 6560.12 | 6145.24 | 6058.87 | 6546.73 | 5824.38 | 197758.15
23 | 6291.14 | 6306.41 | 6480.95 | 6501.47 | 6591.93 | 5864.24 | 6947.8 | 6698.51 | 6506.48 | 6292.65 | 5788.3 | 6115.58 | 6294.23 | 6321.25 | 6053.05 | 5671.9 | 6352.26 | 6579.96 | 6244.01 | 5883.34 | 5812.96 | 6024.93 | 5912.64 | 6228.65 | 6139.88 | 6640.66 | 6216.83 | 6687.74 | 6263.52 | 6178.5 | 6678.72 | 5947.58 | 200518.07
24 | 6527.27 | 6539.85 | 6719.6 | 6742.56 | 6826.31 | 6077.6 | 7187.62 | 6924.61 | 6732.83 | 6516.21 | 5797.23 | 6125.78 | 6306.03 | 6328.55 | 6058.83 | 5680.41 | 6359.45 | 6586.33 | 6249.26 | 5891.02 | 5809.57 | 6241.93 | 6126.39 | 6459.86 | 6360.7 | 6877.93 | 6434.59 | 6923.06 | 6487.83 | 6397.9 | 6929.2 | 6176.95 | 205403.26
25 | 6910.93 | 6917.38 | 7097.27 | 7120.3 | 7219.93 | 6421.61 | 7592.88 | 7295.78 | 7099.99 | 6878.89 | 5807.09 | 6137.75 | 6316.83 | 6336.71 | 6062.62 | 5683.58 | 6366.97 | 6591.64 | 6254.3 | 5893 | 5805.54 | 6617.81 | 6493.63 | 6843.68 | 6740.74 | 7287.54 | 6809.59 | 7329.32 | 6874.43 | 6779.95 | 7362.13 | 6559.33 | 213509.14
26 | 7469.64 | 7467.43 | 7647.25 | 7663.71 | 7780.31 | 6919.16 | 8170.33 | 7841.01 | 7631.18 | 7398.76 | 5817.53 | 6152.18 | 6327.5 | 6345.45 | 6069.39 | 5689.2 | 6375.76 | 6598.61 | 6261.17 | 5897.25 | 5802.73 | 7208.05 | 7081.67 | 7458.28 | 7332.64 | 7925.07 | 7419.57 | 7966.23 | 7459.26 | 7396.16 | 8021.5 | 7148.4 | 225742.38
27 | 8160.08 | 8167.4 | 8370.9 | 8349.71 | 8470.49 | 7534.35 | 8882.74 | 8533.24 | 8284.33 | 8032.41 | 5831.12 | 6166.31 | 6338.62 | 6355.69 | 6078.28 | 5696.19 | 6386.21 | 6606.56 | 6270.04 | 5902.45 | 5804.41 | 8020.8 | 7884.47 | 8281.58 | 8146.47 | 8796 | 8259.12 | 8845.03 | 8263.64 | 8224.79 | 8918 | 7950.97 | 241812.4
28 | 8938.95 | 8940.07 | 9157.29 | 9117.04 | 9253.59 | 8228.29 | 9696.51 | 9300.88 | 9025.56 | 8743.01 | 5846.34 | 6181.26 | 6350.26 | 6367.31 | 6089.06 | 5705.54 | 6397.12 | 6616.91 | 6279.27 | 5908.33 | 5807.68 | 8988.19 | 8806.57 | 9254.92 | 9096.44 | 9829.86 | 9243 | 9871.71 | 9207.47 | 9185.29 | 9960.14 | 8878.05 | 260271.91
29 | 9724.48 | 9721.42 | 9945.35 | 9891.69 | 10045.76 | 8923.73 | 10533.35 | 10081.16 | 9774 | 9457.67 | 5861.89 | 6197.81 | 6364.47 | 6379.85 | 6101.49 | 5715.48 | 6409.38 | 6623.8 | 6289.17 | 5912.23 | 5812.71 | 10001.67 | 9781.79 | 10299.73 | 10111.37 | 10913.9 | 10265 | 10970.35 | 10219.06 | 10210.07 | 11073.74 | 9856.42 | 279469.99
30 | 10465.44 | 10442.24 | 10676.79 | 10610.93 | 10792.81 | 9568.31 | 11313.47 | 10800.46 | 10470.27 | 10124.67 | 5877.96 | 6214.35 | 6378.32 | 6394.2 | 6114.84 | 5724.45 | 6422.03 | 6634.33 | 6299.14 | 5916.6 | 5815.27 | 10956.28 | 10705.64 | 11290.79 | 11091.15 | 11940.75 | 11227.55 | 12012.39 | 11183.45 | 11185.48 | 12127.82 | 10777.91 | 297556.09
31 | 11111.11 | 11071.54 | 11328.38 | 11253.62 | 11445.98 | 10136.39 | 11991.84 | 11442.1 | 11089.03 | 10716.84 | 5893.64 | 6228.93 | 6390.75 | 6408.18 | 6123.54 | 5733.23 | 6430.12 | 6642.13 | 6308.34 | 5920.7 | 5818.2 | 11789.88 | 11521.35 | 12139.27 | 11947.85 | 12836.35 | 12067.57 | 12921.97 | 12035.92 | 12039.21 | 13044.17 | 11580.39 | 313408.52
32 | 11658.72 | 11600.82 | 11878 | 11793.48 | 12001.84 | 10625.88 | 12575.44 | 11983.01 | 11618.31 | 11228.71 | 5910.53 | 6244.92 | 6402.69 | 6421.46 | 6134.95 | 5743.55 | 6441.71 | 6651.16 | 6313.46 | 5924.73 | 5820.07 | 12481.19 | 12203.37 | 12843.96 | 12642 | 13572.72 | 12765.81 | 13674.59 | 12739.18 | 12743.2 | 13799.7 | 12246.71 | 326685.87
33 | 12116.78 | 12037.97 | 12336.49 | 12236.27 | 12462.95 | 11042.72 | 13068.44 | 12432.24 | 12061.07 | 11665.27 | 5921.33 | 6261.99 | 6413.91 | 6433.43 | 6142.44 | 5753.72 | 6453.67 | 6657.49 | 6319.48 | 5929.1 | 5825.8 | 13024 | 12736.66 | 13246.28 | 13190.83 | 14158.21 | 13325.47 | 14276.7 | 13288.96 | 13302.34 | 14402.74 | 12789.4 | 337314.15
34 | 12497.18 | 12398.51 | 12713.17 | 12604.45 | 12841.59 | 11388.25 | 13471.65 | 12805.85 | 12435.05 | 12037.4 | 5933.94 | 6277.87 | 6426.02 | 6443.84 | 6151.26 | 5762.07 | 6464.38 | 6664.79 | 6323.21 | 5932.51 | 5831.32 | 13343.53 | 13050.32 | 13694.4 | 13499.45 | 14488.54 | 13758.24 | 14618.21 | 13708.21 | 13732.8 | 14744.03 | 13095.82 | 345137.86
35 | 12804.07 | 12696.84 | 13016.56 | 12903.47 | 13150.16 | 11668.51 | 13799.03 | 13109.41 | 12737.5 | 12344.87 | 5947.19 | 6292.73 | 6438.22 | 6454.43 | 6160.21 | 5769.96 | 6473.65 | 6671.83 | 6328.27 | 5933.92 | 5833.79 | 13693.48 | 13387.43 | 14064.87 | 13840.52 | 14852.64 | 14012.32 | 14995.56 | 13954.71 | 13985.91 | 15119.59 | 13440.2 | 351881.85
36 | 13050.38 | 12939.79 | 13263.19 | 13141.6 | 13401.1 | 11890.71 | 14068.01 | 13353.18 | 12983.29 | 12593.36 | 5960.86 | 6307.9 | 6449.26 | 6461.43 | 6168.41 | 5777.26 | 6482.25 | 6677.29 | 6332.05 | 5935.27 | 5832.59 | 13900.03 | 13584 | 14273.12 | 14119.54 | 15068.08 | 14281.8 | 15216 | 14217.13 | 14251.96 | 15341.26 | 13729.54 | 357051.64
37 | 13249.46 | 13077.2 | 13404.75 | 13330.71 | 13602.66 | 12064.17 | 14282.14 | 13551.91 | 13184.23 | 12792.78 | 5968.9 | 6317.79 | 6455.61 | 6470.44 | 6176.84 | 5781.83 | 6487.23 | 6683.15 | 6337.78 | 5936.56 | 5831.48 | 14126 | 13733.75 | 14504.64 | 14337.18 | 15304.32 | 14440.62 | 15451.6 | 14431.16 | 14407.11 | 15584.09 | 13893.28 | 361201.37
38 | 13408.75 | 13233.05 | 13563.69 | 13482.93 | 13766.37 | 12201.57 | 14463.02 | 13717.11 | 13342.19 | 12954.67 | 5979.45 | 6329.01 | 6464.34 | 6479.72 | 6182.02 | 5784.4 | 6492.43 | 6688.78 | 6341.88 | 5936.65 | 5831.83 | 14309.89 | 13909.39 | 14638.84 | 14503.2 | 15502.36 | 14602.95 | 15641.2 | 14597.32 | 14580.7 | 15778.21 | 14068.59 | 364776.51
39 | 13536.98 | 13370.99 | 13699.42 | 13606.8 | 13900.06 | 12315.85 | 14611.46 | 13847.58 | 13469.5 | 13090.04 | 5991.65 | 6339.72 | 6473.44 | 6485.39 | 6189.53 | 5789.69 | 6498.26 | 6692.72 | 6345.41 | 5937.11 | 5834.26 | 14414.01 | 14059.57 | 14785.68 | 14599.67 | 15615.55 | 14738.45 | 15791.51 | 14690.67 | 14684.24 | 15926.61 | 14215.47 | 367547.29
40 | 13640.28 | 13482.67 | 13777.31 | 13705.43 | 14007.13 | 12409.86 | 14729.82 | 13953.5 | 13570.54 | 13203.9 | 6003.36 | 6352.31 | 6483.81 | 6492.27 | 6194.77 | 5793.89 | 6502.35 | 6695.13 | 6349.03 | 5938.67 | 5836.85 | 14526.98 | 14180.38 | 14909.33 | 14705.07 | 15731.81 | 14817.41 | 15877.51 | 14788.76 | 14792.56 | 16013.75 | 14334.63 | 369801.07
41 | 13724.46 | 13574.91 | 13862.31 | 13783.99 | 14094.73 | 12488.02 | 14826.63 | 14035.75 | 13652.35 | 13298.62 | 6013.99 | 6365.63 | 6494.81 | 6499.88 | 6202.69 | 5799.69 | 6506.36 | 6698.6 | 6351.5 | 5938.89 | 5840.92 | 14621.33 | 14250.9 | 14982.08 | 14789.33 | 15828.62 | 14903.36 | 15964.15 | 14845.79 | 14855.82 | 16102.38 | 14431.56 | 371630.05
42 | 13794.73 | 13650.49 | 13913.82 | 13847.75 | 14166.3 | 12534.46 | 14883.21 | 14102.4 | 13722.05 | 13373.54 | 6023.94 | 6377.25 | 6505.66 | 6508.47 | 6209.11 | 5806.57 | 6509.04 | 6701.81 | 6354.75 | 5938.99 | 5843.12 | 14697.2 | 14326.11 | 15062.32 | 14836.68 | 15884.55 | 14972.27 | 16033.08 | 14906.05 | 14925.32 | 16153.42 | 14512.15 | 373076.61
43 | 13824.99 | 13682.03 | 13941.8 | 13874.6 | 14196.87 | 12559.32 | 14914.33 | 14131.43 | 13753.57 | 13406.5 | 6028.9 | 6382.69 | 6510.66 | 6512.94 | 6211.42 | 5809.78 | 6511.12 | 6704.25 | 6356.95 | 5939.99 | 5844.74 | 14729.77 | 14360.75 | 15098.05 | 14862.01 | 15914.67 | 15001.44 | 16061.84 | 14932.11 | 14955.17 | 16180.21 | 14546.5 | 373741.4
44 | 13865.53 | 13724.09 | 13981.53 | 13910.5 | 14237.73 | 12594.43 | 14958.67 | 14170.34 | 13796.28 | 13449.91 | 6035.57 | 6390.03 | 6517.32 | 6518.98 | 6214.34 | 5814.14 | 6513.96 | 6706.5 | 6359.97 | 5941.21 | 5846.9 | 14773.31 | 14407.51 | 15146.17 | 14897.72 | 15957.37 | 15040.39 | 16100.2 | 14967.11 | 14995.06 | 16217.95 | 14592.22 | 374642.94
45 | 13888.06 | 13747.44 | 14003.95 | 13930.45 | 14260.43 | 12614.23 | 14983.72 | 14191.97 | 13820.09 | 13473.92 | 6039.27 | 6394.13 | 6521.02 | 6522.35 | 6215.94 | 5816.58 | 6515.56 | 6707.52 | 6361.67 | 5941.87 | 5848.1 | 14797.48 | 14433.55 | 15172.94 | 14917.83 | 15981.46 | 15062.01 | 16121.49 | 14986.59 | 15017.21 | 16239.2 | 14617.58 | 375145.61
Grand Total | 392774.72 | 391568.16 | 400041.23 | 399327.72 | 406790.88 | 360809.22 | 427752.43 | 409543.92 | 397770.85 | 386307.21 | 261182.92 | 276322.27 | 283825.8 | 284984.58 | 272869.05 | 255411.55 | 286627.42 | 296745.44 | 280820.23 | 264702.32 | 260391.31 | 400732.49 | 391761.45 | 411924.81 | 406542.32 | 437548.05 | 411513.16 | 441554.71 | 412164 | 409854.53 | 443887.9 | 396911.64 | 11560964.29

## 50..800_F2amp
Sum of F4 | Well
Cycles | D02 | D03 | D04 | D05 | D06 | D07 | D08 | D09 | D10 | D11 | D12 | E02 | E03 | E04 | E05 | E06 | E07 | E08 | E09 | E10 | E11 | E12 | F02 | F03 | F04 | F05 | F06 | F07 | F08 | F09 | F10 | F11 | F12 | Grand Total
1 | 5565.5 | 6546.6 | 5654 | 6773.6 | 6376.4 | 6258.6 | 6767.9 | 6902.2 | 6686.9 | 5940.5 | 5839.8 | 6022.3 | 5922.8 | 6250.7 | 6364.7 | 6621 | 5266.4 | 6665.4 | 6439.4 | 6741.8 | 7345.6 | 6477.1 | 5753.6 | 6376.4 | 5885.6 | 6815.8 | 6922.3 | 5866.9 | 6927.3 | 6012.6 | 6521.2 | 6842.7 | 6674.8 | 210028.4
2 | 5588.2 | 6624.8 | 5702.78 | 6841.19 | 6431.75 | 6374.38 | 6818.11 | 6949.39 | 6732.56 | 5947.21 | 5785.45 | 6047.36 | 5972.14 | 6310.44 | 6405.01 | 6703.65 | 5340.52 | 6759.04 | 6522.96 | 6816.21 | 7405.56 | 6275.66 | 5803.81 | 6474.59 | 5929.9 | 6889.75 | 6982.96 | 5979.39 | 7036.34 | 6079.63 | 6589.44 | 6958.23 | 6788.72 | 211867.13
3 | 5594.09 | 6641.53 | 5721.4 | 6843.38 | 6446.03 | 6384.68 | 6816.76 | 6952.26 | 6729.57 | 5946.9 | 5776.85 | 6057.46 | 5983.27 | 6317.28 | 6406.74 | 6715.64 | 5351.49 | 6770.52 | 6528.49 | 6831.37 | 7416.07 | 6269.76 | 5822.48 | 6480.68 | 5944.19 | 6898.75 | 7018.47 | 5986.05 | 7053.21 | 6095.55 | 6597.35 | 6962.17 | 6786.19 | 212146.63
4 | 5597.28 | 6652.2 | 5737.82 | 6846.54 | 6454.04 | 6391.56 | 6816.83 | 6953.89 | 6729.64 | 5945.86 | 5766.15 | 6066.48 | 5992.04 | 6325.44 | 6411.65 | 6728.97 | 5361.98 | 6786.73 | 6534.03 | 6846.04 | 7425.31 | 6264.32 | 5839.28 | 6489.43 | 5961.28 | 6907.91 | 7030.84 | 5991.75 | 7065.17 | 6108.63 | 6608.53 | 6967.03 | 6784.26 | 212388.91
5 | 5601.86 | 6666.26 | 5754.22 | 6849.92 | 6465.94 | 6402.19 | 6815.6 | 6954.67 | 6729.69 | 5945.54 | 5759.12 | 6075.04 | 6004.79 | 6335.61 | 6417.8 | 6744.01 | 5376.19 | 6802.29 | 6540.99 | 6854.67 | 7435.99 | 6258.08 | 5853.57 | 6501.77 | 5977.72 | 6913.7 | 7044.03 | 6002.3 | 7075.16 | 6120.41 | 6617.06 | 6972.04 | 6786.49 | 212654.72
6 | 5607.17 | 6676.23 | 5770.58 | 6852.4 | 6474.69 | 6411.13 | 6813.88 | 6956.6 | 6730.73 | 5944.97 | 5754.12 | 6080.66 | 6014.81 | 6345.73 | 6425.73 | 6760.24 | 5387.53 | 6811.08 | 6547.06 | 6866.75 | 7443.82 | 6254.2 | 5865.93 | 6510.42 | 5993.36 | 6922.11 | 7057.12 | 6011.18 | 7081.31 | 6127.48 | 6623.92 | 6978.2 | 6796.77 | 212897.91
7 | 5612.63 | 6684.2 | 5784.84 | 6854.72 | 6483.75 | 6417.59 | 6813.96 | 6957.61 | 6731.24 | 5944.51 | 5749.33 | 6087.51 | 6022.23 | 6357.54 | 6430.93 | 6777.03 | 5395.58 | 6821.01 | 6552.52 | 6877.96 | 7449.35 | 6249.08 | 5877.7 | 6518.99 | 6007.07 | 6932.03 | 7067.64 | 6020.08 | 7087.98 | 6136.52 | 6631.01 | 6983.26 | 6805.88 | 213123.28
8 | 5616.32 | 6689.28 | 5798.09 | 6857.43 | 6491.81 | 6422.14 | 6812.83 | 6958.88 | 6731.93 | 5943.59 | 5743.04 | 6094.48 | 6030.3 | 6365.51 | 6434.82 | 6793.45 | 5403.18 | 6830.17 | 6558.6 | 6887.16 | 7455.57 | 6243.91 | 5887.91 | 6523.05 | 6020.42 | 6941.09 | 7078.14 | 6025.6 | 7091.96 | 6142.76 | 6636.03 | 6987.3 | 6816.97 | 213313.72
9 | 5618.62 | 6695.97 | 5809.27 | 6859.66 | 6496.2 | 6427.13 | 6809.94 | 6960.6 | 6734.15 | 5943.53 | 5735.11 | 6101.7 | 6035.48 | 6374.87 | 6438.06 | 6808.37 | 5410.67 | 6839.27 | 6559.86 | 6895.71 | 7459.64 | 6236.52 | 5897.69 | 6528.85 | 6032.41 | 6949.13 | 7088.9 | 6032.64 | 7098.72 | 6150.71 | 6640.59 | 6989.47 | 6828.07 | 213487.51
10 | 5621.81 | 6700.92 | 5818.56 | 6861.99 | 6500.5 | 6432.1 | 6808.43 | 6962.47 | 6737.21 | 5942.33 | 5726.96 | 6105.78 | 6042.88 | 6383.99 | 6442.3 | 6823.33 | 5417.67 | 6845.55 | 6561.95 | 6901.18 | 7463.92 | 6226.85 | 5908.34 | 6535.66 | 6039.32 | 6954.26 | 7098.32 | 6037.64 | 7103.96 | 6155.75 | 6642.8 | 6992.49 | 6835 | 213632.22
11 | 5626 | 6705.28 | 5826.69 | 6863.94 | 6504.84 | 6435.54 | 6808.45 | 6963.7 | 6739.87 | 5940.51 | 5720.98 | 6109.01 | 6048.56 | 6392.17 | 6447.89 | 6833.05 | 5421.76 | 6850.08 | 6564.28 | 6904.8 | 7466.74 | 6217.22 | 5915.35 | 6541.71 | 6047.16 | 6959.58 | 7104.11 | 6042.92 | 7108.81 | 6159.92 | 6646.97 | 6996.24 | 6844.09 | 213758.22
12 | 5626.94 | 6708.27 | 5831.98 | 6866.05 | 6507.12 | 6437.98 | 6808.19 | 6963.73 | 6739.43 | 5939.77 | 5716.23 | 6113.25 | 6055.27 | 6399.16 | 6452.26 | 6841.19 | 5427.37 | 6852.54 | 6566.41 | 6906.96 | 7471.81 | 6209.72 | 5922.26 | 6548.69 | 6052.8 | 6965.01 | 7110.23 | 6048.14 | 7113.43 | 6164.47 | 6652.73 | 6996.52 | 6853.35 | 213869.26
13 | 5628.74 | 6709.02 | 5834.16 | 6868.03 | 6509.38 | 6441.61 | 6807.23 | 6964.69 | 6738.51 | 5939.06 | 5709.97 | 6118.04 | 6060.27 | 6403.13 | 6454.29 | 6846.62 | 5431.4 | 6854.86 | 6567.78 | 6908.38 | 7477.94 | 6202.71 | 5926.53 | 6555.23 | 6058.26 | 6970.5 | 7117.92 | 6054.14 | 7113.74 | 6167.64 | 6657.02 | 6996.43 | 6862.26 | 213955.49
14 | 5628.51 | 6710.47 | 5838 | 6868.19 | 6510.25 | 6443.29 | 6804.45 | 6966.15 | 6738.72 | 5937.32 | 5704.45 | 6121.62 | 6065.02 | 6407.4 | 6456.99 | 6850.12 | 5435.19 | 6859.89 | 6569.94 | 6910.68 | 7478.66 | 6195.68 | 5928.44 | 6556.55 | 6063.16 | 6975.05 | 7118.01 | 6058.05 | 7115.7 | 6171.27 | 6660.18 | 6996.65 | 6867.32 | 214011.37
15 | 5628.21 | 6710.58 | 5840.12 | 6866.54 | 6510.18 | 6443.65 | 6801.27 | 6965.39 | 6738.33 | 5934.78 | 5695.71 | 6126.21 | 6070.79 | 6410.89 | 6456.94 | 6850.17 | 5437.44 | 6865.98 | 6570.33 | 6914.64 | 7481.21 | 6190.71 | 5929.68 | 6557.01 | 6065.32 | 6977.55 | 7117.15 | 6062.24 | 7116.29 | 6176.36 | 6659.05 | 6997.19 | 6872.63 | 214040.54
16 | 5628.02 | 6713.41 | 5842.72 | 6865.21 | 6510.42 | 6443.7 | 6796.85 | 6964.17 | 6737.59 | 5933.8 | 5685.04 | 6128.12 | 6072.62 | 6411.08 | 6457.56 | 6853.09 | 5439.13 | 6866.14 | 6570.22 | 6918.2 | 7480.69 | 6187.08 | 5931.72 | 6556.55 | 6065.66 | 6981.75 | 7116.13 | 6062.27 | 7116.4 | 6179.46 | 6658 | 6997.59 | 6876.7 | 214047.09
17 | 5628.43 | 6714.22 | 5842.14 | 6864 | 6510.69 | 6444.58 | 6793.75 | 6964.26 | 6734.82 | 5931.65 | 5674.86 | 6130.69 | 6074.8 | 6411.05 | 6458.99 | 6855.24 | 5440.99 | 6866.05 | 6570.03 | 6920.05 | 7480.64 | 6183.89 | 5933.93 | 6556.5 | 6066.51 | 6985.48 | 7116.1 | 6062.08 | 7115.78 | 6183.63 | 6657.61 | 6997.58 | 6879.34 | 214050.36
18 | 5630.06 | 6714.35 | 5841.03 | 6864.06 | 6510.92 | 6447.13 | 6791.73 | 6963.79 | 6732.37 | 5929.55 | 5666.77 | 6130.8 | 6074.54 | 6411.41 | 6459.33 | 6859.3 | 5442.75 | 6866.52 | 6571.78 | 6920.89 | 7481.48 | 6180.77 | 5937.54 | 6555.94 | 6068.6 | 6988.59 | 7116.18 | 6062.36 | 7115.25 | 6186.53 | 6657.14 | 6997.1 | 6879.03 | 214055.59
19 | 5632.21 | 6717.55 | 5840.37 | 6864.34 | 6514.07 | 6448.71 | 6791.48 | 6963.76 | 6733.93 | 5931.12 | 5666.21 | 6136.04 | 6075.44 | 6416.19 | 6461.92 | 6865.9 | 5446.33 | 6869.75 | 6575.61 | 6923.77 | 7484.76 | 6180.95 | 5943.36 | 6559.56 | 6073.74 | 6995.03 | 7116.57 | 6064.36 | 7116.41 | 6189.06 | 6658.7 | 6999.29 | 6877.64 | 214134.13
20 | 5640.08 | 6728.12 | 5846.67 | 6869.03 | 6522.65 | 6455.34 | 6796.29 | 6969.99 | 6740.44 | 5937.07 | 5667.69 | 6146.12 | 6082.79 | 6424.77 | 6469.76 | 6876.25 | 5453.18 | 6875.42 | 6584.01 | 6930.87 | 7494.98 | 6186.24 | 5951.3 | 6568.96 | 6080.48 | 7004.67 | 7123.22 | 6072 | 7123.02 | 6197.25 | 6666.82 | 7005.9 | 6882.69 | 214374.07
21 | 5654.46 | 6746.7 | 5863.05 | 6883.29 | 6539.84 | 6470.9 | 6808.95 | 6985.5 | 6753.88 | 5948.77 | 5676.34 | 6161.94 | 6099.45 | 6440.95 | 6485.08 | 6895.54 | 5468.27 | 6890.33 | 6602.81 | 6946.76 | 7516.21 | 6202.25 | 5967 | 6587.64 | 6096.46 | 7022.91 | 7141.2 | 6087.41 | 7140.23 | 6212.73 | 6683.68 | 7019.28 | 6896.81 | 214896.62
22 | 5684.61 | 6780.99 | 5894.07 | 6917.05 | 6574.69 | 6505.4 | 6839.98 | 7019.68 | 6784.57 | 5974.17 | 5698.78 | 6192.58 | 6132.18 | 6471.85 | 6517.34 | 6934.59 | 5497.62 | 6923 | 6636.87 | 6980.99 | 7557.68 | 6238.99 | 5997.97 | 6622.87 | 6131.57 | 7060.08 | 7174.73 | 6117.24 | 7176.34 | 6242.91 | 6714.78 | 7052.12 | 6925.39 | 215973.68
23 | 5744.26 | 6849.24 | 5954.87 | 6986 | 6644.97 | 6573.24 | 6905.9 | 7086.31 | 6846.98 | 6029.26 | 5751.08 | 6254.93 | 6195.68 | 6533.84 | 6583.82 | 7008.01 | 5552.04 | 6991.17 | 6703.15 | 7055.2 | 7633.46 | 6308.36 | 6057.82 | 6689.81 | 6197.52 | 7133.33 | 7243.79 | 6176.87 | 7248.72 | 6302.63 | 6779.36 | 7118.43 | 6986.75 | 218126.8
24 | 5855.64 | 6980.91 | 6071.94 | 7118.31 | 6770.02 | 6700.16 | 7030.86 | 7214.99 | 6966.7 | 6131.73 | 5855.37 | 6370.11 | 6316.44 | 6656.35 | 6708.55 | 7144.28 | 5653.41 | 7124.78 | 6827.76 | 7198.48 | 7776.84 | 6443.73 | 6170.78 | 6813.25 | 6317.75 | 7267.29 | 7377.97 | 6287.9 | 7381.86 | 6418.58 | 6907.83 | 7241.78 | 7112.86 | 222215.21
25 | 6053.11 | 7222.51 | 6281.11 | 7370.73 | 6999.47 | 6934.37 | 7269.78 | 7452.09 | 7191.39 | 6325.51 | 6051.16 | 6584.98 | 6530.48 | 6879.3 | 6936.6 | 7405.62 | 5842.15 | 7364.58 | 7058.12 | 7450.25 | 8042.86 | 6686.21 | 6377.91 | 7041.09 | 6538.98 | 7514.93 | 7628.14 | 6495.53 | 7625.57 | 6639.73 | 7144.93 | 7471.65 | 7355.56 | 229766.4
26 | 6396.82 | 7640.51 | 6645.98 | 7795.94 | 7403.78 | 7337.06 | 7688.74 | 7864.98 | 7587.48 | 6667.61 | 6395.63 | 6961.11 | 6900.3 | 7267.16 | 7333.77 | 7845.46 | 6182.82 | 7777.95 | 7464.03 | 7883.76 | 8504.45 | 7100.41 | 6750.63 | 7438.9 | 6922.07 | 7946.03 | 8066.2 | 6860.69 | 8051.7 | 7012.87 | 7554.69 | 7879.34 | 7796.32 | 242925.19
27 | 6942.78 | 8318.52 | 7221.86 | 8466.85 | 8047.92 | 7992.06 | 8332.81 | 8545.32 | 8243.22 | 7218.64 | 6959.17 | 7559.84 | 7486.33 | 7898.6 | 7978.95 | 8532.45 | 6711.83 | 8436.68 | 8111.21 | 8573 | 9234.03 | 7742.76 | 7330.3 | 8069.77 | 7523.26 | 8647.44 | 8761.95 | 7441.64 | 8732.53 | 7621.85 | 8228.46 | 8548.66 | 8470.9 | 263931.59
28 | 7711.28 | 9251.43 | 8028.26 | 9410.2 | 8958.85 | 8893.49 | 9239.23 | 9484.75 | 9159.17 | 8005.98 | 7729.53 | 8407.2 | 8316.67 | 8766.46 | 8867.92 | 9497.04 | 7459.01 | 9372.86 | 9022.94 | 9547.43 | 10256.15 | 8628.45 | 8146.5 | 8966.71 | 8365.41 | 9614.27 | 9745.55 | 8270.89 | 9703.99 | 8468.3 | 9158.58 | 9479.38 | 9406.16 | 293340.04
29 | 8645.84 | 10402.47 | 9020.64 | 10584.52 | 10065.27 | 10001.25 | 10373.43 | 10648.52 | 10283.48 | 9022.17 | 8676.28 | 9484.94 | 9330.15 | 9860.56 | 9942.72 | 10665.01 | 8374.05 | 10546.04 | 10156.28 | 10731.74 | 11497.64 | 9713.02 | 9149.79 | 10068.69 | 9384.4 | 10784.96 | 10977.26 | 9288.86 | 10926.02 | 9501.44 | 10282.85 | 10618.5 | 10566.27 | 329575.06
30 | 9711.42 | 11685.03 | 10127.9 | 11900.29 | 11327.3 | 11240.89 | 11647.67 | 11997.07 | 11612.27 | 10179.26 | 9745.1 | 10678.78 | 10489.36 | 11066.19 | 11197.54 | 12023.79 | 9421.45 | 11861.01 | 11436.3 | 12083.84 | 12909.4 | 10935.77 | 10327.59 | 11332.23 | 10544.56 | 12125.83 | 12390.97 | 10479.97 | 12328.98 | 10690.55 | 11573.54 | 11957.61 | 11874.24 | 370903.7
31 | 10823.03 | 13009.87 | 11295.16 | 13281.89 | 12644.62 | 12526.32 | 12966.52 | 13379.67 | 12982.76 | 11357.76 | 10858.62 | 11914.6 | 11717.6 | 12308.46 | 12484.18 | 13412.82 | 10513.37 | 13217.28 | 12766.65 | 13513.85 | 14372.82 | 12161.31 | 11539.52 | 12676.22 | 11748.45 | 13526.77 | 13849.48 | 11699.38 | 13755.86 | 11958.01 | 12917.32 | 13335.76 | 13224.67 | 413740.6
32 | 11902.87 | 14301.07 | 12408.68 | 14620.32 | 13941.83 | 13777.22 | 14266.64 | 14707.79 | 14305.61 | 12513.46 | 11938.93 | 13115.12 | 12896.8 | 13516.03 | 13722.48 | 14743.69 | 11578.39 | 14532.15 | 14058.37 | 14876.11 | 15809.63 | 13350.41 | 12733.6 | 13994.28 | 12911.29 | 14886.78 | 15233.23 | 12879.38 | 15155 | 13174.04 | 14219.73 | 14667.78 | 14529.32 | 455268.03
33 | 12890.36 | 15505 | 13423.12 | 15838.67 | 15108.45 | 14927.81 | 15455.16 | 15937.8 | 15527.81 | 13575.59 | 12931.34 | 14214.32 | 13974.58 | 14625.45 | 14867 | 15973.95 | 12570.36 | 15755.23 | 15260.84 | 16106.76 | 17121.38 | 14428.15 | 13820.87 | 15186.56 | 13978.39 | 16133.47 | 16490.69 | 13968.14 | 16424.07 | 14279.47 | 15412.88 | 15891.84 | 15727.22 | 493332.73
34 | 13748.36 | 16554.93 | 14311.53 | 16875.69 | 16114.61 | 15941.36 | 16480.7 | 16992.01 | 16575.64 | 14487.1 | 13796.03 | 15184.74 | 14911.65 | 15593.34 | 15849.03 | 17029.39 | 13424.14 | 16820.55 | 16291.67 | 17184.76 | 18235.27 | 15362.78 | 14766.78 | 16222.64 | 14920.06 | 17221.25 | 17583.77 | 14920.98 | 17523.5 | 15241.71 | 16460.62 | 16961.18 | 16769.93 | 526357.7
35 | 14469.74 | 17414.86 | 15048.86 | 17728.34 | 16949.41 | 16782.56 | 17316.02 | 17857.93 | 17438.87 | 15240.24 | 14529.23 | 15992.29 | 15695.93 | 16404.73 | 16660.19 | 17909.07 | 14133.74 | 17690.83 | 17141.68 | 18071.71 | 19147.33 | 16141.98 | 15554.52 | 17093.56 | 15695.54 | 18127.93 | 18499.09 | 15718.77 | 18441.27 | 16045.56 | 17333.21 | 17854.05 | 17632.01 | 553761.05
36 | 15042.1 | 18098.89 | 15628.11 | 18410.9 | 17613.22 | 17442.86 | 17979.01 | 18544.51 | 18127.07 | 15838.53 | 15106.13 | 16618.91 | 16328.06 | 17055.7 | 17299.74 | 18411.35 | 14696.27 | 18371 | 17813.4 | 18766.98 | 19859.47 | 16769.6 | 16195.83 | 17786.54 | 16304.82 | 18843.14 | 19234.62 | 16361.01 | 19178.58 | 16690.9 | 18010.03 | 18568.96 | 18321.25 | 575317.49
37 | 15362.86 | 18485.75 | 15970.91 | 18946.98 | 17985.75 | 17813.7 | 18355.43 | 19076.96 | 18515.82 | 16303.09 | 15445.71 | 16987.12 | 16821.39 | 17421.61 | 17677.97 | 18963.82 | 15030.2 | 18772.27 | 18210.97 | 19176.32 | 20278.51 | 17125.56 | 16558.76 | 18179.93 | 16664.48 | 19265.93 | 19798.59 | 16725.03 | 19595.08 | 17188.64 | 18405.61 | 18971.15 | 18711.99 | 588793.89
38 | 15716.3 | 18893.79 | 16335.37 | 19249.02 | 18391.68 | 18200.9 | 18752.69 | 19375.8 | 18936.73 | 16564.59 | 15801.45 | 17381.31 | 17098.76 | 17816.12 | 18083.18 | 19287.22 | 15371.79 | 19197.59 | 18634.13 | 19610.87 | 20714.55 | 17514.45 | 16950.41 | 18593.83 | 17044.14 | 19714.96 | 20128.24 | 17114.2 | 20047.12 | 17480.95 | 18836.36 | 19408.83 | 19137.86 | 601385.19
39 | 15922.16 | 19221.14 | 16622.72 | 19570.01 | 18628.05 | 18511.21 | 18984.73 | 19687.28 | 19184.16 | 16836.69 | 16008.94 | 17608.29 | 17391.29 | 18049.17 | 18319.64 | 19637.93 | 15646.86 | 19445.54 | 18880.44 | 19861.3 | 20964.6 | 17743.99 | 17179.42 | 18929.18 | 17351.13 | 19978.28 | 20376.17 | 17341.12 | 20312.54 | 17793.87 | 19087.65 | 19663.08 | 19385.86 | 610124.44
40 | 16133.08 | 19405.79 | 16792.96 | 19753.47 | 18874.6 | 18685.28 | 19157.21 | 19867.43 | 19366.89 | 16993.33 | 16231.38 | 17853.11 | 17632.8 | 18297.77 | 18572.8 | 19840.1 | 15803.28 | 19713.26 | 19139.45 | 20129.55 | 21153.74 | 17992.89 | 17422.39 | 19119.07 | 17523.42 | 20258.83 | 20656.98 | 17591.3 | 20592.35 | 17975.05 | 19346.26 | 19935.74 | 19570.23 | 617381.79
41 | 16254.65 | 19597.39 | 16960.93 | 19950.23 | 19016.58 | 18818.36 | 19348.64 | 19996.02 | 19566.43 | 17109.69 | 16410.76 | 17995.07 | 17771.62 | 18441.46 | 18719.49 | 19989.99 | 15966.55 | 19868.7 | 19291.76 | 20283.4 | 21303.33 | 18139.17 | 17564.22 | 19325.1 | 17652.02 | 20423.35 | 20822.6 | 17735.56 | 20754.44 | 18107.98 | 19495.27 | 20093.08 | 19778.48 | 622552.32
42 | 16376.33 | 19707.81 | 17057.35 | 20065.57 | 19122.24 | 18962.41 | 19458.89 | 20098.58 | 19682.1 | 17232.02 | 16513.95 | 18136.44 | 17903.56 | 18592.69 | 18825.74 | 20159.54 | 16060.96 | 20033.81 | 19447.31 | 20447.03 | 21422.94 | 18298.03 | 17705.74 | 19443.59 | 17796.86 | 20591.23 | 20995.26 | 17842.91 | 20870.74 | 18255.04 | 19645.91 | 20206.09 | 19950.72 | 626909.39
43 | 16428.2 | 19764.83 | 17108.03 | 20125.21 | 19181.58 | 19023.6 | 19515.6 | 20154.29 | 19740.52 | 17284.55 | 16566.68 | 18196.24 | 17959.49 | 18656.65 | 18885.74 | 20231.65 | 16109.65 | 20104.13 | 19513.82 | 20516.11 | 21490.56 | 18367.29 | 17765.93 | 19503.39 | 17858.82 | 20661.59 | 21069.06 | 17902.55 | 20932.04 | 18317.9 | 19709.59 | 20265.7 | 20022.72 | 628933.71
44 | 16497.49 | 19845.21 | 17179.67 | 20209.26 | 19266.35 | 19105.17 | 19595.34 | 20233.37 | 19822.42 | 17354.87 | 16640.93 | 18276.2 | 18034.27 | 18741.92 | 18971.45 | 20327.78 | 16178.3 | 20198.24 | 19602.87 | 20608.4 | 21587.25 | 18459.98 | 17846.43 | 19587.37 | 17941.8 | 20755.28 | 21167.89 | 17987.53 | 21018.6 | 18402.03 | 19794.64 | 20350.01 | 20118.77 | 631707.09
45 | 16535.97 | 19890.49 | 17220.05 | 20256.6 | 19314.27 | 19150.44 | 19640.22 | 20277.99 | 19868.49 | 17393.94 | 16682.75 | 18320.62 | 18075.81 | 18789.26 | 19019.89 | 20381.14 | 16216.98 | 20250.53 | 19652.35 | 20659.65 | 21641.91 | 18511.47 | 17891.14 | 19634.63 | 17987.91 | 20807.26 | 21222.81 | 18035.53 | 21067.39 | 18448.77 | 19841.86 | 20397.54 | 20172.1 | 633257.76
Grand Total | 404824.4 | 485284.39 | 420408.57 | 495415.56 | 471216.98 | 466681.05 | 485504.08 | 499644.84 | 485718.09 | 426388.12 | 408545.11 | 444449.46 | 438711.49 | 459605.98 | 465264.48 | 496257.85 | 393550.24 | 493817.8 | 476781.43 | 501942.34 | 535131.75 | 451057.46 | 431602.27 | 475403.39 | 437834.07 | 505240.86 | 515180.58 | 437872.48 | 513760.46 | 447377.14 | 482227.79 | 500570.92 | 493042.59 | 15446314.02

## IAC_F2_matrix
Creating a primer matrix from concentrations
| USER_INPUTS | Value | Note
| orig_F_conc | 10 | What is the starting F primer concentration? (in uM) | [5 empty] | orig_F_conc = 10, # What is the starting F primer concentration? (in uM)
| orig_R_conc | 10 | What is the starting R primer concentration? (in uM) | [5 empty] | orig_R_conc = 10, # What is the starting R primer concentration? (in uM)
| orig_P_conc | 10 | What is probe starting concentration? (in uM) | [5 empty] | orig_P_conc = 10, # What is probe starting concentration? (in uM)
| std_F_conc | 300 | What is the constant F primer concentration for standards? This should be guess or from literature. (in nM) | [5 empty] | std_F_conc = 300, # What is the constant F primer concentration for standards? This should be guess or from literature. (in nM)
| std_R_cond | 300 | What is the constant R primer concentration for standards? This should be guess or from literature. (in nM) | [5 empty] | std_R_cond = 300, # What is the constant R primer concentration for standards? This should be guess or from literature. (in nM)
| P_conc | 300 | What is the probe concentration? This should be constant throughout experiment. | [5 empty] | P_conc = 300, # What is the probe concentration? This should be constant throughout experiment.
| tot_stds | 21 | How many wells of standards will be run? (in # wells) | [5 empty] | tot_stds = 21, # How many wells of standards will be run? (in # wells)
| tot_NTCs | 3 | How many wells of NTCs will be run? (in # wells) | [5 empty] | tot_NTCs = 3, # How many wells of NTCs will be run? (in # wells)
| tot_samp | 72 | How many samples with varying conditions will be run? (in # wells) | [5 empty] | tot_samp = 72, # How many samples with varying conditions will be run? (in # wells)
| rxn_base | 8.86 | Everything in PCR buffer (Mg2+, dNTPs, polymerase, enhancers, stabilizers, sugars, etc.). (in uL) | [5 empty] | rxn_base = 8.86, # Everything in PCR buffer (Mg2+, dNTPs, polymerase, enhancers, stabilizers, sugars, etc.). (in uL)
| tot_rxn_vol | 20 | What is total volume of PCR reaction in plate/tube? (in ul) | [5 empty] | tot_rxn_vol = 20, # What is total volume of PCR reaction in plate/tube? (in ul)
| F_concs | [50, 100, 200, 400, 600, 800] | What are the F primer concentrations? (a list in nM) | [5 empty] | F_concs = [50, 100, 200, 400, 600, 800], # What are the F primer concentrations? (a list in nM)
| R_concs | [50, 100, 200, 400, 600, 800] | What are the R primer concentrations? (a list in nM) | [5 empty] | R_concs = [50, 100, 200, 400, 600, 800], # What are the R primer concentrations? (a list in nM)
| F_50 | 50 | F primer at lowest concentration. (in nM) | [5 empty] | F_50 = 50, # F primer at lowest concentration. (in nM)
| F_100 | 100 | F primer at concentration. (in nM) | [5 empty] | F_100 = 100, # F primer at concentration. (in nM)
| F_200 | 200 | F primer at concentration. (in nM) | [5 empty] | F_200 = 200, # F primer at concentration. (in nM)
| F_400 | 400 | F primer at concentration. (in nM) | [5 empty] | F_400 = 400, # F primer at concentration. (in nM)
| F_600 | 600 | F primer at concentration. (in nM) | [5 empty] | F_600 = 600, # F primer at concentration. (in nM)
| F_800 | 800 | F primer at highest concentration. (in nM) | [5 empty] | F_800 = 800, # F primer at highest concentration. (in nM)
| R_50 | 50 | R primer at lowest concentration. (in nM) | [5 empty] | R_50 = 50, # R primer at lowest concentration. (in nM)
| R_100 | 100 | R primer at concentration. (in nM) | [5 empty] | R_100 = 100, # R primer at concentration. (in nM)
| R_200 | 200 | R primer at concentration. (in nM) | [5 empty] | R_200 = 200, # R primer at concentration. (in nM)
| R_400 | 400 | R primer at concentration. (in nM) | [5 empty] | R_400 = 400, # R primer at concentration. (in nM)
| R_600 | 600 | R primer at concentration. (in nM) | [5 empty] | R_600 = 600, # R primer at concentration. (in nM)
| R_800 | 800 | R primer at highest concentration. (in nM) | [5 empty] | R_800 = 800, # R primer at highest concentration. (in nM)
| dna_per_rxn | 2 | How much standard, positive control or NTC to add per well. | [5 empty] | dna_per_rxn = 2, # How much standard, positive control or NTC to add per well.
| std_NTC_reps | 3 | How many standard and NTC replicates? (int) | [5 empty] | std_NTC_reps = 3, # How many standard and NTC replicates? (int)
| R_reps | 12 | How many wells will use R primer at particular concentration? (int) | [5 empty] | R_reps = 12, # How many wells will use R primer at particular concentration? (int)
| F_reps | 12 | How many wells will use F primer at particular concentration? (int) | [5 empty] | F_reps = 12, # How many wells will use F primer at particular concentration? (int)
| F_int_vol | 100 | What is the volume of F intermediate primer in new tube? (in ul) | [5 empty] | F_int_vol = 100, # What is the volume of F intermediate primer in new tube? (in ul)
| percent_waste | 0.2 | What percentage waste? (decimal) | [5 empty] | percent_waste = # What percentage waste? (decimal)
| sN_mix_waste_offset | 0.025 | How much percent_waste offset should sN_mix use? This calculated as percent_waste-sN_mix_overage = percent_waste for sN_mix_overage e.g. (20-7=13%) Should not be 0 otherwise offset = percent_waste. (decimal) | [5 empty] | sN_mix_waste_offset = # How much percent_waste offset should sN_mix use? This calculated as percent_waste-sN_mix_overage = percent_waste for sN_mix_overage e.g. (20-7=13%) Should not be 0 otherwise offset = percent_waste. (decimal)
| R_mix_waste_offset | 0.1 | How much percent_waste offset should R_mix use? This calculated as percent_waste-R_mix_overage = percent_waste for R_mix_overage e.g. (20-=13%) If 0, then offset = percent_waste. (decimal) | [5 empty] | R_mix_waste_offset = # How much percent_waste offset should R_mix use? This calculated as percent_waste-R_mix_overage = percent_waste for R_mix_overage e.g. (20-=13%) If 0, then offset = percent_waste. (decimal)
| std_NTC_waste_offset | 0.028 | How much percent_waste offset should std_NTC use? (decimal) | [5 empty] | std_NTC_waste_offset = # How much percent_waste offset should std_NTC use? (decimal)
| bpw_waste_offset | 0.032 | How much percent_waste offset should bpw_waste use? (decimal) | [5 empty] | bpw_waste_offset = # How much percent_waste offset should bpw_waste use? (decimal)
CALCS
| tot_rxns | 96 | Calc what is total # rxns. (int) | [5 empty] | tot_rxns = # Calc what is total # rxns. (int)
| tot_sds_NTC | 24 | Calc number of standards and nontemplate controls. (int) | [5 empty] | tot_sds_NTC = # Calc number of standards and nontemplate controls. (int)
| rxn_vol_no_dna | 18 | Calc what is volume of PCR rxn with no DNA added. (in ul) | [5 empty] | rxn_vol_no_dna = # Calc what is volume of PCR rxn with no DNA added. (in ul)
| P_per_rxn | 0.6 | Calc this constant probe is needed per pcr reaction. (in ul) | [5 empty] | P_per_rxn = # Calc this constant probe is needed per pcr reaction. (in ul)
| max_vol_F_per_rxn | 1.6 | Calc greatest vol added from starting F primer conc to rxn to satisfy max test F conc? (in ul) | [5 empty] | max_vol_F_per_rxn = # Calc greatest vol added from starting F primer conc to rxn to satisfy max test F conc? (in ul)
| max_vol_R_per_rxn | 1.6 | Calc greatest vol added from starting R primer conc to rxn to satisfy max test R conc? (in ul) | [5 empty] | max_vol_R_per_rxn = # Calc greatest vol added from starting R primer conc to rxn to satisfy max test R conc? (in ul)
| water_per_rxn | 5.34 | Calc this is water added per rxn such that greatest F, R primer conc can be accommodated in total rxn vol. (in ul) | [5 empty] | water_per_rxn = # Calc this is water added per rxn such that greatest F, R primer conc can be accommodated in total rxn vol. (in ul)
| std_vol_F_per_rxn | 0.6 | Calc adding this much F primer to each std rxn | [5 empty] | std_vol_F_per_rxn = # Calc adding this much F primer to each std rxn
| std_vol_R_per_rxn | 0.6 | Calc adding this much F primer to each std rxn | [5 empty] | std_vol_R_per_rxn = # Calc adding this much F primer to each std rxn
| BPW_rxn | 14.8 | Mix base + probe + water (no DNA, F, R primer). (in ul) | [5 empty] | BPW_rxn = # Mix base + probe + water (no DNA, F, R primer). (in ul)
| std_woff_per_sN_rxn | 2 | Calc how much water to add to offset diff in max vol F,R primer per reaction. (Max-Std)*2 (in ul)
| BPW_sN_rxn | 18 | Mix base + probe + water+F,R primer at std conc+water_offset(no DNA). (in ul) | [5 empty] | BPW_sN_rxn = # Mix base + probe + water+F,R primer at std conc+water_offset(no DNA). (in ul)
| BWP_mix_xfer_sN_mix | 417.36 | Calc how much tot_BPW_mix to move to make new tube for stds and NTC. (in ul) | [5 empty] | BWP_mix_xfer_sN_mix = # Calc how much tot_BPW_mix to move to make new tube for stds and NTC. (in ul)
| std_woff_add_to_sN_mix | 56.4 | Cal how much water to add to sdd_NTC mmix with waste | [5 empty] | std_woff_add_to_sN_mix = # Cal how much water to add to sdd_NTC mmix with waste
| F_add_to_sN_mix | 16.92 | Calc how much F primer to add to std_NTC mmix with waste | [5 empty] | F_add_to_sN_mix = # Calc how much F primer to add to std_NTC mmix with waste
| R_add_to_sN_mix | 16.92 | Calc how much R primer to add to std_NTC mmix with waste | [5 empty] | R_add_to_sN_mix = # Calc how much R primer to add to std_NTC mmix with waste
| sN_mix_xfer_to_stds_mix | 63.288 | Calc how much sn_mix to aliquot to tube in preparation for std DNA mixing. (in ul) | [5 empty] | sN_mix_xfer_to_stds_mix = # Calc how much sn_mix to aliquot to tube in preparation for std DNA mixing. (in ul)
| std_DNA_xfer_to_stds_mix | 7.032 | Calc how much DNA from each standard adding to stds_mix with waste. (in ul) | [5 empty] | std_DNA_xfer_to_stds_mix = # Calc how much DNA from each standard adding to stds_mix with waste. (in ul)
| bpw_mix_xfer_bpwd_mix | 1244.6208 | Calc how much bpw_mix needed in new tube for # total samples | [5 empty] | bpw_mix_xfer_bpwd_mix = # Calc how much bpw_mix needed in new tube for # total samples
| std_DNA_xfer_to_bpwd_mix | 168.192 | Calc how much DNA to add to bpw_mix for template in PCR rxns. | [5 empty] | std_DNA_xfer_to_bpwd_mix = # Calc how much DNA to add to bpw_mix for template in PCR rxns.
| dna_per_rxn_10x | 0.2 | 10x concentrated DNA to add to reactions to avoid using too much std sample | [5 empty] | dna_per_rxn_10x = # 10x concentrated DNA to add to reactions to avoid using too much std sample
| bpwd_rxn | 16.8 | Calc vol of bpw reaction + DNA vol | [5 empty] | bpwd_rxn = # Calc vol of bpw reaction + DNA vol
| bpwd_mix_xfer_R_mix | 221.76 | Calc how muc bpwd_mix to transfer to each tube for specific R conc. Less than 12% to avoid missing vol. | [5 empty] | bpwd_mix_xfer_R_mix = # Calc how muc bpwd_mix to transfer to each tube for specific R conc. Less than 12% to avoid missing vol.
| R_50_rxn | 0.1 | Calc how much R primer at lowest concentration to add per reaction? (e.g. 50nM) (in ul) | [5 empty] | R_50_rxn = # Calc how much R primer at lowest concentration to add per reaction? (e.g. 50nM) (in ul)
| R_50_mix | 1.32 | Calc how much R primer to add to R_50 mix. (in ul) | [5 empty] | R_50_mix = # Calc how much R primer to add to R_50 mix. (in ul)
| R_50_woff_rxn | 1.5 | Calc how much water offset to add to each rxn. (in ul) | [5 empty] | R_50_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_50_woff_mix | 19.8 | Calc how much water offset to add to R_50 mix. (in ul) | [5 empty] | R_50_woff_mix = # Calc how much water offset to add to R_50 mix. (in ul)
| R_100_rxn | 0.2 | Calc how much R primer at concentration to add. (in ul) | [5 empty] | R_100_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_100_mix | 2.64 | Calc how much R primer to add to add. (in ul) | [5 empty] | R_100_mix = # Calc how much R primer to add to add. (in ul)
| R_100_woff_rxn | 1.4 | Calc how much water offset to add to each rxn. (in ul) | [5 empty] | R_100_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_100_woff_mix | 18.48 | Calc how much water offset to add to mix. (in ul) | [5 empty] | R_100_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_200_rxn | 0.4 | Calc how much R primer at concentration to add. (in ul) | [5 empty] | R_200_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_200_mix | 5.28 | Calc how much R primer to add to add. (in ul) | [5 empty] | R_200_mix = # Calc how much R primer to add to add. (in ul)
| R_200_woff_rxn | 1.2 | Calc how much water offset to add to each rxn. (in ul) | [5 empty] | R_200_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_200_woff_mix | 15.84 | Calc how much water offset to add to mix. (in ul) | [5 empty] | R_200_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_400_rxn | 0.8 | Calc how much R primer at concentration to add. (in ul) | [5 empty] | R_400_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_400_mix | 10.56 | Calc how much R primer to add to add. (in ul) | [5 empty] | R_400_mix = # Calc how much R primer to add to add. (in ul)
| R_400_woff_rxn | 0.8 | Calc how much water offset to add to each rxn. (in ul) | [5 empty] | R_400_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_400_woff_mix | 10.56 | Calc how much water offset to add to mix. (in ul) | [5 empty] | R_400_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_600_rxn | 1.2 | Calc how much R primer at concentration to add. (in ul) | [5 empty] | R_600_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_600_mix | 15.84 | Calc how much R primer to add to add. (in ul) | [5 empty] | R_600_mix = # Calc how much R primer to add to add. (in ul)
| R_600_woff_rxn | 0.4 | Calc how much water offset to add to each rxn. (in ul) | [5 empty] | R_600_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_600_woff_mix | 5.28 | Calc how much water offset to add to mix. (in ul) | [5 empty] | R_600_woff_mix = # Calc how much water offset to add to mix. (in ul)
| R_800_rxn | 1.6 | Calc how much R primer at concentration to add. (in ul) | [5 empty] | R_800_rxn = # Calc how much R primer at concentration to add. (in ul)
| R_800_mix | 21.12 | Calc how much R primer to add to add. (in ul) | [5 empty] | R_800_mix = # Calc how much R primer to add to add. (in ul)
| R_800_woff_rxn | 0 | Calc how much water offset to add to each rxn. (in ul) | [5 empty] | R_800_woff_rxn = # Calc how much water offset to add to each rxn. (in ul)
| R_800_woff_mix | 0 | Calc how much water offset to add to mix. (in ul) | [5 empty] | R_800_woff_mix = # Calc how much water offset to add to mix. (in ul)
| F_50_int_conc | 0.625 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | F_50_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_100_int_conc | 1.25 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | F_100_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_200_int_conc | 2.5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | F_200_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_400_int_conc | 5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | F_400_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_600_int_conc | 7.5 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | F_600_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_800_int_conc | 10 | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM) | [5 empty] | F_800_int_conc = # What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc in rxn well? (in uM)
| F_50_int_primer | 6.25 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | F_50_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_50_int_water | 93.75 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | F_50_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_100_int_primer | 12.5 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | F_100_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_100_int_water | 87.5 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | F_100_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_200_int_primer | 25 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | F_200_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_200_int_water | 75 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | F_200_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_400_int_primer | 50 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | F_400_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_400_int_water | 50 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | F_400_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_600_int_primer | 75 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | F_600_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_600_int_water | 25 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | F_600_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| F_800_int_primer | 100 | What amount of primer should be added to generate int F primer conc? (in ul) | [5 empty] | F_800_int_primer = # What amount of primer should be added to generate int F primer conc? (in ul)
| F_800_int_water | 0 | What amount of water should be added to generate int F primer conc? (in ul) | [5 empty] | F_800_int_water = # What amount of water should be added to generate int F primer conc? (in ul)
| R_mix_rxn | 18.4 | What is the rxn vol that should be aliquoted to plate wells after R primer addition? (in ul) | [5 empty] | R_mix_rxn = # What is the rxn vol that should be aliquoted to plate wells after R primer addition? (in ul)
MIXES
| BPW_mix_tot | 1704.96 | Mix = base + probe + water (no DNA, F, R primer)*96*waste. (in ul) | [5 empty] | BPW_mix_tot = # Mix = base + probe + water (no DNA, F, R primer)*96*waste. (in ul)
| sN_mix_tot | 507.6 | Mix = base + probe + water + F,R primers at std conc + water offset (no DNA) * number stds_NTC * waste | [5 empty] | sN_mix_tot = # Mix = base + probe + water + F,R primers at std conc + water offset (no DNA) * number stds_NTC * waste
| bpwd_mix_tot | 1412.8128 | Mix = base + probe + DNA (no F, R primer) | [5 empty] | bpwd_mix_tot = # Mix = base + probe + DNA (no F, R primer)
| 50_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) | [5 empty] | 50_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 100_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) | [5 empty] | 100_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 200_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) | [5 empty] | 200_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 400_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) | [5 empty] | 400_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 600_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) | [5 empty] | 600_mix = # Mix = base + probe + DNA+ R primer (No F primer)
| 800_mix | 242.88 | Mix = base + probe + DNA+ R primer (No F primer) | [5 empty] | 800_mix = # Mix = base + probe + DNA+ R primer (No F primer)
Mastermix ordered from Raquel
| Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 800 |  |  | Confirming Volumes and offsets
| Stock Concentrations |  |  | Overage = |  |  | 5 | % |  | BPW_mix | 1704.96 |  | bpwd_mix_tot | 1412.8128 |  | F_10uM use
| dNTP Mix | 25 | mM | [6 empty] | aliquoted (transferred (xfer)) to two mixes |  |  | bpwd_mix_xfer_R_mix | 221.76
| Mg2+ | 1000 | mM | [6 empty] | BWP_mix_xfer_sN_mix | 417.36 |  | # transfers | 6
| AptaTaq | 50 | Units/uL | [6 empty] | bpwd_mix | 1244.6208 |  |  | 1330.56
| Trehalose | 50 | Percent | [7 empty] | 1661.9808
| BSA | 50 | mg/mL
| 10X Taq Buffer (No Mg2+) | 10 | x | [6 empty] | sN_mix has other volumes added and then gets aliquoted to 8 tubes containing standards
| 2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent | [6 empty] | sN_mix_tot | 507.6
[10 empty] | sN_mix_xfer_to_stds_mix | 63.288
[10 empty] | * 8 tubes | 506.304
| Component | Desired |  | Per 20 uL |  | Multiple Reactions
| dNTP Mix | 0.25 | mM | 0.32 | uL | 0.32 | uL
| AptaTaq | 2.5 | Units | 0.05 | uL | 0.05 | uL |  | What intermediate (int) concentration of F primer is needed such that by adding 1.6ul I obtain a final F primer concentration of 50, 100, 200…etc
| Mg2+ | 3.5 | mM | 0.07 | uL | 0.07 | uL |  | uM | ul | nM | ul
| Trehalose | 10 | Percent | 4 | uL | 4 | uL |  | C1 | V1 | C2 | V2
| BSA | 0.5 | mg/mL | 0.2 | uL | 0.2 | uL |  | 0.625 | 1.6 | 50 | 20
| PCR Buffer 10X | 1 | x | 2 | uL | 2 | uL |  | 1.25 | 1.6 | 100 | 20
| 2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 2.2222 | uL |  | 2.5 | 1.6 | 200 | 20
| Water |  |  | 0 | uL | 0 | uL |  | 5 | 1.6 | 400 | 20
|  |  | Total: | 8.8622 | uL | 8.8622 | uL |  | 7.5 | 1.6 | 600 | 20
[9 empty] | 10 | 1.6 | 800 | 20
| tube:"bpw_mix" | 24 wells for std curve and NTCs
| Component | Desired | Units | PROBE RXN | 115.2 |  |  |  | Creating F primer intermediate concentrations from 10uM in 1.5mL tube
| mmix |  |  | 8.8622 | 1020.928 |  |  |  | uM | ul | uM | ul
| Leg pneumo Probe | 500nM |  | 1 | 115.2 |  |  |  | C1 | V1 | C2 | V2 |  | 10uM N2_Fwd | Water
| Leg pneumo ultramer std_5 |  |  | 0.2 | 23.04 |  |  |  | 10 | 6.25 | 0.625 | 100 |  | 6.25 | 93.75
| Leg pneumo F | 400nM |  | 0.8 | 92.16 |  |  |  | 10 | 12.5 | 1.25 | 100 |  | 12.5 | 87.5
| Leg pneumo R | 400nM |  | 0.8 | 92.16 |  |  |  | 10 | 25 | 2.5 | 100 |  | 25 | 75
| IAC Probe | 300nM |  | 0.6 | 69.12 |  |  |  | 10 | 50 | 5 | 100 |  | 50 | 50
[5 empty] | 0 |  |  |  | 10 | 75 | 7.5 | 100 |  | 75 | 25
| water |  |  | 2.5378 | 292.352 |  |  |  | 10 | 100 | 10 | 100 |  | 100 | 0
| TOTAL |  |  | 14.8 | 1704.96
| tube:"sN_mix" | 24 wells for std curve and NTCs
| Component | Desired | Units | PROBE RXN | 27.6
| mmix |  |  | 14.8 | 408.48
| IAC_F2 primer@10uM | 300nM |  | 0.6 | 16.56
| IAC R_primer@10uM | 300nM |  | 0.6 | 16.56
| water |  |  | 2 | 55.2
| TOTAL |  |  | 18 | 496.8
| tube:"bpwd_mix" | 72 wells for samples
| Component | Desired | Units | PROBE RXN | 84.096
| mmix |  |  | 14.8 | 1244.6208
| F_primer@10uM | various |  | 0 | 0
| R_primer@10uM | various |  | 0 | 0
| water |  |  | 0 | 0
| DNA (from std_7) for |  |  | 2 | 168.192
| TOTAL |  |  | 16.8 | 1412.8128
tube:"50_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 50 | nm | 0.1 | 1.32
| water |  |  | 1.5 | 19.8
| TOTAL |  |  | 18.4 | 242.88
tube:"100_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 100 | nm | 0.2 | 2.64
| water |  |  | 1.4 | 18.48
| TOTAL |  |  | 18.4 | 242.88
tube:"200_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 200 | nm | 0.4 | 5.28
| water |  |  | 1.2 | 15.84
| TOTAL |  |  | 18.4 | 242.88
tube:"400_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 400 | nm | 0.8 | 10.56
| water |  |  | 0.4 | 5.28
| TOTAL |  |  | 18 | 237.6
tube:"600_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 600 | nm | 1.2 | 15.84
| water |  |  | 0.4 | 5.28
| TOTAL |  |  | 18.4 | 242.88
tube:"800_mix"
| Component | Desired | Units | PROBE RXN | 13.2
| bpwd_mix |  |  | 16.8 | 221.76
| F_primer@10uM | various |  |  | 0
| R_primer@10uM | 800 | nm | 1.6 | 21.12
| water |  |  | 0 | 0
| TOTAL |  |  | 18.4 | 242.88

## F2_matrix
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | Ct_AVG |  | [F] | [R] | [F]_[R] | Ct_mean | TrendAvg | TrendStDev | 2*stdev | [F] | [R] | [F]_[R] | Ct_mean | TrendAvg | TrendStDev | [Primer] | Constant [R] | Constant [F]
A01 | IAC | Unknown | CY5 | 38.4 | A01 | IAC | Unknown | CY5 | 38.4 | 39.8 |  | 50 | 50 | 50_50 | 39.8 | 37.7075 | 1.2623 |  | 50 | 50 | 50_50 | 39.8 | 40.602 | #DIV/0! | 50 | 37.7075 | 40.602
A01 | Leg pneumo | Unknown | FAM | 29.94 | A02 | IAC | Unknown | CY5 | 41.2 |  |  | 100 | 50 | 100_50 | 36.75 |  |  |  | 50 | 100 | 50_100 | #DIV/0! |  |  | 100 | 35.249 | 36.5667
A02 | IAC | Unknown | CY5 | 41.2 | A03 | IAC | Unknown | CY5 | 36.75 | 36.75 |  | 200 | 50 | 200_50 | 38.07 |  |  |  | 50 | 200 | 50_200 | 42.09 |  |  | 200 | 34.8775 | 33.8175
A02 | Leg pneumo | Unknown | FAM | 30.19 | A04 | IAC | Unknown | CY5 |  |  |  | 400 | 50 | 400_50 | 37.84 |  |  |  | 50 | 400 | 50_400 | 36.81 |  |  | 400 | 33.4008 | 33.0483
A03 | IAC | Unknown | CY5 | 36.75 | A05 | IAC | Unknown | CY5 | 38.73 | 38.07 |  | 600 | 50 | 600_50 | 37.02 |  |  |  | 50 | 600 | 50_600 | 42.07 |  |  | 600 | 33.7583 | 32.6658
A03 | Leg pneumo | Unknown | FAM | 30.7 | A06 | IAC | Unknown | CY5 | 37.41 |  |  | 800 | 50 | 800_50 | 36.765 |  |  |  | 50 | 800 | 50_800 | 42.24 |  |  | 800 | 33.3208 | 32.5058
A04 | IAC | Unknown | CY5 |  | A07 | IAC | Unknown | CY5 | 38.02 | 37.84 |  | 50 | 100 | 50_100 | #DIV/0! | 35.249 | #DIV/0! |  | 100 | 50 | 100_50 | 36.75 | 36.5667 | 1.3579
A04 | Leg pneumo | Unknown | FAM | 30.81 | A08 | IAC | Unknown | CY5 | 37.66 |  |  | 100 | 100 | 100_100 | 38.91 |  |  |  | 100 | 100 | 100_100 | 38.91
A05 | IAC | Unknown | CY5 | 38.73 | A09 | IAC | Unknown | CY5 | 36.82 | 37.02 |  | 200 | 100 | 200_100 | 35.36 |  |  |  | 100 | 200 | 100_200 | 37.125
A05 | Leg pneumo | Unknown | FAM | 31.41 | A10 | IAC | Unknown | CY5 | 37.22 |  |  | 400 | 100 | 400_100 | 34.245 |  |  |  | 100 | 400 | 100_400 | 35.84
A06 | IAC | Unknown | CY5 | 37.41 | A11 | IAC | Unknown | CY5 | 36.21 | 36.765 |  | 600 | 100 | 600_100 | 33.92 |  |  |  | 100 | 600 | 100_600 | 35.58
A06 | Leg pneumo | Unknown | FAM | 30.9 | A12 | IAC | Unknown | CY5 | 37.32 |  |  | 800 | 100 | 800_100 | 33.81 |  |  |  | 100 | 800 | 100_800 | 35.195
A07 | IAC | Unknown | CY5 | 38.02 | B01 | IAC | Unknown | CY5 |  | #DIV/0! |  | 50 | 200 | 50_200 | 42.09 | 34.8775 | 4.0083 |  | 200 | 50 | 200_50 | 38.07 | 33.8175 | 2.489
A07 | Leg pneumo | Unknown | FAM | 31.37 | B02 | IAC | Unknown | CY5 |  |  |  | 100 | 200 | 100_200 | 37.125 |  |  |  | 200 | 100 | 200_100 | 35.36
A08 | IAC | Unknown | CY5 | 37.66 | B03 | IAC | Unknown | CY5 | 38.46 | 38.91 |  | 200 | 200 | 200_200 | 33.32 |  |  |  | 200 | 200 | 200_200 | 33.32
A08 | Leg pneumo | Unknown | FAM | 31.05 | B04 | IAC | Unknown | CY5 | 39.36 |  |  | 400 | 200 | 400_200 | 32.23 |  |  |  | 200 | 400 | 200_400 | 32.695
A09 | IAC | Unknown | CY5 | 36.82 | B05 | IAC | Unknown | CY5 | 35.26 | 35.36 |  | 600 | 200 | 600_200 | 32.35 |  |  |  | 200 | 600 | 200_600 | 32.11
A09 | Leg pneumo | Unknown | FAM | 31.45 | B06 | IAC | Unknown | CY5 | 35.46 |  |  | 800 | 200 | 800_200 | 32.15 |  |  |  | 200 | 800 | 200_800 | 31.35
A10 | IAC | Unknown | CY5 | 37.22 | B07 | IAC | Unknown | CY5 | 34.52 | 34.245 |  | 50 | 400 | 50_400 | 36.81 | 33.4008 | 2.3208 |  | 400 | 50 | 400_50 | 37.84 | 33.0483 | 2.6312
A10 | Leg pneumo | Unknown | FAM | 31.4 | B08 | IAC | Unknown | CY5 | 33.97 |  |  | 100 | 400 | 100_400 | 35.84 |  |  |  | 400 | 100 | 400_100 | 34.245
A11 | IAC | Unknown | CY5 | 36.21 | B09 | IAC | Unknown | CY5 | 33.94 | 33.92 |  | 200 | 400 | 200_400 | 32.695 |  |  |  | 400 | 200 | 400_200 | 32.23 | [5 empty] | y = 930.2554 + (32.44879 - 930.2554)/(1 + (x/0.002186158)^-0.9452625)^98.95066
A11 | Leg pneumo | Unknown | FAM | 30.76 | B10 | IAC | Unknown | CY5 | 33.9 |  |  | 400 | 400 | 400_400 | 31.795 |  |  |  | 400 | 400 | 400_400 | 31.795
A12 | IAC | Unknown | CY5 | 37.32 | B11 | IAC | Unknown | CY5 | 33.92 | 33.81 |  | 600 | 400 | 600_400 | 31.75 |  |  |  | 400 | 600 | 400_600 | 31.485
A12 | Leg pneumo | Unknown | FAM | 31.8 | B12 | IAC | Unknown | CY5 | 33.7 |  |  | 800 | 400 | 800_400 | 31.515 |  |  |  | 400 | 800 | 400_800 | 30.695
B01 | IAC | Unknown | CY5 |  | C01 | IAC | Unknown | CY5 | 42.09 | 42.09 |  | 50 | 600 | 50_600 | 42.07 | 33.7583 | 4.463 |  | 600 | 50 | 600_50 | 37.02 | 32.6658 | 2.4968
B01 | Leg pneumo | Unknown | FAM | 31.31 | C02 | IAC | Unknown | CY5 |  |  |  | 100 | 600 | 100_600 | 35.58 |  |  |  | 600 | 100 | 600_100 | 33.92
B02 | IAC | Unknown | CY5 |  | C03 | IAC | Unknown | CY5 | 36.24 | 37.125 |  | 200 | 600 | 200_600 | 32.11 |  |  |  | 600 | 200 | 600_200 | 32.35
B02 | Leg pneumo | Unknown | FAM | 31.28 | C04 | IAC | Unknown | CY5 | 38.01 |  |  | 400 | 600 | 400_600 | 31.485 |  |  |  | 600 | 400 | 600_400 | 31.75
B03 | IAC | Unknown | CY5 | 38.46 | C05 | IAC | Unknown | CY5 | 33.35 | 33.32 |  | 600 | 600 | 600_600 | 30.725 |  |  |  | 600 | 600 | 600_600 | 30.725
B03 | Leg pneumo | Unknown | FAM | 31.09 | C06 | IAC | Unknown | CY5 | 33.29 |  |  | 800 | 600 | 800_600 | 30.58 |  |  |  | 600 | 800 | 600_800 | 30.23
B04 | IAC | Unknown | CY5 | 39.36 | C07 | IAC | Unknown | CY5 | 32.32 | 32.23 |  | 50 | 800 | 50_800 | 42.24 | 33.3208 | 4.7548 |  | 800 | 50 | 800_50 | 36.765 | 32.5058 | 2.4457
B04 | Leg pneumo | Unknown | FAM | 30.92 | C08 | IAC | Unknown | CY5 | 32.14 |  |  | 100 | 800 | 100_800 | 35.195 |  |  |  | 800 | 100 | 800_100 | 33.81
B05 | IAC | Unknown | CY5 | 35.26 | C09 | IAC | Unknown | CY5 | 32.36 | 32.35 |  | 200 | 800 | 200_800 | 31.35 |  |  |  | 800 | 200 | 800_200 | 32.15
B05 | Leg pneumo | Unknown | FAM | 31.36 | C10 | IAC | Unknown | CY5 | 32.34 |  |  | 400 | 800 | 400_800 | 30.695 |  |  |  | 800 | 400 | 800_400 | 31.515
B06 | IAC | Unknown | CY5 | 35.46 | C11 | IAC | Unknown | CY5 | 31.94 | 32.15 |  | 600 | 800 | 600_800 | 30.23 |  |  |  | 800 | 600 | 800_600 | 30.58
B06 | Leg pneumo | Unknown | FAM | 31.11 | C12 | IAC | Unknown | CY5 | 32.36 |  |  | 800 | 800 | 800_800 | 30.215 |  |  |  | 800 | 800 | 800_800 | 30.215
B07 | IAC | Unknown | CY5 | 34.52 | D01 | IAC | Unknown | CY5 | 36.81 | 36.81
B07 | Leg pneumo | Unknown | FAM | 31.36 | D02 | IAC | Unknown | CY5
B08 | IAC | Unknown | CY5 | 33.97 | D03 | IAC | Unknown | CY5 | 35.58 | 35.84
B08 | Leg pneumo | Unknown | FAM | 31.02 | D04 | IAC | Unknown | CY5 | 36.1
B09 | IAC | Unknown | CY5 | 33.94 | D05 | IAC | Unknown | CY5 | 32.76 | 32.695
B09 | Leg pneumo | Unknown | FAM | 31.07 | D06 | IAC | Unknown | CY5 | 32.63
B10 | IAC | Unknown | CY5 | 33.9 | D07 | IAC | Unknown | CY5 | 31.72 | 31.795
B10 | Leg pneumo | Unknown | FAM | 31.18 | D08 | IAC | Unknown | CY5 | 31.87
B11 | IAC | Unknown | CY5 | 33.92 | D09 | IAC | Unknown | CY5 | 31.76 | 31.75
B11 | Leg pneumo | Unknown | FAM | 31.11 | D10 | IAC | Unknown | CY5 | 31.74
B12 | IAC | Unknown | CY5 | 33.7 | D11 | IAC | Unknown | CY5 | 31.43 | 31.515
B12 | Leg pneumo | Unknown | FAM | 31.36 | D12 | IAC | Unknown | CY5 | 31.6
C01 | IAC | Unknown | CY5 | 42.09 | E01 | IAC | Unknown | CY5 | 42.07 | 42.07
C01 | Leg pneumo | Unknown | FAM | 30.89 | E02 | IAC | Unknown | CY5
C02 | IAC | Unknown | CY5 |  | E03 | IAC | Unknown | CY5 | 35.19 | 35.58
C02 | Leg pneumo | Unknown | FAM | 30.67 | E04 | IAC | Unknown | CY5 | 35.97
C03 | IAC | Unknown | CY5 | 36.24 | E05 | IAC | Unknown | CY5 | 31.88 | 32.11
C03 | Leg pneumo | Unknown | FAM | 31.31 | E06 | IAC | Unknown | CY5 | 32.34
C04 | IAC | Unknown | CY5 | 38.01 | E07 | IAC | Unknown | CY5 | 31.66 | 31.485
C04 | Leg pneumo | Unknown | FAM | 30.92 | E08 | IAC | Unknown | CY5 | 31.31
C05 | IAC | Unknown | CY5 | 33.35 | E09 | IAC | Unknown | CY5 | 30.65 | 30.725
C05 | Leg pneumo | Unknown | FAM | 31.34 | E10 | IAC | Unknown | CY5 | 30.8 |  |  |  |  | IDT DNA Data
C06 | IAC | Unknown | CY5 | 33.29 | E11 | IAC | Unknown | CY5 | 30.42 | 30.58 |  | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/ul | Log(Copy#) | Cq_mean | Cq_stdev
C06 | Leg pneumo | Unknown | FAM | 31.51 | E12 | IAC | Unknown | CY5 | 30.74 |  |  |  | IDT Source | 100 | 100000000
C07 | IAC | Unknown | CY5 | 32.32 | F01 | IAC | Unknown | CY5 | 42.24 | 42.24 |  | Tube 2 |  | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | 11.3833 | 0.094516
C07 | Leg pneumo | Unknown | FAM | 31.25 | F02 | IAC | Unknown | CY5 |  |  |  | Tube 3 |  | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 9.0808 | 15.4133 | 0.066583
C08 | IAC | Unknown | CY5 | 32.14 | F03 | IAC | Unknown | CY5 | 34.48 | 35.195 |  | Tube 4 |  | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808 | 19.5033 | 0.10017
C08 | Leg pneumo | Unknown | FAM | 30.85 | F04 | IAC | Unknown | CY5 | 35.91 |  |  | Tube 5 |  | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808 | 22.66 | 0.14933
C09 | IAC | Unknown | CY5 | 32.36 | F05 | IAC | Unknown | CY5 | 30.96 | 31.35 |  | Tube 6 |  | 1e-06 | 1 | 602200000 | 1204400 | 6.0808 | 24.17 | 0.19519
C09 | Leg pneumo | Unknown | FAM | 31.39 | F06 | IAC | Unknown | CY5 | 31.74 |  |  | Tube 7 |  | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808 | 27.7033 | 0.33171
C10 | IAC | Unknown | CY5 | 32.34 | F07 | IAC | Unknown | CY5 | 30.95 | 30.695 |  | Tube 8 |  | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808 | 31.39 | 1.1018
C10 | Leg pneumo | Unknown | FAM | 30.9 | F08 | IAC | Unknown | CY5 | 30.44
C11 | IAC | Unknown | CY5 | 31.94 | F09 | IAC | Unknown | CY5 | 30.24 | 30.23 | [11 empty] | m | b
C11 | Leg pneumo | Unknown | FAM | 30.44 | F10 | IAC | Unknown | CY5 | 30.22 | [12 empty] | -3.1881 | 44.3204
C12 | IAC | Unknown | CY5 | 32.36 | F11 | IAC | Unknown | CY5 | 30.06 | 30.215
C12 | Leg pneumo | Unknown | FAM | 31.38 | F12 | IAC | Unknown | CY5 | 30.37 | [12 empty] | PCR Efficiency | 105.9%
D01 | IAC | Unknown | CY5 | 36.81 | G01 | IAC | Unknown | CY5 | 11.35 | 11.3833 | 0.094516
D01 | Leg pneumo | Unknown | FAM | 31.61 | G02 | IAC | Unknown | CY5 | 11.49
D02 | IAC | Unknown | CY5 |  | G03 | IAC | Unknown | CY5 | 11.31
D02 | Leg pneumo | Unknown | FAM | 30.99 | G04 | IAC | Unknown | CY5 | 15.37 | 15.4133 | 0.066583
D03 | IAC | Unknown | CY5 | 35.58 | G05 | IAC | Unknown | CY5 | 15.49
D03 | Leg pneumo | Unknown | FAM | 31.19 | G06 | IAC | Unknown | CY5 | 15.38
D04 | IAC | Unknown | CY5 | 36.1 | G07 | IAC | Unknown | CY5 | 19.58 | 19.5033 | 0.10017
D04 | Leg pneumo | Unknown | FAM | 31.42 | G08 | IAC | Unknown | CY5 | 19.54
D05 | IAC | Unknown | CY5 | 32.76 | G09 | IAC | Unknown | CY5 | 19.39
D05 | Leg pneumo | Unknown | FAM | 31.14 | G10 | IAC | Unknown | CY5 | 22.55 | 22.66 | 0.14933
D06 | IAC | Unknown | CY5 | 32.63 | G11 | IAC | Unknown | CY5 | 22.83
D06 | Leg pneumo | Unknown | FAM | 30.96 | G12 | IAC | Unknown | CY5 | 22.6
D07 | IAC | Unknown | CY5 | 31.72 | H01 | IAC | Unknown | CY5 | 23.98 | 24.17 | 0.19519
D07 | Leg pneumo | Unknown | FAM | 31.23 | H02 | IAC | Unknown | CY5 | 24.37
D08 | IAC | Unknown | CY5 | 31.87 | H03 | IAC | Unknown | CY5 | 24.16 |  |  |  |  | TubeName | Conc (uM) | pM | Copy Number | Copy Number/rxn | Log(copy#/rxn)
D08 | Leg pneumo | Unknown | FAM | 31.22 | H04 | IAC | Unknown | CY5 | 27.99 | 27.7033 | 0.33171 |  |  | IDT Source | 100 | 100000000
D09 | IAC | Unknown | CY5 | 31.76 | H05 | IAC | Unknown | CY5 | 27.78 |  |  |  |  | std_0 | 1 | 1000000 | 6.022e+14 | 1.2044e+12
D09 | Leg pneumo | Unknown | FAM | 30.77 | H06 | IAC | Unknown | CY5 | 27.34 |  |  |  |  | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808
D10 | IAC | Unknown | CY5 | 31.74 | H07 | IAC | Unknown | CY5 | 30.69 | 31.39 | 1.1018 |  |  | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808
D10 | Leg pneumo | Unknown | FAM | 30.73 | H08 | IAC | Unknown | CY5 | 30.82 |  |  |  |  | std_3 | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 9.0808
D11 | IAC | Unknown | CY5 | 31.43 | H09 | IAC | Unknown | CY5 | 32.66 |  |  |  |  | std_4 | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808
D11 | Leg pneumo | Unknown | FAM | 31.05 | H10 | IAC | Unknown | CY5 |  | #DIV/0! | #DIV/0! |  |  | std_5 | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808
D12 | IAC | Unknown | CY5 | 31.6 | H11 | IAC | Unknown | CY5 | [5 empty] | std_6 | 1e-06 | 1 | 602200000 | 1204400 | 6.0808
D12 | Leg pneumo | Unknown | FAM | 31.11 | H12 | IAC | Unknown | CY5 | [5 empty] | std_7 | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808
E01 | IAC | Unknown | CY5 | 42.07 | A01 | Leg pneumo | Unknown | FAM | 29.94 | 30.065 |  |  |  | std_8 | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808
E01 | Leg pneumo | Unknown | FAM | 31.15 | A02 | Leg pneumo | Unknown | FAM | 30.19
E02 | IAC | Unknown | CY5 |  | A03 | Leg pneumo | Unknown | FAM | 30.7 | 30.755 |  |  |  | Leg pneumo Variations in presence of IAC F, R Concentrations
E02 | Leg pneumo | Unknown | FAM | 31.26 | A04 | Leg pneumo | Unknown | FAM | 30.81 |  |  |  |  | [F] | [R] | [F]_[R] | Ct_mean
E03 | IAC | Unknown | CY5 | 35.19 | A05 | Leg pneumo | Unknown | FAM | 31.41 | 31.155 |  |  |  | 50 | 50 | 50_50 | 30.065
E03 | Leg pneumo | Unknown | FAM | 31 | A06 | Leg pneumo | Unknown | FAM | 30.9 |  |  |  |  | 100 | 50 | 100_50 | 30.755
E04 | IAC | Unknown | CY5 | 35.97 | A07 | Leg pneumo | Unknown | FAM | 31.37 | 31.21 |  |  |  | 200 | 50 | 200_50 | 31.155
E04 | Leg pneumo | Unknown | FAM | 31.24 | A08 | Leg pneumo | Unknown | FAM | 31.05 |  |  |  |  | 400 | 50 | 400_50 | 31.21
E05 | IAC | Unknown | CY5 | 31.88 | A09 | Leg pneumo | Unknown | FAM | 31.45 | 31.425 |  |  |  | 600 | 50 | 600_50 | 31.425
E05 | Leg pneumo | Unknown | FAM | 31.55 | A10 | Leg pneumo | Unknown | FAM | 31.4 |  |  |  |  | 800 | 50 | 800_50 | 31.28
E06 | IAC | Unknown | CY5 | 32.34 | A11 | Leg pneumo | Unknown | FAM | 30.76 | 31.28 |  |  |  | 50 | 100 | 50_100 | 31.295
E06 | Leg pneumo | Unknown | FAM | 31.07 | A12 | Leg pneumo | Unknown | FAM | 31.8 |  |  |  |  | 100 | 100 | 100_100 | 31.005
E07 | IAC | Unknown | CY5 | 31.66 | B01 | Leg pneumo | Unknown | FAM | 31.31 | 31.295 |  |  |  | 200 | 100 | 200_100 | 31.235
E07 | Leg pneumo | Unknown | FAM | 31.91 | B02 | Leg pneumo | Unknown | FAM | 31.28 |  |  |  |  | 400 | 100 | 400_100 | 31.19
E08 | IAC | Unknown | CY5 | 31.31 | B03 | Leg pneumo | Unknown | FAM | 31.09 | 31.005 |  |  |  | 600 | 100 | 600_100 | 31.125
E08 | Leg pneumo | Unknown | FAM | 31.1 | B04 | Leg pneumo | Unknown | FAM | 30.92 |  |  |  |  | 800 | 100 | 800_100 | 31.235
E09 | IAC | Unknown | CY5 | 30.65 | B05 | Leg pneumo | Unknown | FAM | 31.36 | 31.235 |  |  |  | 50 | 200 | 50_200 | 30.78
E09 | Leg pneumo | Unknown | FAM | 30.93 | B06 | Leg pneumo | Unknown | FAM | 31.11 |  |  |  |  | 100 | 200 | 100_200 | 31.115
E10 | IAC | Unknown | CY5 | 30.8 | B07 | Leg pneumo | Unknown | FAM | 31.36 | 31.19 |  |  |  | 200 | 200 | 200_200 | 31.425
E10 | Leg pneumo | Unknown | FAM | 30.83 | B08 | Leg pneumo | Unknown | FAM | 31.02 |  |  |  |  | 400 | 200 | 400_200 | 31.05
E11 | IAC | Unknown | CY5 | 30.42 | B09 | Leg pneumo | Unknown | FAM | 31.07 | 31.125 |  |  |  | 600 | 200 | 600_200 | 31.145
E11 | Leg pneumo | Unknown | FAM | 30.63 | B10 | Leg pneumo | Unknown | FAM | 31.18 |  |  |  |  | 800 | 200 | 800_200 | 30.91
E12 | IAC | Unknown | CY5 | 30.74 | B11 | Leg pneumo | Unknown | FAM | 31.11 | 31.235 |  |  |  | 50 | 400 | 50_400 | 31.3
E12 | Leg pneumo | Unknown | FAM | 31.13 | B12 | Leg pneumo | Unknown | FAM | 31.36 |  |  |  |  | 100 | 400 | 100_400 | 31.305
F01 | IAC | Unknown | CY5 | 42.24 | C01 | Leg pneumo | Unknown | FAM | 30.89 | 30.78 |  |  |  | 200 | 400 | 200_400 | 31.05
F01 | Leg pneumo | Unknown | FAM | 31.39 | C02 | Leg pneumo | Unknown | FAM | 30.67 |  |  |  |  | 400 | 400 | 400_400 | 31.225
F02 | IAC | Unknown | CY5 |  | C03 | Leg pneumo | Unknown | FAM | 31.31 | 31.115 |  |  |  | 600 | 400 | 600_400 | 30.75
F02 | Leg pneumo | Unknown | FAM | 30.99 | C04 | Leg pneumo | Unknown | FAM | 30.92 |  |  |  |  | 800 | 400 | 800_400 | 31.08
F03 | IAC | Unknown | CY5 | 34.48 | C05 | Leg pneumo | Unknown | FAM | 31.34 | 31.425 |  |  |  | 50 | 600 | 50_600 | 31.205
F03 | Leg pneumo | Unknown | FAM | 31.42 | C06 | Leg pneumo | Unknown | FAM | 31.51 |  |  |  |  | 100 | 600 | 100_600 | 31.12
F04 | IAC | Unknown | CY5 | 35.91 | C07 | Leg pneumo | Unknown | FAM | 31.25 | 31.05 |  |  |  | 200 | 600 | 200_600 | 31.31
F04 | Leg pneumo | Unknown | FAM | 31.51 | C08 | Leg pneumo | Unknown | FAM | 30.85 |  |  |  |  | 400 | 600 | 400_600 | 31.505
F05 | IAC | Unknown | CY5 | 30.96 | C09 | Leg pneumo | Unknown | FAM | 31.39 | 31.145 |  |  |  | 600 | 600 | 600_600 | 30.88
F05 | Leg pneumo | Unknown | FAM | 31.04 | C10 | Leg pneumo | Unknown | FAM | 30.9 |  |  |  |  | 800 | 600 | 800_600 | 30.88
F06 | IAC | Unknown | CY5 | 31.74 | C11 | Leg pneumo | Unknown | FAM | 30.44 | 30.91 |  |  |  | 50 | 800 | 50_800 | 31.19
F06 | Leg pneumo | Unknown | FAM | 31.67 | C12 | Leg pneumo | Unknown | FAM | 31.38 |  |  |  |  | 100 | 800 | 100_800 | 31.465
F07 | IAC | Unknown | CY5 | 30.95 | D01 | Leg pneumo | Unknown | FAM | 31.61 | 31.3 |  |  |  | 200 | 800 | 200_800 | 31.355
F07 | Leg pneumo | Unknown | FAM | 31.43 | D02 | Leg pneumo | Unknown | FAM | 30.99 |  |  |  |  | 400 | 800 | 400_800 | 31.205
F08 | IAC | Unknown | CY5 | 30.44 | D03 | Leg pneumo | Unknown | FAM | 31.19 | 31.305 |  |  |  | 600 | 800 | 600_800 | 31.605
F08 | Leg pneumo | Unknown | FAM | 30.98 | D04 | Leg pneumo | Unknown | FAM | 31.42 |  |  |  |  | 800 | 800 | 800_800 | 31.02
F09 | IAC | Unknown | CY5 | 30.24 | D05 | Leg pneumo | Unknown | FAM | 31.14 | 31.05
F09 | Leg pneumo | Unknown | FAM | 31.86 | D06 | Leg pneumo | Unknown | FAM | 30.96
F10 | IAC | Unknown | CY5 | 30.22 | D07 | Leg pneumo | Unknown | FAM | 31.23 | 31.225
F10 | Leg pneumo | Unknown | FAM | 31.35 | D08 | Leg pneumo | Unknown | FAM | 31.22
F11 | IAC | Unknown | CY5 | 30.06 | D09 | Leg pneumo | Unknown | FAM | 30.77 | 30.75
F11 | Leg pneumo | Unknown | FAM | 30.81 | D10 | Leg pneumo | Unknown | FAM | 30.73
F12 | IAC | Unknown | CY5 | 30.37 | D11 | Leg pneumo | Unknown | FAM | 31.05 | 31.08
F12 | Leg pneumo | Unknown | FAM | 31.23 | D12 | Leg pneumo | Unknown | FAM | 31.11
G01 | IAC | Unknown | CY5 | 11.35 | E01 | Leg pneumo | Unknown | FAM | 31.15 | 31.205
G01 | Leg pneumo | Unknown | FAM | 31.53 | E02 | Leg pneumo | Unknown | FAM | 31.26
G02 | IAC | Unknown | CY5 | 11.49 | E03 | Leg pneumo | Unknown | FAM | 31 | 31.12
G02 | Leg pneumo | Unknown | FAM | 31.94 | E04 | Leg pneumo | Unknown | FAM | 31.24
G03 | IAC | Unknown | CY5 | 11.31 | E05 | Leg pneumo | Unknown | FAM | 31.55 | 31.31
G03 | Leg pneumo | Unknown | FAM | 32.44 | E06 | Leg pneumo | Unknown | FAM | 31.07
G04 | IAC | Unknown | CY5 | 15.37 | E07 | Leg pneumo | Unknown | FAM | 31.91 | 31.505
G04 | Leg pneumo | Unknown | FAM | 31.85 | E08 | Leg pneumo | Unknown | FAM | 31.1
G05 | IAC | Unknown | CY5 | 15.49 | E09 | Leg pneumo | Unknown | FAM | 30.93 | 30.88
G05 | Leg pneumo | Unknown | FAM | 31.98 | E10 | Leg pneumo | Unknown | FAM | 30.83
G06 | IAC | Unknown | CY5 | 15.38 | E11 | Leg pneumo | Unknown | FAM | 30.63 | 30.88
G06 | Leg pneumo | Unknown | FAM | 32.25 | E12 | Leg pneumo | Unknown | FAM | 31.13
G07 | IAC | Unknown | CY5 | 19.58 | F01 | Leg pneumo | Unknown | FAM | 31.39 | 31.19
G07 | Leg pneumo | Unknown | FAM | 31.84 | F02 | Leg pneumo | Unknown | FAM | 30.99
G08 | IAC | Unknown | CY5 | 19.54 | F03 | Leg pneumo | Unknown | FAM | 31.42 | 31.465
G08 | Leg pneumo | Unknown | FAM | 31.86 | F04 | Leg pneumo | Unknown | FAM | 31.51
G09 | IAC | Unknown | CY5 | 19.39 | F05 | Leg pneumo | Unknown | FAM | 31.04 | 31.355
G09 | Leg pneumo | Unknown | FAM | 31.78 | F06 | Leg pneumo | Unknown | FAM | 31.67
G10 | IAC | Unknown | CY5 | 22.55 | F07 | Leg pneumo | Unknown | FAM | 31.43 | 31.205
G10 | Leg pneumo | Unknown | FAM | 30.87 | F08 | Leg pneumo | Unknown | FAM | 30.98
G11 | IAC | Unknown | CY5 | 22.83 | F09 | Leg pneumo | Unknown | FAM | 31.86 | 31.605
G11 | Leg pneumo | Unknown | FAM | 31.4 | F10 | Leg pneumo | Unknown | FAM | 31.35
G12 | IAC | Unknown | CY5 | 22.6 | F11 | Leg pneumo | Unknown | FAM | 30.81 | 31.02
G12 | Leg pneumo | Unknown | FAM | 31.76 | F12 | Leg pneumo | Unknown | FAM | 31.23
H01 | IAC | Unknown | CY5 | 23.98 | G01 | Leg pneumo | Unknown | FAM | 31.53 | 31.97
H01 | Leg pneumo | Unknown | FAM | 30.44 | G02 | Leg pneumo | Unknown | FAM | 31.94
H02 | IAC | Unknown | CY5 | 24.37 | G03 | Leg pneumo | Unknown | FAM | 32.44
H02 | Leg pneumo | Unknown | FAM | 31.29 | G04 | Leg pneumo | Unknown | FAM | 31.85 | 32.0267
H03 | IAC | Unknown | CY5 | 24.16 | G05 | Leg pneumo | Unknown | FAM | 31.98
H03 | Leg pneumo | Unknown | FAM | 31.19 | G06 | Leg pneumo | Unknown | FAM | 32.25
H04 | IAC | Unknown | CY5 | 27.99 | G07 | Leg pneumo | Unknown | FAM | 31.84 | 31.8267
H04 | Leg pneumo | Unknown | FAM | 31.2 | G08 | Leg pneumo | Unknown | FAM | 31.86
H05 | IAC | Unknown | CY5 | 27.78 | G09 | Leg pneumo | Unknown | FAM | 31.78
H05 | Leg pneumo | Unknown | FAM | 31.01 | G10 | Leg pneumo | Unknown | FAM | 30.87 | 31.3433
H06 | IAC | Unknown | CY5 | 27.34 | G11 | Leg pneumo | Unknown | FAM | 31.4
H06 | Leg pneumo | Unknown | FAM | 31.98 | G12 | Leg pneumo | Unknown | FAM | 31.76
H07 | IAC | Unknown | CY5 | 30.69 | H01 | Leg pneumo | Unknown | FAM | 30.44 | 30.9733
H07 | Leg pneumo | Unknown | FAM | 31.49 | H02 | Leg pneumo | Unknown | FAM | 31.29
H08 | IAC | Unknown | CY5 | 30.82 | H03 | Leg pneumo | Unknown | FAM | 31.19
H08 | Leg pneumo | Unknown | FAM | 31.32 | H04 | Leg pneumo | Unknown | FAM | 31.2 | 31.3967
H09 | IAC | Unknown | CY5 | 32.66 | H05 | Leg pneumo | Unknown | FAM | 31.01
H09 | Leg pneumo | Unknown | FAM | 32.89 | H06 | Leg pneumo | Unknown | FAM | 31.98
H10 | IAC | Unknown | CY5 |  | H07 | Leg pneumo | Unknown | FAM | 31.49 | 31.9
H10 | Leg pneumo | Unknown | FAM |  | H08 | Leg pneumo | Unknown | FAM | 31.32
H11 | IAC | Unknown | CY5 |  | H09 | Leg pneumo | Unknown | FAM | 32.89
H11 | Leg pneumo | Unknown | FAM |  | H10 | Leg pneumo | Unknown | FAM |  | #DIV/0!
H12 | IAC | Unknown | CY5 |  | H11 | Leg pneumo | Unknown | FAM
H12 | Leg pneumo | Unknown | FAM |  | H12 | Leg pneumo | Unknown | FAM
CONCLUSIONS:
Pattern suggests that 400nM F and 400nM is sufficient for IAC amplification. Concentrations greater than 400nM will be superflous.
PCR efficiency okay, within recommended limits: 90 to 110%.
Leg pneumo fluorescence unaffected by IAC F2, R primer concentrations from 50 to 800nM.
20210521
After probe data, adopting 300nM concentration.

## F2_probe_data
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | Ct_mean | Ct_stdDev | F_conc | R_conc | P_conc | Name | Ct_mean | Ct_stdev | 2*stdev | GroupMean_const[P] | Group_SD | F_conc | R_conc | P_conc | Name | Ct_mean | Ct_stdev | 2*stdev | GroupMean_cons[F,R] | Group_SD |  | Name | GroupMean_const[P] | Group_SD | 2*stdev | Name | GroupMean_cons[F,R] | Group_SD | 2*stdev
A01 | IAC | Unknown | CY5 | 35.13 | A01 | IAC | Unknown | CY5 | 35.13 | 36.07 | 0.63896 | 500 | 500 | 50 | 500_500_50 | 36.07 | 0.63896 | 1.2779 | 36.2208 | 1.2034 | 500 | 500 | 50 | 500_500_50 | 36.07 | 0.63896 | 1.2779 | 34.4129 | 0.83128 |  | 500_500_50 | 36.2208 | 1.2034 | 2.4067 | 500_500_50 | 34.4129 | 0.83128 | 1.6626
A01 | Leg pneumoe | Unknown | FAM | 23.22 | A02 | IAC | Unknown | CY5 | 36.35 |  |  | 400 | 400 | 50 | 400_400_50 | 37.4925 | 0.77466 | 1.5493 |  |  | 500 | 500 | 100 | 500_500_100 | 34.415 | 0.46307 | 0.92614 |  |  |  | 500_500_100 | 34.6775 | 0.99993 | 1.9999 | 400_400_50 | 35.2567 | 1.2745 | 2.549
A02 | IAC | Unknown | CY5 | 36.35 | A03 | IAC | Unknown | CY5 | 36.55 |  |  | 300 | 300 | 50 | 300_300_50 | 35.1 | 0.4706 | 0.9412 |  |  | 500 | 500 | 200 | 500_500_200 | 34.11 | 0.68838 | 1.3768 |  |  |  | 500_500_200 | 34.1667 | 0.8589 | 1.7178 | 300_300_50 | 33.4288 | 0.9761 | 1.9522
A02 | Leg pneumoe | Unknown | FAM | 23.58 | A04 | IAC | Unknown | CY5 | 36.25 |  |  | 500 | 500 | 100 | 500_500_100 | 34.415 | 0.46307 | 0.92614 | 34.6775 | 0.99993 | 500 | 500 | 400 | 500_500_400 | 33.9825 | 0.22485 | 0.4497 |  |  |  | 500_500_400 | 34.0692 | 0.89565 | 1.7913
A03 | IAC | Unknown | CY5 | 36.55 | A05 | IAC | Unknown | CY5 | 36.54 | 37.4925 | 0.77466 | 400 | 400 | 100 | 400_400_100 | 35.7825 | 0.7762 | 1.5524 |  |  | 500 | 500 | 600 | 500_500_600 | 33.9075 | 0.48986 | 0.97971 |  |  |  | 500_500_600 | 33.5475 | 0.61705 | 1.2341
A03 | Leg pneumoe | Unknown | FAM | 23.09 | A06 | IAC | Unknown | CY5 | 37.7 |  |  | 300 | 300 | 100 | 300_300_100 | 33.835 | 0.52798 | 1.056 |  |  | 500 | 500 | 800 | 500_500_800 | 33.9925 | 0.43722 | 0.87443 |  |  |  | 500_500_800 | 33.515 | 1.1111 | 2.2221
A04 | IAC | Unknown | CY5 | 36.25 | A07 | IAC | Unknown | CY5 | 38.4 |  |  | 500 | 500 | 200 | 500_500_200 | 34.11 | 0.68838 | 1.3768 | 34.1667 | 0.8589 | 400 | 400 | 50 | 400_400_50 | 37.4925 | 0.77466 | 1.5493 | 35.2567 | 1.2745
A04 | Leg pneumoe | Unknown | FAM | 24.28 | A08 | IAC | Unknown | CY5 | 37.33 |  |  | 400 | 400 | 200 | 400_400_200 | 35.0525 | 0.41274 | 0.82549 |  |  | 400 | 400 | 100 | 400_400_100 | 35.7825 | 0.7762 | 1.5524
A05 | IAC | Unknown | CY5 | 36.54 | A09 | IAC | Unknown | CY5 | 34.73 | 35.1 | 0.4706 | 300 | 300 | 200 | 300_300_200 | 33.3375 | 0.46807 | 0.93614 |  |  | 400 | 400 | 200 | 400_400_200 | 35.0525 | 0.41274 | 0.82549
A05 | Leg pneumoe | Unknown | FAM | 23.15 | A10 | IAC | Unknown | CY5 | 35.73 |  |  | 500 | 500 | 400 | 500_500_400 | 33.9825 | 0.22485 | 0.4497 | 34.0692 | 0.89565 | 400 | 400 | 400 | 400_400_400 | 35.005 | 0.73396 | 1.4679
A06 | IAC | Unknown | CY5 | 37.7 | A11 | IAC | Unknown | CY5 | 35.19 |  |  | 400 | 400 | 400 | 400_400_400 | 35.005 | 0.73396 | 1.4679 |  |  | 400 | 400 | 600 | 400_400_600 | 33.9 | 0.60349 | 1.207
A06 | Leg pneumoe | Unknown | FAM | 23.3 | A12 | IAC | Unknown | CY5 | 34.75 |  |  | 300 | 300 | 400 | 300_300_400 | 33.22 | 0.42032 | 0.84063 |  |  | 400 | 400 | 800 | 400_400_800 | 34.3075 | 0.30761 | 0.61522
A07 | IAC | Unknown | CY5 | 38.4 | B01 | IAC | Unknown | CY5 | 33.96 | 34.415 | 0.46307 | 500 | 500 | 600 | 500_500_600 | 33.9075 | 0.48986 | 0.97971 | 33.5475 | 0.61705 | 300 | 300 | 50 | 300_300_50 | 35.1 | 0.4706 | 0.9412 | 33.4288 | 0.9761
A07 | Leg pneumoe | Unknown | FAM | 23.72 | B02 | IAC | Unknown | CY5 | 34.92 |  |  | 400 | 400 | 600 | 400_400_600 | 33.9 | 0.60349 | 1.207 |  |  | 300 | 300 | 100 | 300_300_100 | 33.835 | 0.52798 | 1.056
A08 | IAC | Unknown | CY5 | 37.33 | B03 | IAC | Unknown | CY5 | 34.09 |  |  | 300 | 300 | 600 | 300_300_600 | 32.835 | 0.28525 | 0.5705 |  |  | 300 | 300 | 200 | 300_300_200 | 33.3375 | 0.46807 | 0.93614
A08 | Leg pneumoe | Unknown | FAM | 23.91 | B04 | IAC | Unknown | CY5 | 34.69 |  |  | 500 | 500 | 800 | 500_500_800 | 33.9925 | 0.43722 | 0.87443 | 33.515 | 1.1111 | 300 | 300 | 400 | 300_300_400 | 33.22 | 0.42032 | 0.84063
A09 | IAC | Unknown | CY5 | 34.73 | B05 | IAC | Unknown | CY5 | 34.69 | 35.7825 | 0.7762 | 400 | 400 | 800 | 400_400_800 | 34.3075 | 0.30761 | 0.61522 |  |  | 300 | 300 | 600 | 300_300_600 | 32.835 | 0.28525 | 0.5705
A09 | Leg pneumoe | Unknown | FAM | 22.75 | B06 | IAC | Unknown | CY5 | 35.87 |  |  | 300 | 300 | 800 | 300_300_800 | 32.245 | 0.31964 | 0.63927 |  |  | 300 | 300 | 800 | 300_300_800 | 32.245 | 0.31964 | 0.63927
A10 | IAC | Unknown | CY5 | 35.73 | B07 | IAC | Unknown | CY5 | 36.51
A10 | Leg pneumoe | Unknown | FAM | 23.96 | B08 | IAC | Unknown | CY5 | 36.06
A11 | IAC | Unknown | CY5 | 35.19 | B09 | IAC | Unknown | CY5 | 33.06 | 33.835 | 0.52798
A11 | Leg pneumoe | Unknown | FAM | 23.36 | B10 | IAC | Unknown | CY5 | 33.98
A12 | IAC | Unknown | CY5 | 34.75 | B11 | IAC | Unknown | CY5 | 34.24
A12 | Leg pneumoe | Unknown | FAM | 23.24 | B12 | IAC | Unknown | CY5 | 34.06
B01 | IAC | Unknown | CY5 | 33.96 | C01 | IAC | Unknown | CY5 | 33.19 | 34.11 | 0.68838
B01 | Leg pneumoe | Unknown | FAM | 22.73 | C02 | IAC | Unknown | CY5 | 33.99
B02 | IAC | Unknown | CY5 | 34.92 | C03 | IAC | Unknown | CY5 | 34.53
B02 | Leg pneumoe | Unknown | FAM | 23.35 | C04 | IAC | Unknown | CY5 | 34.73
B03 | IAC | Unknown | CY5 | 34.09 | C05 | IAC | Unknown | CY5 | 34.68 | 35.0525 | 0.41274
B03 | Leg pneumoe | Unknown | FAM | 23.26 | C06 | IAC | Unknown | CY5 | 34.9
B04 | IAC | Unknown | CY5 | 34.69 | C07 | IAC | Unknown | CY5 | 34.99
B04 | Leg pneumoe | Unknown | FAM | 23.57 | C08 | IAC | Unknown | CY5 | 35.64
B05 | IAC | Unknown | CY5 | 34.69 | C09 | IAC | Unknown | CY5 | 32.76 | 33.3375 | 0.46807
B05 | Leg pneumoe | Unknown | FAM | 23.63 | C10 | IAC | Unknown | CY5 | 33.56
B06 | IAC | Unknown | CY5 | 35.87 | C11 | IAC | Unknown | CY5 | 33.19
B06 | Leg pneumoe | Unknown | FAM | 23.26 | C12 | IAC | Unknown | CY5 | 33.84
B07 | IAC | Unknown | CY5 | 36.51 | D01 | IAC | Unknown | CY5 | 33.7 | 33.9825 | 0.22485
B07 | Leg pneumoe | Unknown | FAM | 23.48 | D02 | IAC | Unknown | CY5 | 34.09
B08 | IAC | Unknown | CY5 | 36.06 | D03 | IAC | Unknown | CY5 | 33.92
B08 | Leg pneumoe | Unknown | FAM | 23.2 | D04 | IAC | Unknown | CY5 | 34.22
B09 | IAC | Unknown | CY5 | 33.06 | D05 | IAC | Unknown | CY5 | 34.25 | 35.005 | 0.73396
B09 | Leg pneumoe | Unknown | FAM | 22.58 | D06 | IAC | Unknown | CY5 | 34.52
B10 | IAC | Unknown | CY5 | 33.98 | D07 | IAC | Unknown | CY5 | 35.48
B10 | Leg pneumoe | Unknown | FAM | 22.9 | D08 | IAC | Unknown | CY5 | 35.77
B11 | IAC | Unknown | CY5 | 34.24 | D09 | IAC | Unknown | CY5 | 32.73 | 33.22 | 0.42032
B11 | Leg pneumoe | Unknown | FAM | 23.11 | D10 | IAC | Unknown | CY5 | 33.01
B12 | IAC | Unknown | CY5 | 34.06 | D11 | IAC | Unknown | CY5 | 33.55
B12 | Leg pneumoe | Unknown | FAM | 24.02 | D12 | IAC | Unknown | CY5 | 33.59
C01 | IAC | Unknown | CY5 | 33.19 | E01 | IAC | Unknown | CY5 | 33.46 | 33.9075 | 0.48986
C01 | Leg pneumoe | Unknown | FAM | 22.5 | E02 | IAC | Unknown | CY5 | 33.65
C02 | IAC | Unknown | CY5 | 33.99 | E03 | IAC | Unknown | CY5 | 33.94
C02 | Leg pneumoe | Unknown | FAM | 23.45 | E04 | IAC | Unknown | CY5 | 34.58
C03 | IAC | Unknown | CY5 | 34.53 | E05 | IAC | Unknown | CY5 | 33.37 | 33.9 | 0.60349
C03 | Leg pneumoe | Unknown | FAM | 23.23 | E06 | IAC | Unknown | CY5 | 33.4
C04 | IAC | Unknown | CY5 | 34.73 | E07 | IAC | Unknown | CY5 | 34.29
C04 | Leg pneumoe | Unknown | FAM | 23.33 | E08 | IAC | Unknown | CY5 | 34.54
C05 | IAC | Unknown | CY5 | 34.68 | E09 | IAC | Unknown | CY5 | 32.43 | 32.835 | 0.28525
C05 | Leg pneumoe | Unknown | FAM | 23.28 | E10 | IAC | Unknown | CY5 | 32.9
C06 | IAC | Unknown | CY5 | 34.9 | E11 | IAC | Unknown | CY5 | 32.91
C06 | Leg pneumoe | Unknown | FAM | 22.77 | E12 | IAC | Unknown | CY5 | 33.1
C07 | IAC | Unknown | CY5 | 34.99 | F01 | IAC | Unknown | CY5 | 33.36 | 33.9925 | 0.43722
C07 | Leg pneumoe | Unknown | FAM | 23.27 | F02 | IAC | Unknown | CY5 | 34.29
C08 | IAC | Unknown | CY5 | 35.64 | F03 | IAC | Unknown | CY5 | 34.04
C08 | Leg pneumoe | Unknown | FAM | 23.52 | F04 | IAC | Unknown | CY5 | 34.28
C09 | IAC | Unknown | CY5 | 32.76 | F05 | IAC | Unknown | CY5 | 34.13 | 34.3075 | 0.30761
C09 | Leg pneumoe | Unknown | FAM | 22.88 | F06 | IAC | Unknown | CY5 | 34.1
C10 | IAC | Unknown | CY5 | 33.56 | F07 | IAC | Unknown | CY5 | 34.76
C10 | Leg pneumoe | Unknown | FAM | 23.52 | F08 | IAC | Unknown | CY5 | 34.24
C11 | IAC | Unknown | CY5 | 33.19 | F09 | IAC | Unknown | CY5 | 32.08 | 32.245 | 0.31964
C11 | Leg pneumoe | Unknown | FAM | 22.87 | F10 | IAC | Unknown | CY5 | 31.91
C12 | IAC | Unknown | CY5 | 33.84 | F11 | IAC | Unknown | CY5 | 32.64
C12 | Leg pneumoe | Unknown | FAM | 23.41 | F12 | IAC | Unknown | CY5 | 32.35 | [15 empty] | this is pasted from other tab; confirming values on left.
D01 | IAC | Unknown | CY5 | 33.7 | G01 | IAC | Unknown | CY5 | 12.48 | 11.9567 | 0.69357 | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/ul | Log(Copy#) | Cq_mean | Cq_stdev |  |  |  |  | TubeName | Conc (uM) | pM | Copy Number | Copy Number/rxn | Log(copy#/rxn)
D01 | Leg pneumoe | Unknown | FAM | 22.99 | G02 | IAC | Unknown | CY5 | 12.22 |  |  |  | IDT Source | 100 | 100000000 | [9 empty] | IDT Source | 100 | 100000000
D02 | IAC | Unknown | CY5 | 34.09 | G03 | IAC | Unknown | CY5 | 11.17 |  |  | Tube 2 |  | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | 11.9567 | 0.69357 |  |  |  |  | std_0 | 1 | 1000000 | 6.022e+14 | 1.2044e+12
D02 | Leg pneumoe | Unknown | FAM | 23.77 | G04 | IAC | Unknown | CY5 | 15.69 | 15.7067 | 0.086217 | Tube 3 |  | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 9.0808 | 15.7067 | 0.086217 |  |  |  |  | std_1 | 0.1 | 100000 | 6.022e+13 | 120440000000 | 11.0808
D03 | IAC | Unknown | CY5 | 33.92 | G05 | IAC | Unknown | CY5 | 15.8 |  |  | Tube 4 |  | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808 | 20.23 | 0.14 |  |  |  |  | std_2 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808
D03 | Leg pneumoe | Unknown | FAM | 22.95 | G06 | IAC | Unknown | CY5 | 15.63 |  |  | Tube 5 |  | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808 | 23.49 | 0.11136 |  |  |  |  | std_3 | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 9.0808
D04 | IAC | Unknown | CY5 | 34.22 | G07 | IAC | Unknown | CY5 | 20.37 | 20.23 | 0.14 | Tube 6 |  | 1e-06 | 1 | 602200000 | 1204400 | 6.0808 | 24.6633 | 0.1106 |  |  |  |  | std_4 | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808
D04 | Leg pneumoe | Unknown | FAM | 23.45 | G08 | IAC | Unknown | CY5 | 20.23 |  |  | Tube 7 |  | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808 | 28.3767 | 0.21127 |  |  |  |  | std_5 | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808
D05 | IAC | Unknown | CY5 | 34.25 | G09 | IAC | Unknown | CY5 | 20.09 |  |  | Tube 8 |  | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808 | 31.0867 | 2.5174 |  |  |  |  | std_6 | 1e-06 | 1 | 602200000 | 1204400 | 6.0808
D05 | Leg pneumoe | Unknown | FAM | 22.69 | G10 | IAC | Unknown | CY5 | 23.39 | 23.49 | 0.11136 | [13 empty] | std_7 | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808
D06 | IAC | Unknown | CY5 | 34.52 | G11 | IAC | Unknown | CY5 | 23.47 | [15 empty] | std_8 | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808
D06 | Leg pneumoe | Unknown | FAM | 21.72 | G12 | IAC | Unknown | CY5 | 23.61
D07 | IAC | Unknown | CY5 | 35.48 | H01 | IAC | Unknown | CY5 | 24.78 | 24.6633 | 0.1106 | [10 empty] | m | b
D07 | Leg pneumoe | Unknown | FAM | 23.03 | H02 | IAC | Unknown | CY5 | 24.65 | [12 empty] | -3.113 | 44.258
D08 | IAC | Unknown | CY5 | 35.77 | H03 | IAC | Unknown | CY5 | 24.56
D08 | Leg pneumoe | Unknown | FAM | 23.23 | H04 | IAC | Unknown | CY5 | 28.6 | 28.3767 | 0.21127 | [10 empty] | PCR Efficiency | 109.5%
D09 | IAC | Unknown | CY5 | 32.73 | H05 | IAC | Unknown | CY5 | 28.35
D09 | Leg pneumoe | Unknown | FAM | 22.29 | H06 | IAC | Unknown | CY5 | 28.18
D10 | IAC | Unknown | CY5 | 33.01 | H07 | IAC | Unknown | CY5 | 32.51 | 31.0867 | 2.5174
D10 | Leg pneumoe | Unknown | FAM | 22.53 | H08 | IAC | Unknown | CY5 | 32.57
D11 | IAC | Unknown | CY5 | 33.55 | H09 | IAC | Unknown | CY5 | 31.75
D11 | Leg pneumoe | Unknown | FAM | 23.18 | H10 | IAC | Unknown | CY5 | 36.91
D12 | IAC | Unknown | CY5 | 33.59 | H11 | IAC | Unknown | CY5 | 38.57
D12 | Leg pneumoe | Unknown | FAM | 23.37 | H12 | IAC | Unknown | CY5
E01 | IAC | Unknown | CY5 | 33.46 | A01 | Leg pneumoe | Unknown | FAM | 23.22 | 23.5425 | 0.53357
E01 | Leg pneumoe | Unknown | FAM | 22.17 | A02 | Leg pneumoe | Unknown | FAM | 23.58
E02 | IAC | Unknown | CY5 | 33.65 | A03 | Leg pneumoe | Unknown | FAM | 23.09
E02 | Leg pneumoe | Unknown | FAM | 22.91 | A04 | Leg pneumoe | Unknown | FAM | 24.28
E03 | IAC | Unknown | CY5 | 33.94 | A05 | Leg pneumoe | Unknown | FAM | 23.15 | 23.52 | 0.35468
E03 | Leg pneumoe | Unknown | FAM | 23.14 | A06 | Leg pneumoe | Unknown | FAM | 23.3
E04 | IAC | Unknown | CY5 | 34.58 | A07 | Leg pneumoe | Unknown | FAM | 23.72 |  |  |  | F_conc | R_conc | P_conc | Name | Ct_mean | Ct_stdev | 2*stdev
E04 | Leg pneumoe | Unknown | FAM | 23.24 | A08 | Leg pneumoe | Unknown | FAM | 23.91 |  |  |  | 500 | 500 | 50 | 500_500_50 | 23.5425 | 0.53357 | 1.0671
E05 | IAC | Unknown | CY5 | 33.37 | A09 | Leg pneumoe | Unknown | FAM | 22.75 | 23.3275 | 0.49742 |  | 400 | 400 | 50 | 400_400_50 | 23.52 | 0.35468 | 0.70937
E05 | Leg pneumoe | Unknown | FAM | 22.47 | A10 | Leg pneumoe | Unknown | FAM | 23.96 |  |  |  | 300 | 300 | 50 | 300_300_50 | 23.3275 | 0.49742 | 0.99484
E06 | IAC | Unknown | CY5 | 33.4 | A11 | Leg pneumoe | Unknown | FAM | 23.36 |  |  |  | 500 | 500 | 100 | 500_500_100 | 23.2275 | 0.35631 | 0.71262
E06 | Leg pneumoe | Unknown | FAM | 23.55 | A12 | Leg pneumoe | Unknown | FAM | 23.24 |  |  |  | 400 | 400 | 100 | 400_400_100 | 23.3925 | 0.19889 | 0.39779
E07 | IAC | Unknown | CY5 | 34.29 | B01 | Leg pneumoe | Unknown | FAM | 22.73 | 23.2275 | 0.35631 |  | 300 | 300 | 100 | 300_300_100 | 23.1525 | 0.61803 | 1.2361
E07 | Leg pneumoe | Unknown | FAM | 23.6 | B02 | Leg pneumoe | Unknown | FAM | 23.35 |  |  |  | 500 | 500 | 200 | 500_500_200 | 23.1275 | 0.42789 | 0.85578
E08 | IAC | Unknown | CY5 | 34.54 | B03 | Leg pneumoe | Unknown | FAM | 23.26 |  |  |  | 400 | 400 | 200 | 400_400_200 | 23.21 | 0.31528 | 0.63056
E08 | Leg pneumoe | Unknown | FAM | 23.1 | B04 | Leg pneumoe | Unknown | FAM | 23.57 |  |  |  | 300 | 300 | 200 | 300_300_200 | 23.17 | 0.34361 | 0.68722
E09 | IAC | Unknown | CY5 | 32.43 | B05 | Leg pneumoe | Unknown | FAM | 23.63 | 23.3925 | 0.19889 |  | 500 | 500 | 400 | 500_500_400 | 23.29 | 0.39226 | 0.78452
E09 | Leg pneumoe | Unknown | FAM | 23.03 | B06 | Leg pneumoe | Unknown | FAM | 23.26 |  |  |  | 400 | 400 | 400 | 400_400_400 | 22.6675 | 0.66984 | 1.3397
E10 | IAC | Unknown | CY5 | 32.9 | B07 | Leg pneumoe | Unknown | FAM | 23.48 |  |  |  | 300 | 300 | 400 | 300_300_400 | 22.8425 | 0.51481 | 1.0296
E10 | Leg pneumoe | Unknown | FAM | 22.85 | B08 | Leg pneumoe | Unknown | FAM | 23.2 |  |  |  | 500 | 500 | 600 | 500_500_600 | 22.865 | 0.48349 | 0.96699
E11 | IAC | Unknown | CY5 | 32.91 | B09 | Leg pneumoe | Unknown | FAM | 22.58 | 23.1525 | 0.61803 |  | 400 | 400 | 600 | 400_400_600 | 23.18 | 0.52402 | 1.048
E11 | Leg pneumoe | Unknown | FAM | 22.91 | B10 | Leg pneumoe | Unknown | FAM | 22.9 |  |  |  | 300 | 300 | 600 | 300_300_600 | 23.025 | 0.20421 | 0.40841
E12 | IAC | Unknown | CY5 | 33.1 | B11 | Leg pneumoe | Unknown | FAM | 23.11 |  |  |  | 500 | 500 | 800 | 500_500_800 | 23.5975 | 0.59763 | 1.1953
E12 | Leg pneumoe | Unknown | FAM | 23.31 | B12 | Leg pneumoe | Unknown | FAM | 24.02 |  |  |  | 400 | 400 | 800 | 400_400_800 | 23.9575 | 0.31127 | 0.62255
F01 | IAC | Unknown | CY5 | 33.36 | C01 | Leg pneumoe | Unknown | FAM | 22.5 | 23.1275 | 0.42789 |  | 300 | 300 | 800 | 300_300_800 | 23.415 | 0.4846 | 0.96919
F01 | Leg pneumoe | Unknown | FAM | 23.09 | C02 | Leg pneumoe | Unknown | FAM | 23.45
F02 | IAC | Unknown | CY5 | 34.29 | C03 | Leg pneumoe | Unknown | FAM | 23.23
F02 | Leg pneumoe | Unknown | FAM | 24.33 | C04 | Leg pneumoe | Unknown | FAM | 23.33
F03 | IAC | Unknown | CY5 | 34.04 | C05 | Leg pneumoe | Unknown | FAM | 23.28 | 23.21 | 0.31528
F03 | Leg pneumoe | Unknown | FAM | 23.13 | C06 | Leg pneumoe | Unknown | FAM | 22.77
F04 | IAC | Unknown | CY5 | 34.28 | C07 | Leg pneumoe | Unknown | FAM | 23.27
F04 | Leg pneumoe | Unknown | FAM | 23.84 | C08 | Leg pneumoe | Unknown | FAM | 23.52
F05 | IAC | Unknown | CY5 | 34.13 | C09 | Leg pneumoe | Unknown | FAM | 22.88 | 23.17 | 0.34361
F05 | Leg pneumoe | Unknown | FAM | 23.59 | C10 | Leg pneumoe | Unknown | FAM | 23.52
F06 | IAC | Unknown | CY5 | 34.1 | C11 | Leg pneumoe | Unknown | FAM | 22.87
F06 | Leg pneumoe | Unknown | FAM | 23.84 | C12 | Leg pneumoe | Unknown | FAM | 23.41
F07 | IAC | Unknown | CY5 | 34.76 | D01 | Leg pneumoe | Unknown | FAM | 22.99 | 23.29 | 0.39226
F07 | Leg pneumoe | Unknown | FAM | 24.09 | D02 | Leg pneumoe | Unknown | FAM | 23.77
F08 | IAC | Unknown | CY5 | 34.24 | D03 | Leg pneumoe | Unknown | FAM | 22.95
F08 | Leg pneumoe | Unknown | FAM | 24.31 | D04 | Leg pneumoe | Unknown | FAM | 23.45
F09 | IAC | Unknown | CY5 | 32.08 | D05 | Leg pneumoe | Unknown | FAM | 22.69 | 22.6675 | 0.66984
F09 | Leg pneumoe | Unknown | FAM | 23.65 | D06 | Leg pneumoe | Unknown | FAM | 21.72
F10 | IAC | Unknown | CY5 | 31.91 | D07 | Leg pneumoe | Unknown | FAM | 23.03
F10 | Leg pneumoe | Unknown | FAM | 23.18 | D08 | Leg pneumoe | Unknown | FAM | 23.23
F11 | IAC | Unknown | CY5 | 32.64 | D09 | Leg pneumoe | Unknown | FAM | 22.29 | 22.8425 | 0.51481
F11 | Leg pneumoe | Unknown | FAM | 22.87 | D10 | Leg pneumoe | Unknown | FAM | 22.53
F12 | IAC | Unknown | CY5 | 32.35 | D11 | Leg pneumoe | Unknown | FAM | 23.18
F12 | Leg pneumoe | Unknown | FAM | 23.96 | D12 | Leg pneumoe | Unknown | FAM | 23.37
G01 | IAC | Unknown | CY5 | 12.48 | E01 | Leg pneumoe | Unknown | FAM | 22.17 | 22.865 | 0.48349
G01 | Leg pneumoe | Unknown | FAM | 23.41 | E02 | Leg pneumoe | Unknown | FAM | 22.91
G02 | IAC | Unknown | CY5 | 12.22 | E03 | Leg pneumoe | Unknown | FAM | 23.14
G02 | Leg pneumoe | Unknown | FAM | 23.3 | E04 | Leg pneumoe | Unknown | FAM | 23.24
G03 | IAC | Unknown | CY5 | 11.17 | E05 | Leg pneumoe | Unknown | FAM | 22.47 | 23.18 | 0.52402
G03 | Leg pneumoe | Unknown | FAM | 23.59 | E06 | Leg pneumoe | Unknown | FAM | 23.55
G04 | IAC | Unknown | CY5 | 15.69 | E07 | Leg pneumoe | Unknown | FAM | 23.6
G04 | Leg pneumoe | Unknown | FAM | 23.05 | E08 | Leg pneumoe | Unknown | FAM | 23.1
G05 | IAC | Unknown | CY5 | 15.8 | E09 | Leg pneumoe | Unknown | FAM | 23.03 | 23.025 | 0.20421
G05 | Leg pneumoe | Unknown | FAM | 23.39 | E10 | Leg pneumoe | Unknown | FAM | 22.85
G06 | IAC | Unknown | CY5 | 15.63 | E11 | Leg pneumoe | Unknown | FAM | 22.91
G06 | Leg pneumoe | Unknown | FAM | 23.44 | E12 | Leg pneumoe | Unknown | FAM | 23.31
G07 | IAC | Unknown | CY5 | 20.37 | F01 | Leg pneumoe | Unknown | FAM | 23.09 | 23.5975 | 0.59763
G07 | Leg pneumoe | Unknown | FAM | 22.96 | F02 | Leg pneumoe | Unknown | FAM | 24.33
G08 | IAC | Unknown | CY5 | 20.23 | F03 | Leg pneumoe | Unknown | FAM | 23.13
G08 | Leg pneumoe | Unknown | FAM | 23.64 | F04 | Leg pneumoe | Unknown | FAM | 23.84
G09 | IAC | Unknown | CY5 | 20.09 | F05 | Leg pneumoe | Unknown | FAM | 23.59 | 23.9575 | 0.31127
G09 | Leg pneumoe | Unknown | FAM | 23.57 | F06 | Leg pneumoe | Unknown | FAM | 23.84
G10 | IAC | Unknown | CY5 | 23.39 | F07 | Leg pneumoe | Unknown | FAM | 24.09
G10 | Leg pneumoe | Unknown | FAM | 22.94 | F08 | Leg pneumoe | Unknown | FAM | 24.31
G11 | IAC | Unknown | CY5 | 23.47 | F09 | Leg pneumoe | Unknown | FAM | 23.65 | 23.415 | 0.4846
G11 | Leg pneumoe | Unknown | FAM | 22.61 | F10 | Leg pneumoe | Unknown | FAM | 23.18
G12 | IAC | Unknown | CY5 | 23.61 | F11 | Leg pneumoe | Unknown | FAM | 22.87
G12 | Leg pneumoe | Unknown | FAM | 23.51 | F12 | Leg pneumoe | Unknown | FAM | 23.96
H01 | IAC | Unknown | CY5 | 24.78 | G01 | Leg pneumoe | Unknown | FAM | 23.41
H01 | Leg pneumoe | Unknown | FAM | 22.33 | G02 | Leg pneumoe | Unknown | FAM | 23.3
H02 | IAC | Unknown | CY5 | 24.65 | G03 | Leg pneumoe | Unknown | FAM | 23.59
H02 | Leg pneumoe | Unknown | FAM | 22.82 | G04 | Leg pneumoe | Unknown | FAM | 23.05
H03 | IAC | Unknown | CY5 | 24.56 | G05 | Leg pneumoe | Unknown | FAM | 23.39
H03 | Leg pneumoe | Unknown | FAM | 22.73 | G06 | Leg pneumoe | Unknown | FAM | 23.44
H04 | IAC | Unknown | CY5 | 28.6 | G07 | Leg pneumoe | Unknown | FAM | 22.96
H04 | Leg pneumoe | Unknown | FAM | 23.52 | G08 | Leg pneumoe | Unknown | FAM | 23.64
H05 | IAC | Unknown | CY5 | 28.35 | G09 | Leg pneumoe | Unknown | FAM | 23.57
H05 | Leg pneumoe | Unknown | FAM | 23.45 | G10 | Leg pneumoe | Unknown | FAM | 22.94
H06 | IAC | Unknown | CY5 | 28.18 | G11 | Leg pneumoe | Unknown | FAM | 22.61
H06 | Leg pneumoe | Unknown | FAM | 23.84 | G12 | Leg pneumoe | Unknown | FAM | 23.51
H07 | IAC | Unknown | CY5 | 32.51 | H01 | Leg pneumoe | Unknown | FAM | 22.33
H07 | Leg pneumoe | Unknown | FAM | 23.62 | H02 | Leg pneumoe | Unknown | FAM | 22.82
H08 | IAC | Unknown | CY5 | 32.57 | H03 | Leg pneumoe | Unknown | FAM | 22.73
H08 | Leg pneumoe | Unknown | FAM | 23.5 | H04 | Leg pneumoe | Unknown | FAM | 23.52
H09 | IAC | Unknown | CY5 | 31.75 | H05 | Leg pneumoe | Unknown | FAM | 23.45
H09 | Leg pneumoe | Unknown | FAM | 23.5 | H06 | Leg pneumoe | Unknown | FAM | 23.84
H10 | IAC | Unknown | CY5 | 36.91 | H07 | Leg pneumoe | Unknown | FAM | 23.62
H10 | Leg pneumoe | Unknown | FAM | 23.25 | H08 | Leg pneumoe | Unknown | FAM | 23.5
H11 | IAC | Unknown | CY5 | 38.57 | H09 | Leg pneumoe | Unknown | FAM | 23.5
H11 | Leg pneumoe | Unknown | FAM | 23.15 | H10 | Leg pneumoe | Unknown | FAM | 23.25
H12 | IAC | Unknown | CY5 |  | H11 | Leg pneumoe | Unknown | FAM | 23.15
H12 | Leg pneumoe | Unknown | FAM | 23.88 | H12 | Leg pneumoe | Unknown | FAM | 23.88
Note on AC75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: confirmed that beginning concentration = 100uM.
CONCLUSIONS
Previously conclued in F2_matrix that optimized F, R primer concentrations would be ~400nM. However, this primer/probe data provides additional data suggestion that 300nM for primer concentrations may be more appropriate coupled with probe concentration between 200-400nM.
When used with Leg pneumo primers and probe, optimized IAC:
[F]:300nM
[R]:300nM
[P]:250nM
Also, remember that the error bars are 95% CI so overlap between neighboring values informative.
The chart above plots averaged Cq values when probe is held constant at 50, 100, 200nM etc while forward and reverse primers are allowed to vary between 500, 400 and 300nM. Standard deviation (stdev) is expected to be large as dramatic swings in F, R concentration occur. A reduction in stdev may reflect a saturation of probe. Another way of saying, at these three concentrations of F, R primer, the probe conc doesn't much matter. Does this reflect stability? Will need to keep thinking about this one. May also not be relevant figure.
Not a lot of variability in Leg pneumo amplificaiton when IAC primer and probe allowed to vary. Data supports no interplay between Leg pneumo F and IAC F2 or R.

## sample_test_IAC_opt
Now that I have optimized for Leg pneumo and IAC in the presence of Leg pneumo, I can test final optimized concentrations and see if they can amplify Leg pneumo samples
Mastermix ordered from Raquel
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 800
Stock Concentrations |  |  | Overage = |  |  | 5 | %
dNTP Mix | 25 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Component | Desired |  | Per 20 uL |  | Multiple Reactions
dNTP Mix | 0.25 | mM | 0.32 | uL | 0.32 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 0.05 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 0.07 | uL
Trehalose | 10 | Percent | 4 | uL | 4 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 0.2 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 2 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 2.2222 | uL
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 8.8622 | uL | 8.8622 | uL
Component | Desired | Units | PROBE RXN | 115.2 |  | Tube | TubeName | Conc (uM) | pM | Copy Number | Copy Number/0.2ul | Copy Number/2ul | Log(Copy#)
mmix |  |  | 8.8622 | 1020.928 |  |  | IDT Source | 100 | 100000000
Leg pneumo F | 400nM |  | 0.8 | 92.16 |  | Tube 2 |  | 0.01 | 10000 | 6.022e+12 | 12044000000 | 1204400000 | 10.0808
Leg pneumo R | 400nM |  | 0.8 | 92.16 |  | Tube 3 |  | 0.001 | 1000 | 602200000000.0002 | 1204400000 | 120440000 | 9.0808
Leg pneumo probe | 500nM |  | 1 | 115.2 |  | Tube 4 |  | 0.0001 | 100 | 60220000000 | 120440000 | 12044000 | 8.0808
IAC F2 | 300nM |  | 0.6 | 69.12 |  | Tube 5 |  | 1e-05 | 10 | 6022000000 | 12044000 | 1204400 | 7.0808
IAC R | 300nM |  | 0.6 | 69.12 |  | Tube 6 |  | 1e-06 | 1 | 602200000 | 1204400 | 120440 | 6.0808
IAC probe | 250nM |  | 0.5 | 57.6 |  | Tube 7 |  | 1e-07 | 0.1 | 60220000 | 120440 | 12044 | 5.0808
IAC std_6 |  |  | 0.2 | 23.04 |  | Tube 8 |  | 1e-08 | 0.01 | 6022000 | 12044 | 1204.4 | 4.0808
water |  |  | 4.6378 | 534.272
TOTAL |  |  | 18 | 2073.6
Used 6 Leg pneumo standards.
Some samples positive when should have been negative. NTCs also positive = contamination. Water?
Not much difference observed between samples.
IAC_no_dextran
IAC_with_dextran
some depression in fluorescence observed in samples containing 3% dextran but Cq appear unaffected.
Note on A34: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: In every 0.2ul added from std_6 to 20ul reaction, there are 100,000 copies. Expected Cq =33.3 +/- 0.9

## lyophilized
Have Raquel make master mix:
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 500
Stock Concentrations |  |  | Overage = |  |  | 5%
dNTP Mix | 10 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Dextran | 25 | Percent
Leg pneumo F | 100 | uM
Leg pneumo R | 100 | uM
Leg pneumo P | 100 | uM
IAC_F2 | 100 | uM
IAC_R | 100 | uM
IAC_P | 100 | uM
Component | Desired |  | Per 20 uL |  | Multiple Reactions
dNTP Mix | 0.25 | mM | 0.8 | uL | 420 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 26.25 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 36.75 | uL
Trehalose | 10 | Percent | 4 | uL | 2100 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 105 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 1050 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 1166.6667 | uL
dextran | 3 | Percent | 2.4 | uL | 1260 | uL
Leg pneumo F | 0.4 | uM | 0.08 | uL | 42 | uL
Leg pneumo R | 0.4 | uM | 0.08 | uL | 42 | uL
Leg pneumo P | 0.5 | uM | 0.1 | uL | 52.5 | uL
IAC_F2 | 0.3 | uM | 0 | uL | 0 | uL
IAC_R | 0.3 | uM | 0.06 | uL | 31.5 | uL
IAC_P | 0.25 | uM | 0.05 | uL | 26.25 | uL
IAC Ultramer | 0 |  | 0 | uL | 0 | uL
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 12.1122 | uL | 6358.9167 | uL
Procedure
1) Add reagents together according to the above schedule
2) Aliquot 12.11ul into each well of an 8-well strip tube. There should be 24, 8-well strip tubes for a total of 192 wells.
3) The remaining volume (6358.92-(12.11*192) = 4033.8ul) should be place in a 5mL conical tube and sent downstairs.
4) The 24 strip cap tubes should be placed in lyophilizer and processed. When finished, they may be sent downstairs.
I have the IAC F2 and the IAC ultramer down here. I'll add them in when I want to use the mix.
using program in Exp800.14.
Need to add IAC_F2 and Ultramer to water solution so that every well can receive.
C1 | V1 | C2 | V2
10 | 0.6 | 0.3 | 20
lyophilized mastermix tube containing everything but DNA.
Component | Desired | Units | PROBE RXN | 60
mmix |  |  | 0 | 0
sample |  |  | 2 | 120
IAC_F2 | 0.3uM |  | 0.6 | 36
IAC ultramer "std_7" |  |  | 0.2 | 12
water |  |  | 17.2 | 1032
TOTAL |  |  | 20 | 1200
307.2
Q96 | [5 empty] | Q16 | [6 empty] | miniS
Well | Target | Property | Dye | Ct |  | Well | Sample | Type | Target | Ct | Result |  | Well | Target | Property | Dye | Ct
C03 | Internal Control | Unknown | CY5 | 41.2 |  | A1-1 | A1 | Unknown | Legionella sp. | 11.87 | Detected |  | A01 | Internal Control | Unknown | cy5
C03 | Leg pneumo | Unknown | FAM | 11.82 |  | A1-4 | A1 | Unknown | Internal Control | 36.67 | Pass |  | A01 | Legionella sp | Unknown | FAM | 11.21
C04 | Internal Control | Unknown | CY5 | 39.5 |  | A2-1 | A2 | Unknown | Legionella sp. | 23.84 | Detected |  | A02 | Internal Control | Unknown | cy5 | 39.56
C04 | Leg pneumo | Unknown | FAM | 22.77 |  | A2-4 | A2 | Unknown | Internal Control | 36.84 | Pass |  | A02 | Legionella sp | Unknown | FAM | 22.99
C05 | Internal Control | Unknown | CY5 | 44.12 |  | A3-1 | A3 | Unknown | Legionella sp. | 22.41 | Detected |  | A03 | Internal Control | Unknown | cy5 | 38.28
C05 | Leg pneumo | Unknown | FAM | 21.77 |  | A3-4 | A3 | Unknown | Internal Control | 0 | Fail |  | A03 | Legionella sp | Unknown | FAM | 21.59
C06 | Internal Control | Unknown | CY5 | 38.63 |  | A4-1 | A4 | Unknown | Legionella sp. | 27.94 | Detected |  | A04 | Internal Control | Unknown | cy5 | 36.61
C06 | Leg pneumo | Unknown | FAM | 29.1 |  | A4-4 | A4 | Unknown | Internal Control | 32.64 | Pass |  | A04 | Legionella sp | Unknown | FAM | 28.9
C07 | Internal Control | Unknown | CY5 | 35.92 |  | A5-1 | A5 | Unknown | Legionella sp. | 28.91 | Detected |  | A05 | Internal Control | Unknown | cy5 | 38.4
C07 | Leg pneumo | Unknown | FAM | 28.37 |  | A5-4 | A5 | Unknown | Internal Control | 34.19 | Pass |  | A05 | Legionella sp | Unknown | FAM | 28.62
C08 | Internal Control | Unknown | CY5 | 36.95 |  | A6-1 | A6 | Unknown | Legionella sp. | 30.49 | Detected |  | A06 | Internal Control | Unknown | cy5 | 37.64
C08 | Leg pneumo | Unknown | FAM | 29.07 |  | A6-4 | A6 | Unknown | Internal Control | 33.51 | Pass |  | A06 | Legionella sp | Unknown | FAM | 29.74
C09 | Internal Control | Unknown | CY5 | 34.86 |  | A7-1 | A7 | Unknown | Legionella sp. | 31.67 | Detected |  | A07 | Internal Control | Unknown | cy5 | 36.32
C09 | Leg pneumo | Unknown | FAM | 32.15 |  | A7-4 | A7 | Unknown | Internal Control | 34.57 | Pass |  | A07 | Legionella sp | Unknown | FAM | 29.49
C10 | Internal Control | Unknown | CY5 |  |  | A8-1 | A8 | Unknown | Legionella sp. | 0 | Undetected |  | A08 | Internal Control | Unknown | cy5 | 34.97
C10 | Leg pneumo | Unknown | FAM |  |  | A8-4 | A8 | Unknown | Internal Control | 35.52 | Pass |  | A08 | Legionella sp | Unknown | FAM | 18.37
D03 | Internal Control | Unknown | CY5 | 40.42 |  | B1-1 | B1 | Unknown | Legionella sp. | 11.76 | Detected |  | B01 | Internal Control | Unknown | cy5 | 38.93
D03 | Leg pneumo | Unknown | FAM | 13.15 |  | B1-4 | B1 | Unknown | Internal Control | 34.31 | Pass |  | B01 | Legionella sp | Unknown | FAM | 11.35
D04 | Internal Control | Unknown | CY5 | 36.69 |  | B2-1 | B2 | Unknown | Legionella sp. | 23.92 | Detected |  | B02 | Internal Control | Unknown | cy5 | 38.75
D04 | Leg pneumo | Unknown | FAM | 22.38 |  | B2-4 | B2 | Unknown | Internal Control | 35.54 | Pass |  | B02 | Legionella sp | Unknown | FAM | 22.88
D05 | Internal Control | Unknown | CY5 | 38.8 |  | B3-1 | B3 | Unknown | Legionella sp. | 22.49 | Detected |  | B03 | Internal Control | Unknown | cy5 | 39.96
D05 | Leg pneumo | Unknown | FAM | 21.31 |  | B3-4 | B3 | Unknown | Internal Control | 36 | Pass |  | B03 | Legionella sp | Unknown | FAM | 22.56
D06 | Internal Control | Unknown | CY5 | 38.85 |  | B4-1 | B4 | Unknown | Legionella sp. | 29.09 | Detected |  | B04 | Internal Control | Unknown | cy5 | 36.88
D06 | Leg pneumo | Unknown | FAM | 28.54 |  | B4-4 | B4 | Unknown | Internal Control | 33.95 | Pass |  | B04 | Legionella sp | Unknown | FAM | 29.45
D07 | Internal Control | Unknown | CY5 |  |  | B5-1 | B5 | Unknown | Legionella sp. | 29.21 | Detected |  | B05 | Internal Control | Unknown | cy5 | 37.39
D07 | Leg pneumo | Unknown | FAM | 28.08 |  | B5-4 | B5 | Unknown | Internal Control | 33.11 | Pass |  | B05 | Legionella sp | Unknown | FAM | 28.75
D08 | Internal Control | Unknown | CY5 | 36.42 |  | B6-1 | B6 | Unknown | Legionella sp. | 30.26 | Detected |  | B06 | Internal Control | Unknown | cy5 | 38.6
D08 | Leg pneumo | Unknown | FAM | 29.27 |  | B6-4 | B6 | Unknown | Internal Control | 0 | Fail |  | B06 | Legionella sp | Unknown | FAM | 29.39
D09 | Internal Control | Unknown | CY5 | 38.9 |  | B7-1 | B7 | Unknown | Legionella sp. | 32.27 | Detected |  | B07 | Internal Control | Unknown | cy5 | 34.31
D09 | Leg pneumo | Unknown | FAM | 32.53 |  | B7-4 | B7 | Unknown | Internal Control | 0 | Fail |  | B07 | Legionella sp | Unknown | FAM | 31.51
D10 | Internal Control | Unknown | CY5 |  |  | B8-1 | B8 | Unknown | Legionella sp. | 0 | Undetected |  | B08 | Internal Control | Unknown | cy5 | 34.63
D10 | Leg pneumo | Unknown | FAM | 37.96 |  | B8-4 | B8 | Unknown | Internal Control | 32.45 | Pass |  | B08 | Legionella sp | Unknown | FAM
ThermalCycler | Well | Well_Q96 | Well_Q16 | Sample_Q16 | Type_Q16 | Target | Property | Dye | Ct |  | ThermalCycler | Well | Well_Q96 | Well_Q16 | Sample_Q16 | Type_Q16 | Target | Property | Dye | Ct | ThermalCycler | Well | Well_Q96 | Well_Q16 | Sample_Q16 | Type_Q16 | Target | Property | Dye | Ct
Q96 | A01 | C03 |  |  |  | Internal Control | Unknown | CY5 | 41.2 |  | Q96 | A08 | C10 |  |  |  | Internal Control | Unknown | CY5 | 0 | miniS | 1 |  |  |  |  | Internal Control | Unknown | CY5 | 0
Q96 | A01 | C03 |  |  |  | Leg pneumo | Unknown | FAM | 11.82 |  | Q96 | A08 | C10 |  |  |  | Leg pneumo | Unknown | FAM | 0 | miniS | 1 |  |  |  |  | Legionella sp | Unknown | FAM | 11.21
Q96 | A02 | C04 |  |  |  | Internal Control | Unknown | CY5 | 39.5 |  | Q96 | B05 | D07 |  |  |  | Internal Control | Unknown | CY5 | 0 | miniS | 2 |  |  |  |  | Legionella sp | Unknown | FAM | 22.99
Q96 | A02 | C04 |  |  |  | Leg pneumo | Unknown | FAM | 22.77 |  | Q96 | B08 | D10 |  |  |  | Internal Control | Unknown | CY5 | 0 | miniS | 2 |  |  |  |  | Internal Control | Unknown | CY5 | 39.56
Q96 | A03 | C05 |  |  |  | Internal Control | Unknown | CY5 | 44.12 |  | Q16 | A03 |  | A3-4 | A3 | Unknown | Internal Control |  | CY5 | 0 | miniS | 3 |  |  |  |  | Legionella sp | Unknown | FAM | 21.59
Q96 | A03 | C05 |  |  |  | Leg pneumo | Unknown | FAM | 21.77 |  | Q16 | A08 |  | A8-1 | A8 | Unknown | Legionella sp. |  | FAM | 0 | miniS | 3 |  |  |  |  | Internal Control | Unknown | CY5 | 38.28
Q96 | A04 | C06 |  |  |  | Internal Control | Unknown | CY5 | 38.63 |  | Q16 | B06 |  | B6-4 | B6 | Unknown | Internal Control |  | CY5 | 0 | miniS | 4 |  |  |  |  | Legionella sp | Unknown | FAM | 28.9
Q96 | A04 | C06 |  |  |  | Leg pneumo | Unknown | FAM | 29.1 |  | Q16 | B07 |  | B7-4 | B7 | Unknown | Internal Control |  | CY5 | 0 | miniS | 4 |  |  |  |  | Internal Control | Unknown | CY5 | 36.61
Q96 | A05 | C07 |  |  |  | Internal Control | Unknown | CY5 | 35.92 |  | Q16 | B08 |  | B8-1 | B8 | Unknown | Legionella sp. |  | FAM | 0 | miniS | 5 |  |  |  |  | Legionella sp | Unknown | FAM | 28.62
Q96 | A05 | C07 |  |  |  | Leg pneumo | Unknown | FAM | 28.37 |  | miniS | A01 |  |  |  |  | Internal Control | Unknown | CY5 | 0 | miniS | 5 |  |  |  |  | Internal Control | Unknown | CY5 | 38.4
Q96 | A06 | C08 |  |  |  | Internal Control | Unknown | CY5 | 36.95 |  | miniS | B08 |  |  |  |  | Legionella sp | Unknown | FAM | 0 | miniS | 6 |  |  |  |  | Legionella sp | Unknown | FAM | 29.74
Q96 | A06 | C08 |  |  |  | Leg pneumo | Unknown | FAM | 29.07 |  | miniS | A01 |  |  |  |  | Legionella sp | Unknown | FAM | 11.21 | miniS | 6 |  |  |  |  | Internal Control | Unknown | CY5 | 37.64
Q96 | A07 | C09 |  |  |  | Internal Control | Unknown | CY5 | 34.86 |  | miniS | B01 |  |  |  |  | Legionella sp | Unknown | FAM | 11.35 | miniS | 7 |  |  |  |  | Legionella sp | Unknown | FAM | 29.49
Q96 | A07 | C09 |  |  |  | Leg pneumo | Unknown | FAM | 32.15 |  | Q16 | B01 |  | B1-1 | B1 | Unknown | Legionella sp. |  | FAM | 11.76 | miniS | 7 |  |  |  |  | Internal Control | Unknown | CY5 | 36.32
Q96 | A08 | C10 |  |  |  | Internal Control | Unknown | CY5 | 0 |  | Q96 | A01 | C03 |  |  |  | Leg pneumo | Unknown | FAM | 11.82 | miniS | 8 |  |  |  |  | Legionella sp | Unknown | FAM | 18.37
Q96 | A08 | C10 |  |  |  | Leg pneumo | Unknown | FAM | 0 |  | Q16 | A01 |  | A1-1 | A1 | Unknown | Legionella sp. |  | FAM | 11.87 | miniS | 8 |  |  |  |  | Internal Control | Unknown | CY5 | 34.97
Q96 | B01 | D03 |  |  |  | Internal Control | Unknown | CY5 | 40.42 |  | Q96 | B01 | D03 |  |  |  | Leg pneumo | Unknown | FAM | 13.15 | miniS | 1 |  |  |  |  | Legionella sp | Unknown | FAM | 11.35
Q96 | B01 | D03 |  |  |  | Leg pneumo | Unknown | FAM | 13.15 |  | miniS | A08 |  |  |  |  | Legionella sp | Unknown | FAM | 18.37 | miniS | 1 |  |  |  |  | Internal Control | Unknown | CY5 | 38.93
Q96 | B02 | D04 |  |  |  | Internal Control | Unknown | CY5 | 36.69 |  | Q96 | B03 | D05 |  |  |  | Leg pneumo | Unknown | FAM | 21.31 | miniS | 2 |  |  |  |  | Legionella sp | Unknown | FAM | 22.88
Q96 | B02 | D04 |  |  |  | Leg pneumo | Unknown | FAM | 22.38 |  | miniS | A03 |  |  |  |  | Legionella sp | Unknown | FAM | 21.59 | miniS | 2 |  |  |  |  | Internal Control | Unknown | CY5 | 38.75
Q96 | B03 | D05 |  |  |  | Internal Control | Unknown | CY5 | 38.8 |  | Q96 | A03 | C05 |  |  |  | Leg pneumo | Unknown | FAM | 21.77 | miniS | 3 |  |  |  |  | Legionella sp | Unknown | FAM | 22.56
Q96 | B03 | D05 |  |  |  | Leg pneumo | Unknown | FAM | 21.31 |  | Q96 | B02 | D04 |  |  |  | Leg pneumo | Unknown | FAM | 22.38 | miniS | 3 |  |  |  |  | Internal Control | Unknown | CY5 | 39.96
Q96 | B04 | D06 |  |  |  | Internal Control | Unknown | CY5 | 38.85 |  | Q16 | A03 |  | A3-1 | A3 | Unknown | Legionella sp. |  | FAM | 22.41 | miniS | 4 |  |  |  |  | Legionella sp | Unknown | FAM | 29.45
Q96 | B04 | D06 |  |  |  | Leg pneumo | Unknown | FAM | 28.54 |  | Q16 | B03 |  | B3-1 | B3 | Unknown | Legionella sp. |  | FAM | 22.49 | miniS | 4 |  |  |  |  | Internal Control | Unknown | CY5 | 36.88
Q96 | B05 | D07 |  |  |  | Internal Control | Unknown | CY5 | 0 |  | miniS | B03 |  |  |  |  | Legionella sp | Unknown | FAM | 22.56 | miniS | 5 |  |  |  |  | Legionella sp | Unknown | FAM | 28.75
Q96 | B05 | D07 |  |  |  | Leg pneumo | Unknown | FAM | 28.08 |  | Q96 | A02 | C04 |  |  |  | Leg pneumo | Unknown | FAM | 22.77 | miniS | 5 |  |  |  |  | Internal Control | Unknown | CY5 | 37.39
Q96 | B06 | D08 |  |  |  | Internal Control | Unknown | CY5 | 36.42 |  | miniS | B02 |  |  |  |  | Legionella sp | Unknown | FAM | 22.88 | miniS | 6 |  |  |  |  | Legionella sp | Unknown | FAM | 29.39
Q96 | B06 | D08 |  |  |  | Leg pneumo | Unknown | FAM | 29.27 |  | miniS | A02 |  |  |  |  | Legionella sp | Unknown | FAM | 22.99 | miniS | 6 |  |  |  |  | Internal Control | Unknown | CY5 | 38.6
Q96 | B07 | D09 |  |  |  | Internal Control | Unknown | CY5 | 38.9 |  | Q16 | A02 |  | A2-1 | A2 | Unknown | Legionella sp. |  | FAM | 23.84 | miniS | 7 |  |  |  |  | Legionella sp | Unknown | FAM | 31.51
Q96 | B07 | D09 |  |  |  | Leg pneumo | Unknown | FAM | 32.53 |  | Q16 | B02 |  | B2-1 | B2 | Unknown | Legionella sp. |  | FAM | 23.92 | miniS | 7 |  |  |  |  | Internal Control | Unknown | CY5 | 34.31
Q96 | B08 | D10 |  |  |  | Internal Control | Unknown | CY5 | 0 |  | Q16 | A04 |  | A4-1 | A4 | Unknown | Legionella sp. |  | FAM | 27.94 | miniS | 8 |  |  |  |  | Legionella sp | Unknown | FAM | 0
Q96 | B08 | D10 |  |  |  | Leg pneumo | Unknown | FAM | 37.96 |  | Q96 | B05 | D07 |  |  |  | Leg pneumo | Unknown | FAM | 28.08 | miniS | 8 |  |  |  |  | Internal Control | Unknown | CY5 | 34.63
Q16 | A01 |  | A1-1 | A1 | Unknown | Legionella sp. |  | FAM | 11.87 |  | Q96 | A05 | C07 |  |  |  | Leg pneumo | Unknown | FAM | 28.37 | Q16 | 1 |  | A1-1 | A1 | Unknown | Legionella sp. |  | FAM | 11.87
Q16 | A01 |  | A1-4 | A1 | Unknown | Internal Control |  | CY5 | 36.67 |  | Q96 | B04 | D06 |  |  |  | Leg pneumo | Unknown | FAM | 28.54 | Q16 | 1 |  | A1-4 | A1 | Unknown | Internal Control |  | CY5 | 36.67
Q16 | A02 |  | A2-1 | A2 | Unknown | Legionella sp. |  | FAM | 23.84 |  | miniS | A05 |  |  |  |  | Legionella sp | Unknown | FAM | 28.62 | Q16 | 2 |  | A2-1 | A2 | Unknown | Legionella sp. |  | FAM | 23.84
Q16 | A02 |  | A2-4 | A2 | Unknown | Internal Control |  | CY5 | 36.84 |  | miniS | B05 |  |  |  |  | Legionella sp | Unknown | FAM | 28.75 | Q16 | 2 |  | A2-4 | A2 | Unknown | Internal Control |  | CY5 | 36.84
Q16 | A03 |  | A3-1 | A3 | Unknown | Legionella sp. |  | FAM | 22.41 |  | miniS | A04 |  |  |  |  | Legionella sp | Unknown | FAM | 28.9 | Q16 | 3 |  | A3-4 | A3 | Unknown | Internal Control |  | CY5 | 0
Q16 | A03 |  | A3-4 | A3 | Unknown | Internal Control |  | CY5 | 0 |  | Q16 | A05 |  | A5-1 | A5 | Unknown | Legionella sp. |  | FAM | 28.91 | Q16 | 3 |  | A3-1 | A3 | Unknown | Legionella sp. |  | FAM | 22.41
Q16 | A04 |  | A4-1 | A4 | Unknown | Legionella sp. |  | FAM | 27.94 |  | Q96 | A06 | C08 |  |  |  | Leg pneumo | Unknown | FAM | 29.07 | Q16 | 4 |  | A4-1 | A4 | Unknown | Legionella sp. |  | FAM | 27.94
Q16 | A04 |  | A4-4 | A4 | Unknown | Internal Control |  | CY5 | 32.64 |  | Q16 | B04 |  | B4-1 | B4 | Unknown | Legionella sp. |  | FAM | 29.09 | Q16 | 4 |  | A4-4 | A4 | Unknown | Internal Control |  | CY5 | 32.64
Q16 | A05 |  | A5-1 | A5 | Unknown | Legionella sp. |  | FAM | 28.91 |  | Q96 | A04 | C06 |  |  |  | Leg pneumo | Unknown | FAM | 29.1 | Q16 | 5 |  | A5-1 | A5 | Unknown | Legionella sp. |  | FAM | 28.91
Q16 | A05 |  | A5-4 | A5 | Unknown | Internal Control |  | CY5 | 34.19 |  | Q16 | B05 |  | B5-1 | B5 | Unknown | Legionella sp. |  | FAM | 29.21 | Q16 | 5 |  | A5-4 | A5 | Unknown | Internal Control |  | CY5 | 34.19
Q16 | A06 |  | A6-1 | A6 | Unknown | Legionella sp. |  | FAM | 30.49 |  | Q96 | B06 | D08 |  |  |  | Leg pneumo | Unknown | FAM | 29.27 | Q16 | 6 |  | A6-1 | A6 | Unknown | Legionella sp. |  | FAM | 30.49
Q16 | A06 |  | A6-4 | A6 | Unknown | Internal Control |  | CY5 | 33.51 |  | miniS | B06 |  |  |  |  | Legionella sp | Unknown | FAM | 29.39 | Q16 | 6 |  | A6-4 | A6 | Unknown | Internal Control |  | CY5 | 33.51
Q16 | A07 |  | A7-1 | A7 | Unknown | Legionella sp. |  | FAM | 31.67 |  | miniS | B04 |  |  |  |  | Legionella sp | Unknown | FAM | 29.45 | Q16 | 7 |  | A7-1 | A7 | Unknown | Legionella sp. |  | FAM | 31.67
Q16 | A07 |  | A7-4 | A7 | Unknown | Internal Control |  | CY5 | 34.57 |  | miniS | A07 |  |  |  |  | Legionella sp | Unknown | FAM | 29.49 | Q16 | 7 |  | A7-4 | A7 | Unknown | Internal Control |  | CY5 | 34.57
Q16 | A08 |  | A8-1 | A8 | Unknown | Legionella sp. |  | FAM | 0 |  | miniS | A06 |  |  |  |  | Legionella sp | Unknown | FAM | 29.74 | Q16 | 8 |  | A8-1 | A8 | Unknown | Legionella sp. |  | FAM | 0
Q16 | A08 |  | A8-4 | A8 | Unknown | Internal Control |  | CY5 | 35.52 |  | Q16 | B06 |  | B6-1 | B6 | Unknown | Legionella sp. |  | FAM | 30.26 | Q16 | 8 |  | A8-4 | A8 | Unknown | Internal Control |  | CY5 | 35.52
Q16 | B01 |  | B1-1 | B1 | Unknown | Legionella sp. |  | FAM | 11.76 |  | Q16 | A06 |  | A6-1 | A6 | Unknown | Legionella sp. |  | FAM | 30.49 | Q16 | 1 |  | B1-1 | B1 | Unknown | Legionella sp. |  | FAM | 11.76
Q16 | B01 |  | B1-4 | B1 | Unknown | Internal Control |  | CY5 | 34.31 |  | miniS | B07 |  |  |  |  | Legionella sp | Unknown | FAM | 31.51 | Q16 | 1 |  | B1-4 | B1 | Unknown | Internal Control |  | CY5 | 34.31
Q16 | B02 |  | B2-1 | B2 | Unknown | Legionella sp. |  | FAM | 23.92 |  | Q16 | A07 |  | A7-1 | A7 | Unknown | Legionella sp. |  | FAM | 31.67 | Q16 | 2 |  | B2-1 | B2 | Unknown | Legionella sp. |  | FAM | 23.92
Q16 | B02 |  | B2-4 | B2 | Unknown | Internal Control |  | CY5 | 35.54 |  | Q96 | A07 | C09 |  |  |  | Leg pneumo | Unknown | FAM | 32.15 | Q16 | 2 |  | B2-4 | B2 | Unknown | Internal Control |  | CY5 | 35.54
Q16 | B03 |  | B3-1 | B3 | Unknown | Legionella sp. |  | FAM | 22.49 |  | Q16 | B07 |  | B7-1 | B7 | Unknown | Legionella sp. |  | FAM | 32.27 | Q16 | 3 |  | B3-1 | B3 | Unknown | Legionella sp. |  | FAM | 22.49
Q16 | B03 |  | B3-4 | B3 | Unknown | Internal Control |  | CY5 | 36 |  | Q16 | B08 |  | B8-4 | B8 | Unknown | Internal Control |  | CY5 | 32.45 | Q16 | 3 |  | B3-4 | B3 | Unknown | Internal Control |  | CY5 | 36
Q16 | B04 |  | B4-1 | B4 | Unknown | Legionella sp. |  | FAM | 29.09 |  | Q96 | B07 | D09 |  |  |  | Leg pneumo | Unknown | FAM | 32.53 | Q16 | 4 |  | B4-1 | B4 | Unknown | Legionella sp. |  | FAM | 29.09
Q16 | B04 |  | B4-4 | B4 | Unknown | Internal Control |  | CY5 | 33.95 |  | Q16 | A04 |  | A4-4 | A4 | Unknown | Internal Control |  | CY5 | 32.64 | Q16 | 4 |  | B4-4 | B4 | Unknown | Internal Control |  | CY5 | 33.95
Q16 | B05 |  | B5-1 | B5 | Unknown | Legionella sp. |  | FAM | 29.21 |  | Q16 | B05 |  | B5-4 | B5 | Unknown | Internal Control |  | CY5 | 33.11 | Q16 | 5 |  | B5-1 | B5 | Unknown | Legionella sp. |  | FAM | 29.21
Q16 | B05 |  | B5-4 | B5 | Unknown | Internal Control |  | CY5 | 33.11 |  | Q16 | A06 |  | A6-4 | A6 | Unknown | Internal Control |  | CY5 | 33.51 | Q16 | 5 |  | B5-4 | B5 | Unknown | Internal Control |  | CY5 | 33.11
Q16 | B06 |  | B6-1 | B6 | Unknown | Legionella sp. |  | FAM | 30.26 |  | Q16 | B04 |  | B4-4 | B4 | Unknown | Internal Control |  | CY5 | 33.95 | Q16 | 6 |  | B6-4 | B6 | Unknown | Internal Control |  | CY5 | 0
Q16 | B06 |  | B6-4 | B6 | Unknown | Internal Control |  | CY5 | 0 |  | Q16 | A05 |  | A5-4 | A5 | Unknown | Internal Control |  | CY5 | 34.19 | Q16 | 6 |  | B6-1 | B6 | Unknown | Legionella sp. |  | FAM | 30.26
Q16 | B07 |  | B7-1 | B7 | Unknown | Legionella sp. |  | FAM | 32.27 |  | Q16 | B01 |  | B1-4 | B1 | Unknown | Internal Control |  | CY5 | 34.31 | Q16 | 7 |  | B7-4 | B7 | Unknown | Internal Control |  | CY5 | 0
Q16 | B07 |  | B7-4 | B7 | Unknown | Internal Control |  | CY5 | 0 |  | miniS | B07 |  |  |  |  | Internal Control | Unknown | CY5 | 34.31 | Q16 | 7 |  | B7-1 | B7 | Unknown | Legionella sp. |  | FAM | 32.27
Q16 | B08 |  | B8-1 | B8 | Unknown | Legionella sp. |  | FAM | 0 |  | Q16 | A07 |  | A7-4 | A7 | Unknown | Internal Control |  | CY5 | 34.57 | Q16 | 8 |  | B8-1 | B8 | Unknown | Legionella sp. |  | FAM | 0
Q16 | B08 |  | B8-4 | B8 | Unknown | Internal Control |  | CY5 | 32.45 |  | miniS | B08 |  |  |  |  | Internal Control | Unknown | CY5 | 34.63 | Q16 | 8 |  | B8-4 | B8 | Unknown | Internal Control |  | CY5 | 32.45
miniS | A01 |  |  |  |  | Internal Control | Unknown | CY5 | 0 |  | Q96 | A07 | C09 |  |  |  | Internal Control | Unknown | CY5 | 34.86 | Q96 | 1 | C03 |  |  |  | Leg pneumo | Unknown | FAM | 11.82
miniS | A01 |  |  |  |  | Legionella sp | Unknown | FAM | 11.21 |  | miniS | A08 |  |  |  |  | Internal Control | Unknown | CY5 | 34.97 | Q96 | 1 | C03 |  |  |  | Internal Control | Unknown | CY5 | 41.2
miniS | A02 |  |  |  |  | Internal Control | Unknown | CY5 | 39.56 |  | Q16 | A08 |  | A8-4 | A8 | Unknown | Internal Control |  | CY5 | 35.52 | Q96 | 2 | C04 |  |  |  | Leg pneumo | Unknown | FAM | 22.77
miniS | A02 |  |  |  |  | Legionella sp | Unknown | FAM | 22.99 |  | Q16 | B02 |  | B2-4 | B2 | Unknown | Internal Control |  | CY5 | 35.54 | Q96 | 2 | C04 |  |  |  | Internal Control | Unknown | CY5 | 39.5
miniS | A03 |  |  |  |  | Internal Control | Unknown | CY5 | 38.28 |  | Q96 | A05 | C07 |  |  |  | Internal Control | Unknown | CY5 | 35.92 | Q96 | 3 | C05 |  |  |  | Leg pneumo | Unknown | FAM | 21.77
miniS | A03 |  |  |  |  | Legionella sp | Unknown | FAM | 21.59 |  | Q16 | B03 |  | B3-4 | B3 | Unknown | Internal Control |  | CY5 | 36 | Q96 | 3 | C05 |  |  |  | Internal Control | Unknown | CY5 | 44.12
miniS | A04 |  |  |  |  | Internal Control | Unknown | CY5 | 36.61 |  | miniS | A07 |  |  |  |  | Internal Control | Unknown | CY5 | 36.32 | Q96 | 4 | C06 |  |  |  | Leg pneumo | Unknown | FAM | 29.1
miniS | A04 |  |  |  |  | Legionella sp | Unknown | FAM | 28.9 |  | Q96 | B06 | D08 |  |  |  | Internal Control | Unknown | CY5 | 36.42 | Q96 | 4 | C06 |  |  |  | Internal Control | Unknown | CY5 | 38.63
miniS | A05 |  |  |  |  | Internal Control | Unknown | CY5 | 38.4 |  | miniS | A04 |  |  |  |  | Internal Control | Unknown | CY5 | 36.61 | Q96 | 5 | C07 |  |  |  | Leg pneumo | Unknown | FAM | 28.37
miniS | A05 |  |  |  |  | Legionella sp | Unknown | FAM | 28.62 |  | Q16 | A01 |  | A1-4 | A1 | Unknown | Internal Control |  | CY5 | 36.67 | Q96 | 5 | C07 |  |  |  | Internal Control | Unknown | CY5 | 35.92
miniS | A06 |  |  |  |  | Internal Control | Unknown | CY5 | 37.64 |  | Q96 | B02 | D04 |  |  |  | Internal Control | Unknown | CY5 | 36.69 | Q96 | 6 | C08 |  |  |  | Leg pneumo | Unknown | FAM | 29.07
miniS | A06 |  |  |  |  | Legionella sp | Unknown | FAM | 29.74 |  | Q16 | A02 |  | A2-4 | A2 | Unknown | Internal Control |  | CY5 | 36.84 | Q96 | 6 | C08 |  |  |  | Internal Control | Unknown | CY5 | 36.95
miniS | A07 |  |  |  |  | Internal Control | Unknown | CY5 | 36.32 |  | miniS | B04 |  |  |  |  | Internal Control | Unknown | CY5 | 36.88 | Q96 | 7 | C09 |  |  |  | Leg pneumo | Unknown | FAM | 32.15
miniS | A07 |  |  |  |  | Legionella sp | Unknown | FAM | 29.49 |  | Q96 | A06 | C08 |  |  |  | Internal Control | Unknown | CY5 | 36.95 | Q96 | 7 | C09 |  |  |  | Internal Control | Unknown | CY5 | 34.86
miniS | A08 |  |  |  |  | Internal Control | Unknown | CY5 | 34.97 |  | miniS | B05 |  |  |  |  | Internal Control | Unknown | CY5 | 37.39 | Q96 | 8 | C10 |  |  |  | Internal Control | Unknown | CY5 | 0
miniS | A08 |  |  |  |  | Legionella sp | Unknown | FAM | 18.37 |  | miniS | A06 |  |  |  |  | Internal Control | Unknown | CY5 | 37.64 | Q96 | 8 | C10 |  |  |  | Leg pneumo | Unknown | FAM | 0
miniS | B01 |  |  |  |  | Internal Control | Unknown | CY5 | 38.93 |  | Q96 | B08 | D10 |  |  |  | Leg pneumo | Unknown | FAM | 37.96 | Q96 | 1 | D03 |  |  |  | Leg pneumo | Unknown | FAM | 13.15
miniS | B01 |  |  |  |  | Legionella sp | Unknown | FAM | 11.35 |  | miniS | A03 |  |  |  |  | Internal Control | Unknown | CY5 | 38.28 | Q96 | 1 | D03 |  |  |  | Internal Control | Unknown | CY5 | 40.42
miniS | B02 |  |  |  |  | Internal Control | Unknown | CY5 | 38.75 |  | miniS | A05 |  |  |  |  | Internal Control | Unknown | CY5 | 38.4 | Q96 | 2 | D04 |  |  |  | Leg pneumo | Unknown | FAM | 22.38
miniS | B02 |  |  |  |  | Legionella sp | Unknown | FAM | 22.88 |  | miniS | B06 |  |  |  |  | Internal Control | Unknown | CY5 | 38.6 | Q96 | 2 | D04 |  |  |  | Internal Control | Unknown | CY5 | 36.69
miniS | B03 |  |  |  |  | Internal Control | Unknown | CY5 | 39.96 |  | Q96 | A04 | C06 |  |  |  | Internal Control | Unknown | CY5 | 38.63 | Q96 | 3 | D05 |  |  |  | Leg pneumo | Unknown | FAM | 21.31
miniS | B03 |  |  |  |  | Legionella sp | Unknown | FAM | 22.56 |  | miniS | B02 |  |  |  |  | Internal Control | Unknown | CY5 | 38.75 | Q96 | 3 | D05 |  |  |  | Internal Control | Unknown | CY5 | 38.8
miniS | B04 |  |  |  |  | Internal Control | Unknown | CY5 | 36.88 |  | Q96 | B03 | D05 |  |  |  | Internal Control | Unknown | CY5 | 38.8 | Q96 | 4 | D06 |  |  |  | Leg pneumo | Unknown | FAM | 28.54
miniS | B04 |  |  |  |  | Legionella sp | Unknown | FAM | 29.45 |  | Q96 | B04 | D06 |  |  |  | Internal Control | Unknown | CY5 | 38.85 | Q96 | 4 | D06 |  |  |  | Internal Control | Unknown | CY5 | 38.85
miniS | B05 |  |  |  |  | Internal Control | Unknown | CY5 | 37.39 |  | Q96 | B07 | D09 |  |  |  | Internal Control | Unknown | CY5 | 38.9 | Q96 | 5 | D07 |  |  |  | Internal Control | Unknown | CY5 | 0
miniS | B05 |  |  |  |  | Legionella sp | Unknown | FAM | 28.75 |  | miniS | B01 |  |  |  |  | Internal Control | Unknown | CY5 | 38.93 | Q96 | 5 | D07 |  |  |  | Leg pneumo | Unknown | FAM | 28.08
miniS | B06 |  |  |  |  | Internal Control | Unknown | CY5 | 38.6 |  | Q96 | A02 | C04 |  |  |  | Internal Control | Unknown | CY5 | 39.5 | Q96 | 6 | D08 |  |  |  | Leg pneumo | Unknown | FAM | 29.27
miniS | B06 |  |  |  |  | Legionella sp | Unknown | FAM | 29.39 |  | miniS | A02 |  |  |  |  | Internal Control | Unknown | CY5 | 39.56 | Q96 | 6 | D08 |  |  |  | Internal Control | Unknown | CY5 | 36.42
miniS | B07 |  |  |  |  | Internal Control | Unknown | CY5 | 34.31 |  | miniS | B03 |  |  |  |  | Internal Control | Unknown | CY5 | 39.96 | Q96 | 7 | D09 |  |  |  | Leg pneumo | Unknown | FAM | 32.53
miniS | B07 |  |  |  |  | Legionella sp | Unknown | FAM | 31.51 |  | Q96 | B01 | D03 |  |  |  | Internal Control | Unknown | CY5 | 40.42 | Q96 | 7 | D09 |  |  |  | Internal Control | Unknown | CY5 | 38.9
miniS | B08 |  |  |  |  | Internal Control | Unknown | CY5 | 34.63 |  | Q96 | A01 | C03 |  |  |  | Internal Control | Unknown | CY5 | 41.2 | Q96 | 8 | D10 |  |  |  | Internal Control | Unknown | CY5 | 0
miniS | B08 |  |  |  |  | Legionella sp | Unknown | FAM | 0 |  | Q96 | A03 | C05 |  |  |  | Internal Control | Unknown | CY5 | 44.12 | Q96 | 8 | D10 |  |  |  | Leg pneumo | Unknown | FAM | 37.96
Average | [12 empty] | Std | Well | Dye | Copy Number/rxn (2ul) | log(copy#) | miniS | Q16 | Q96
ThermalCycler | Well | Dye | Ct | Ct_AVG |  |  | ThermalCycler | Well | Dye | Ct_AVG |  |  | std_2 | 1 | FAM | 1204400000 | 9.0808 | 11.28 | 11.815 | 12.485
miniS | A01 | CY5 |  | 38.93 |  |  | miniS | 1 | CY5 | 38.93 |  |  | std_3 | 2 | FAM | 120440000 | 8.0808 | 22.935 | 23.88 | 22.575
miniS | A02 | CY5 | 39.56 | 39.155 |  |  | miniS | 2 | CY5 | 39.155 |  |  | std_4 | 3 | FAM | 12044000 | 7.0808 | 22.075 | 22.45 | 21.54
miniS | A03 | CY5 | 38.28 | 39.12 |  |  | miniS | 3 | CY5 | 39.12 |  |  | std_5 | 4 | FAM | 1204400 | 6.0808 | 29.175 | 28.515 | 28.82
miniS | A04 | CY5 | 36.61 | 36.745 |  |  | miniS | 4 | CY5 | 36.745 |  |  | std_6 | 5 | FAM | 120440 | 5.0808 | 28.685 | 29.06 | 28.225
miniS | A05 | CY5 | 38.4 | 37.895 |  |  | miniS | 5 | CY5 | 37.895 |  |  | std_7 | 6 | FAM | 12044 | 4.0808 | 29.565 | 30.375 | 29.17
miniS | A06 | CY5 | 37.64 | 38.12 |  |  | miniS | 6 | CY5 | 38.12 |  |  | std_8 | 7 | FAM | 1204.4 | 3.0808 | 30.5 | 31.97 | 32.34
miniS | A07 | CY5 | 36.32 | 35.315 |  |  | miniS | 7 | CY5 | 35.315 |  |  | NTC | NTC | FAM | 0 | #NUM! | 18.37 | 0 | 37.96
miniS | A08 | CY5 | 34.97 | 34.8 |  |  | miniS | NTC | CY5 | 34.8
miniS | B01 | CY5 | 38.93 |  |  |  | miniS | 1 | FAM | 11.28
miniS | B02 | CY5 | 38.75 |  |  |  | miniS | 2 | FAM | 22.935 |  |  | TubeName | Conc (uM) | ng | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul)
miniS | B03 | CY5 | 39.96 |  |  |  | miniS | 3 | FAM | 22.075 |  |  | IDT | 100 | 122000 |  | 2.4088e+15
miniS | B04 | CY5 | 36.88 |  |  |  | miniS | 4 | FAM | 29.175
miniS | B05 | CY5 | 37.39 |  |  |  | miniS | 5 | FAM | 28.685 |  |  | Tube_0 | 10 | 30500 |  | 6.022e+14
miniS | B06 | CY5 | 38.6 |  |  |  | miniS | 6 | FAM | 29.565
miniS | B07 | CY5 | 34.31 |  |  |  | miniS | 7 | FAM | 30.5 |  |  | Tube_1 | 0.1 | 305 | 6.2742e+12 | 6.022e+12 | 12044000000
miniS | B08 | CY5 | 34.63 |  |  |  | miniS | NTC | FAM | 18.37 |  |  | Tube_2 | 0.01 | 30.5 | 627418093126.6887 | 602200000000.0001 | 1204400000
miniS | A01 | FAM | 11.21 | 11.28 |  |  | Q16 | 1 | CY5 | 35.49 |  |  | Tube_3 | 0.001 | 3.05 | 62741809312.6689 | 60220000000 | 120440000
miniS | A02 | FAM | 22.99 | 22.935 |  |  | Q16 | 2 | CY5 | 36.19 |  |  | Tube_4 | 0.0001 | 0.305 | 6274180931.2669 | 6022000000 | 12044000
miniS | A03 | FAM | 21.59 | 22.075 |  |  | Q16 | 3 | CY5 | 36 |  |  | Tube_5 | 1e-05 | 0.0305 | 627418093.1267 | 602200000 | 1204400
miniS | A04 | FAM | 28.9 | 29.175 |  |  | Q16 | 4 | CY5 | 33.295 |  |  | Tube_6 | 1e-06 | 0.00305 | 62741809.3127 | 60220000 | 120440
miniS | A05 | FAM | 28.62 | 28.685 |  |  | Q16 | 5 | CY5 | 33.65 |  |  | Tube_7 | 1e-07 | 0.000305 | 6274180.9313 | 6022000 | 12044
miniS | A06 | FAM | 29.74 | 29.565 |  |  | Q16 | 6 | CY5 | 33.51 |  |  | Tube_8 | 1e-08 | 3.05e-05 | 627418.0931 | 602200 | 1204.4
miniS | A07 | FAM | 29.49 | 30.5 |  |  | Q16 | 7 | CY5 | 34.57 |  |  | Tube_9 | 1e-09 | 3.05e-06 | 62741.8093 | 60220 | 120.44
miniS | A08 | FAM | 18.37 | 18.37 |  |  | Q16 | NTC | CY5 | 33.985 |  |  | Tube_10 | 1e-10 | 3.05e-07 | 6274.1809 | 6022 | 12.044
miniS | B01 | FAM | 11.35 |  |  |  | Q16 | 1 | FAM | 11.815 |  |  | Tube_11 | 1e-11 | 3.05e-08 | 627.4181 | 602.2 | 1.2044
miniS | B02 | FAM | 22.88 |  |  |  | Q16 | 2 | FAM | 23.88 |  |  | Tube_12 | 1e-12 | 3.05e-09 | 62.7418 | 60.22 | 0.12044
miniS | B03 | FAM | 22.56 |  |  |  | Q16 | 3 | FAM | 22.45 |  |  | Tube_13 | 1e-13 | 3.05e-10 | 6.2742 | 6.022 | 0.012044
miniS | B04 | FAM | 29.45 |  |  |  | Q16 | 4 | FAM | 28.515 |  |  | Tube_14 | 1e-14 | 3.05e-11 | 0.62742 | 0.6022 | 0.0012044
miniS | B05 | FAM | 28.75 |  |  |  | Q16 | 5 | FAM | 29.06 |  |  | Tube_15 | 1e-15 | 3.05e-12 | 0.062742 | 0.06022 | 0.00012044
miniS | B06 | FAM | 29.39 |  |  |  | Q16 | 6 | FAM | 30.375
miniS | B07 | FAM | 31.51 |  |  |  | Q16 | 7 | FAM | 31.97
miniS | B08 | FAM |  |  |  |  | Q16 | NTC | FAM | 0
Q16 | A01 | CY5 | 36.67 | 35.49 |  |  | Q96 | 1 | CY5 | 40.81
Q16 | A02 | CY5 | 36.84 | 36.19 |  |  | Q96 | 2 | CY5 | 38.095
Q16 | A03 | CY5 |  | 36 |  |  | Q96 | 3 | CY5 | 41.46
Q16 | A04 | CY5 | 32.64 | 33.295 |  |  | Q96 | 4 | CY5 | 38.74
Q16 | A05 | CY5 | 34.19 | 33.65 |  |  | Q96 | 5 | CY5 | 35.92
Q16 | A06 | CY5 | 33.51 | 33.51 |  |  | Q96 | 6 | CY5 | 36.685
Q16 | A07 | CY5 | 34.57 | 34.57 |  |  | Q96 | 7 | CY5 | 36.88
Q16 | A08 | CY5 | 35.52 | 33.985 |  |  | Q96 | NTC | CY5 | 0
Q16 | B01 | CY5 | 34.31 |  |  |  | Q96 | 1 | FAM | 12.485
Q16 | B02 | CY5 | 35.54 |  |  |  | Q96 | 2 | FAM | 22.575
Q16 | B03 | CY5 | 36 |  |  |  | Q96 | 3 | FAM | 21.54
Q16 | B04 | CY5 | 33.95 |  |  |  | Q96 | 4 | FAM | 28.82
Q16 | B05 | CY5 | 33.11 |  |  |  | Q96 | 5 | FAM | 28.225
Q16 | B06 | CY5 |  |  |  |  | Q96 | 6 | FAM | 29.17
Q16 | B07 | CY5 |  |  |  |  | Q96 | 7 | FAM | 32.34
Q16 | B08 | CY5 | 32.45 |  |  |  | Q96 | NTC | FAM | 37.96
Q16 | A01 | FAM | 11.87 | 11.815
Q16 | A02 | FAM | 23.84 | 23.88
Q16 | A03 | FAM | 22.41 | 22.45
Q16 | A04 | FAM | 27.94 | 28.515
Q16 | A05 | FAM | 28.91 | 29.06
Q16 | A06 | FAM | 30.49 | 30.375
Q16 | A07 | FAM | 31.67 | 31.97
Q16 | A08 | FAM | 0 | 0
Q16 | B01 | FAM | 11.76
Q16 | B02 | FAM | 23.92
Q16 | B03 | FAM | 22.49
Q16 | B04 | FAM | 29.09
Q16 | B05 | FAM | 29.21
Q16 | B06 | FAM | 30.26
Q16 | B07 | FAM | 32.27
Q16 | B08 | FAM | 0
Q96 | A01 | CY5 | 41.2 | 40.81
Q96 | A02 | CY5 | 39.5 | 38.095
Q96 | A03 | CY5 | 44.12 | 41.46
Q96 | A04 | CY5 | 38.63 | 38.74
Q96 | A05 | CY5 | 35.92 | 35.92
Q96 | A06 | CY5 | 36.95 | 36.685
Q96 | A07 | CY5 | 34.86 | 36.88
Q96 | A08 | CY5 |  | 0
Q96 | B01 | CY5 | 40.42
Q96 | B02 | CY5 | 36.69
Q96 | B03 | CY5 | 38.8
Q96 | B04 | CY5 | 38.85
Q96 | B05 | CY5
Q96 | B06 | CY5 | 36.42
Q96 | B07 | CY5 | 38.9
Q96 | B08 | CY5
Q96 | A01 | FAM | 11.82 | 12.485
Q96 | A02 | FAM | 22.77 | 22.575
Q96 | A03 | FAM | 21.77 | 21.54
Q96 | A04 | FAM | 29.1 | 28.82
Q96 | A05 | FAM | 28.37 | 28.225
Q96 | A06 | FAM | 29.07 | 29.17
Q96 | A07 | FAM | 32.15 | 32.34
Q96 | A08 | FAM |  | 37.96
Q96 | B01 | FAM | 13.15
Q96 | B02 | FAM | 22.38
Q96 | B03 | FAM | 21.31
Q96 | B04 | FAM | 28.54
Q96 | B05 | FAM | 28.08
Q96 | B06 | FAM | 29.27
Q96 | B07 | FAM | 32.53
Q96 | B08 | FAM | 37.96
creating liquid master mix to test liquid vs lyophilized.
lyophilized mastermix tube containing everything but DNA.
Component | Desired | Units | PROBE RXN | 19.2
mmix |  |  | 12.1122 | 232.5547
sample |  |  | 0 | 0
IAC_F2 | 0.3uM |  | 0.6 | 11.52
IAC ultramer "std_7" |  |  | 0.2 | 3.84
water |  |  | 5.0878 | 97.6853
TOTAL |  |  | 18 | 345.6
make mastermix
then aliquot to 8 tubes. Add dna. Mix. Aliquot to strip tubes.
Component | Desired | Units | PROBE RXN | 2.3
mmix |  |  | 18 | 41.4
sample |  |  | 2 | 4.6
IAC_F2 | 0.3uM |  | 0 | 0
IAC ultramer "std_7" |  |  | 0 | 0
water |  |  | 0 | 0
TOTAL |  |  | 20 | 46
see log for experimental details
Used OT-2 program to run 6 strip tubes with stds2..7.
20210526
Run 1
Two were each run on Q16 (annitoa), Q96 and miniS (16-well format from BioER).
Q16 failed with fatal error. Greg reported. Didn't pull data from miniS or from Q96
Run 2
late that day, 20210526
Prepared 6 strip tubes.
Ran on Q16, Q96 and miniS.
Analyzed data and copied above.
Run 3
Prepared 6 strip tubes. Used stds 8, 9, 10, 11, 12, 13, 14
Also prepared 2 strip tubes with liquid formulations according to above schedule.
Ran the liquid strip tube on both Q96 and Q16. In each case it's the top tube (Row A, Row C) High to low (left to right)
4 tubes (1 liquid, 3 lyhophilized) on Q96.
keeping in this folder: "C:\Users\HarleyKing\OneDrive - LuminUltra Technologies Ltd\Documents\Exp800-automation\Exp800.10 Optimal Concentration of IAC in Leg pneumo Assay\qPCR\Lyophilization_Q96_vs_Q16_vs_miniS"
miniS time is incorrect. Will update.
Run 4
these runs conducted 5/29/21, Saturday
20210529
Need LOD. Made new Leg pneumo std tubes with new program that uses new tips for dilutions
Have 6, 8-strip tubes remaining. Using 4 for testing stds7..std13 with NTC.
Remaining 2 strip tubes will be used for sample testing.
20210604
Run 5
Ran std7..std14 on Leg spp batch 004 and Leg pneumo. Observe similar LODs
Component | Desired | Units | PROBE RXN | 48
mmix |  |  | 0 | 0
sample |  |  | 0 | 0
IAC_F2 | 0.3uM |  | 0.6 | 28.8
IAC ultramer "std_7" |  |  | 0.2 | 9.6
water |  |  | 17.2 | 825.6
TOTAL |  |  | 18 | 864
Component | Desired | Units | PROBE RXN | 10
mmix |  |  | 0 | 0
sample |  |  | 0 | 0
IAC_F2 | 0.3uM |  | 0.6 | 6
IAC ultramer "std_7" |  |  | 0.2 | 2
water |  |  | 17.2 | 172
TOTAL |  |  | 18 | 180
Sample | H2O | BAC tablet | Wei (Leg) | gDNA | std_5 | IAC_F2 | IAC_std_8 | TOTAL |  | Sample | H2O | BAC tablet | Wei (Leg) | gDNA | std_5 | IAC_F2 | IAC_std_8 | TOTAL
| 0 | 19.2 |  |  |  | 0.6 | 0.2 | 20 |  | 1 | 0 | 192 | 0 | 0 | 0 | 6 | 2 | 200
| 9.2 | 10 |  |  |  | 0.6 | 0.2 | 20 |  | 2 | 92 | 100 | 0 | 0 | 0 | 6 | 2 | 200
| 14.2 | 5 |  |  |  | 0.6 | 0.2 | 20 |  | 3 | 142 | 50 | 0 | 0 | 0 | 6 | 2 | 200
| 18.2 | 1 |  |  |  | 0.6 | 0.2 | 20 |  | 4 | 182 | 10 | 0 | 0 | 0 | 6 | 2 | 200
| 14.2 |  | 5 |  |  | 0.6 | 0.2 | 20 |  | 5 | 142 | 0 | 50 | 0 | 0 | 6 | 2 | 200
| 14.2 |  |  | 5 |  | 0.6 | 0.2 | 20 |  | 6 | 142 | 0 | 0 | 50 | 0 | 6 | 2 | 200
| 17.2 |  |  |  | 2 | 0.6 | 0.2 | 20 |  | 7 | 172 | 0 | 0 | 0 | 20 | 6 | 2 | 200
| 19.2 |  |  |  |  | 0.6 | 0.2 | 20 |  | 8 | 192 | 0 | 0 | 0 | 0 | 6 | 2 | 200
20210601
presented Lyophilzed Leg pneumo results in lab meeting. In short, able to calc LOD and do sample detection.
Greg and Wei both validate that difficult to detect Leg pneumo in tablet produced by ieLABs.
Wei recommended I ask Neil to order more
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 500
Stock Concentrations |  |  | Overage = |  |  | 5%
dNTP Mix | 25 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 20 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Dextran | 25 | Percent
Leg pneumo F | 100 | uM
Leg pneumo R | 100 | uM
Leg pneumo P | 100 | uM
IAC_F2 | 10 | uM
IAC_R | 100 | uM
IAC_P | 100 | uM
IAC_Ultramer (std_7) | 10000 | copies/ul
Component | Desired |  | Per 20 uL |  | Multiple Reactions
dNTP Mix | 0.25 | mM | 0.32 | uL | 168 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 0 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 0 | uL
Trehalose | 10 | Percent | 4 | uL | 0 | uL
BSA | 0.5 | mg/mL | 0.5 | uL | 0 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 0 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 0 | uL
dextran | 3 | Percent | 2.4 | uL | 0 | uL
Leg pneumo F | 0.4 | uM | 0.08 | uL | 0 | uL
Leg pneumo R | 0.4 | uM | 0.08 | uL | 0 | uL
Leg pneumo P | 0.5 | uM | 0.1 | uL | 0 | uL
IAC_F2 | 0.3 | uM | 0.6 | uL | 0 | uL
IAC_R | 0.3 | uM | 0.06 | uL | 0 | uL
IAC_P | 0.25 | uM | 0.05 | uL | 0 | uL
IAC Ultramer | 10000 | copies/ul | 0.2 | uL | 0 | uL
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 12.7322 | uL | 0 | uL
20210610
LOQ doesn't look nice in tab "New.Leg.pne_std_curve"
Think it was due to Raquel not spinnign down reaction mix in sample wells.
Repeating.
﻿Perform on Q96. Think I ﻿use 4 tubes with stds 1.7.
﻿Then 4 tubes with stds 2..8
﻿Then 4 tubes with stds 3..9
﻿Hope that gets something in the linear range.
will use the OT-2 to do in batches of 4, but I will make the samples
Component | Desired | Units | PROBE RXN | 24
mmix |  |  | 0 | 0
DNA std |  |  | 2 | 48
water |  |  | 18 | 432
TOTAL |  |  | 20 | 480
Conclusions
LOQ doesn't look so straight. Want to repeat with new set of Leg pneumo stds. Also, having Raquel spin down samples before lyophilization may have made a difference. Used stds2-8 to generate linear curve and determine LOQ. Will repeat.

## pivot
Sum of Ct_AVG | Column Labels
Row Labels | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Grand Total
FAM | 35.58 | 69.39 | 66.065 | 86.51 | 85.97 | 89.11 | 94.81 | 56.33 | 583.765
miniS | 11.28 | 22.935 | 22.075 | 29.175 | 28.685 | 29.565 | 30.5 | 18.37 | 192.585
Q16 | 11.815 | 23.88 | 22.45 | 28.515 | 29.06 | 30.375 | 31.97 | 0 | 178.065
Q96 | 12.485 | 22.575 | 21.54 | 28.82 | 28.225 | 29.17 | 32.34 | 37.96 | 213.115
CY5 | 115.23 | 113.44 | 116.58 | 108.78 | 107.465 | 108.315 | 106.765 | 68.785 | 845.36
miniS | 38.93 | 39.155 | 39.12 | 36.745 | 37.895 | 38.12 | 35.315 | 34.8 | 300.08
Q16 | 35.49 | 36.19 | 36 | 33.295 | 33.65 | 33.51 | 34.57 | 33.985 | 276.69
Q96 | 40.81 | 38.095 | 41.46 | 38.74 | 35.92 | 36.685 | 36.88 | 0 | 268.59
Grand Total | 150.81 | 182.83 | 182.645 | 195.29 | 193.435 | 197.425 | 201.575 | 125.115 | 1429.125
Better plot a few tabs down
uses #8-->NTC
and shows conc in legend.

## New.Leg.pne_std_curve
Well | Target | Property | Dye | Ct | CT_MEAN | CT_STDEV |  | IDT DNA Data
A01 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0! | TubeName | Conc (uM) | pM | Copy Number | Copy Number/2ul | Log(CopyNumber) | CT_MEAN | CT_STDEV | 2*stdev
A02 | Legionella sp. | Unknown | FAM |  |  |  | IDT Source | 100 | 100000000 | 2.4088e+15
A03 | Legionella sp. | Unknown | FAM |  |  |  | std_0 | 1 | 1000000 | 6.022e+14
A04 | Legionella sp. | Unknown | FAM |  |  |  | std_1 | 0.01 | 10000 | 6.022e+12 | 12044000000 | 10.0808 | #DIV/0! | #DIV/0! | #DIV/0!
A05 | Legionella sp. | Unknown | FAM |  |  |  | std_2 | 0.001 | 1000 | 602200000000.0001 | 1204400000 | 9.0808 | #DIV/0! | #DIV/0! | #DIV/0!
A06 | Legionella sp. | Unknown | FAM | 29.53 | 30.01 | 0.9113 | std_3 | 0.0001 | 100 | 60220000000 | 120440000 | 8.0808 | 7.7967 | 0.11719 | 0.23438
A07 | Legionella sp. | Unknown | FAM | 30.91 |  |  | std_4 | 1e-05 | 10 | 6022000000 | 12044000 | 7.0808 | 12.0375 | 0.93239 | 1.8648
A08 | Legionella sp. | Unknown | FAM | 30.2 |  |  | std_5 | 1e-06 | 1 | 602200000 | 1204400 | 6.0808 | 15.6625 | 0.65556 | 1.3111
A09 | Legionella sp. | Unknown | FAM | 30.19 |  |  | std_6 | 1e-07 | 0.1 | 60220000 | 120440 | 5.0808 | 19.1275 | 0.9211 | 1.8422
A10 | Legionella sp. | Unknown | FAM | 28.74 |  |  | std_7 | 1e-08 | 0.01 | 6022000 | 12044 | 4.0808 | 22.475 | 0.70543 | 1.4109
A11 | Legionella sp. | Unknown | FAM |  |  |  | std_8 | 1e-09 | 0.001 | 602200 | 1204.4 | 3.0808 | 26.495 | 1.0081 | 2.0162
A12 | Legionella sp. | Unknown | FAM |  |  |  | std_9 | 1e-10 | 0.0001 | 60220 | 120.44 | 2.0808 | 30.01 | 0.9113 | 1.8226
B01 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0! | std_10 | 1e-11 | 1e-05 | 6022 | 12.044 | 1.0808 | 32.63 | 1.17 | 2.34
B02 | Legionella sp. | Unknown | FAM |  |  |  | std_11 | 1e-12 | 1e-06 | 602.2 | 1.2044 | 0.080771 | 33.11 | #DIV/0! | #DIV/0!
B03 | Legionella sp. | Unknown | FAM |  |  |  | std_12 | 1e-13 | 1e-07 | 60.22 | 0.12044 | -0.91923 | #DIV/0! | #DIV/0! | #DIV/0!
B04 | Legionella sp. | Unknown | FAM |  |  |  | std_13 | 1e-14 | 1e-08 | 6.022 | 0.012044 | -1.9192 | #DIV/0! | #DIV/0! | #DIV/0!
B05 | Legionella sp. | Unknown | FAM |  |  |  | std_14 | 1e-15 | 1e-09 | 0.6022 | 0.0012044 | -2.9192 | #DIV/0! | #DIV/0! | #DIV/0!
B06 | Legionella sp. | Unknown | FAM | 32.71 | 32.63 | 1.17 | std_15 | 1e-16 | 1e-10 | 0.06022 | 0.00012044 | -3.9192 | #DIV/0! | #DIV/0! | #DIV/0!
B07 | Legionella sp. | Unknown | FAM | 33.26
B08 | Legionella sp. | Unknown | FAM | [8 empty] | PCR Efficiency | #DIV/0!
B09 | Legionella sp. | Unknown | FAM | 33.35
B10 | Legionella sp. | Unknown | FAM | 31.28
B11 | Legionella sp. | Unknown | FAM
B12 | Legionella sp. | Unknown | FAM
C01 | Legionella sp. | Unknown | FAM | 8.72 | 7.7967 | 0.11719
C02 | Legionella sp. | Unknown | FAM | 7.93
C03 | Legionella sp. | Unknown | FAM | 7.75
C04 | Legionella sp. | Unknown | FAM | 7.71
C05 | Legionella sp. | Unknown | FAM
C06 | Legionella sp. | Unknown | FAM |  | 33.11 | #DIV/0!
C07 | Legionella sp. | Unknown | FAM
C08 | Legionella sp. | Unknown | FAM | 33.11
C09 | Legionella sp. | Unknown | FAM
C10 | Legionella sp. | Unknown | FAM
C11 | Legionella sp. | Unknown | FAM
C12 | Legionella sp. | Unknown | FAM
D01 | Legionella sp. | Unknown | FAM | 12.92 | 12.0375 | 0.93239
D02 | Legionella sp. | Unknown | FAM | 12.55
D03 | Legionella sp. | Unknown | FAM | 12.46
D04 | Legionella sp. | Unknown | FAM | 12.5
D05 | Legionella sp. | Unknown | FAM | 10.64
D06 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0!
D07 | Legionella sp. | Unknown | FAM
D08 | Legionella sp. | Unknown | FAM
D09 | Legionella sp. | Unknown | FAM
D10 | Legionella sp. | Unknown | FAM
D11 | Legionella sp. | Unknown | FAM
D12 | Legionella sp. | Unknown | FAM
E01 | Legionella sp. | Unknown | FAM | 15.64 | 15.6625 | 0.65556
E02 | Legionella sp. | Unknown | FAM | 15.92
E03 | Legionella sp. | Unknown | FAM | 16.17
E04 | Legionella sp. | Unknown | FAM | 15.86
E05 | Legionella sp. | Unknown | FAM | 14.7
E06 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0!
E07 | Legionella sp. | Unknown | FAM
E08 | Legionella sp. | Unknown | FAM
E09 | Legionella sp. | Unknown | FAM
E10 | Legionella sp. | Unknown | FAM
E11 | Legionella sp. | Unknown | FAM
E12 | Legionella sp. | Unknown | FAM
F01 | Legionella sp. | Unknown | FAM | 19.8 | 19.1275 | 0.9211
F02 | Legionella sp. | Unknown | FAM | 19.82
F03 | Legionella sp. | Unknown | FAM | 19.49
F04 | Legionella sp. | Unknown | FAM | 19.43
F05 | Legionella sp. | Unknown | FAM | 17.77
F06 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0!
F07 | Legionella sp. | Unknown | FAM
F08 | Legionella sp. | Unknown | FAM
F09 | Legionella sp. | Unknown | FAM
F10 | Legionella sp. | Unknown | FAM
F11 | Legionella sp. | Unknown | FAM
F12 | Legionella sp. | Unknown | FAM
G01 | Legionella sp. | Unknown | FAM | 22.89 | 22.475 | 0.70543
G02 | Legionella sp. | Unknown | FAM | 22.96
G03 | Legionella sp. | Unknown | FAM | 22.92
G04 | Legionella sp. | Unknown | FAM | 22.57
G05 | Legionella sp. | Unknown | FAM | 21.45
G06 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0!
G07 | Legionella sp. | Unknown | FAM
G08 | Legionella sp. | Unknown | FAM
G09 | Legionella sp. | Unknown | FAM
G10 | Legionella sp. | Unknown | FAM
G11 | Legionella sp. | Unknown | FAM
G12 | Legionella sp. | Unknown | FAM
H01 | Legionella sp. | Unknown | FAM | 9.43 | 26.495 | 1.0081
H02 | Legionella sp. | Unknown | FAM | 26.99
H03 | Legionella sp. | Unknown | FAM | 27.39
H04 | Legionella sp. | Unknown | FAM | 26.52
H05 | Legionella sp. | Unknown | FAM | 25.08
H06 | Legionella sp. | Unknown | FAM |  | #DIV/0! | #DIV/0!
H07 | Legionella sp. | Unknown | FAM
H08 | Legionella sp. | Unknown | FAM
H09 | Legionella sp. | Unknown | FAM
H10 | Legionella sp. | Unknown | FAM
H11 | Legionella sp. | Unknown | FAM
H12 | Legionella sp. | Unknown | FAM
Conclusions:
This plate was started and then a power outage shut off the power about 10-15min into start. I began the program anew after deleting the 3min init denaturation. Turned cycles down from 45 to 40. Not being too critical of the data or stressing over the curves. In fact, not doing them.
Instead, this plate was run to validate the hypothesis that the reason I keep detecting DNA in tubes with fewer than 1 copy of DNA is because I"m using the same tip. Re-wrote the dilution protocol to shuck tips and use new ones to perform dilution. This plate and it's failures above std_11 suggest that the tip does matter and can contaminate.
Moving on to next test: sample with std curve using std9..std14 to calculate LOD.

## new_Legpne_std_liquid_LOD
20210529
Testing out Leg pneumo std curve with clean tips between mixes in preparation for new std curve
Does dna carry over because I'm using the same tip?
Blank - Master Mix for Lyophilization |  |  | Number of Reactions = |  |  | 500
Stock Concentrations |  |  | Overage = |  |  | 5%
dNTP Mix | 10 | mM
Mg2+ | 1000 | mM
AptaTaq | 50 | Units/uL
Trehalose | 50 | Percent
BSA | 50 | mg/mL
10X Taq Buffer (No Mg2+) | 10 | x
2-Hydroxypropyl-β-Cyclodextrin | 45 | Percent
Dextran | 25 | Percent
Leg pneumo F | 100 | uM
Leg pneumo R | 100 | uM
Leg pneumo P | 100 | uM
IAC_F2 | 100 | uM
IAC_R | 100 | uM
IAC_P | 100 | uM
Component | Desired |  | Per 20 uL |  | Multiple Reactions
dNTP Mix | 0.25 | mM | 0.8 | uL | 0 | uL
AptaTaq | 2.5 | Units | 0.05 | uL | 0 | uL
Mg2+ | 3.5 | mM | 0.07 | uL | 0 | uL
Trehalose | 10 | Percent | 4 | uL | 0 | uL
BSA | 0.5 | mg/mL | 0.2 | uL | 0 | uL
PCR Buffer 10X | 1 | x | 2 | uL | 0 | uL
2-Hydroxypropyl-β-Cyclodextrin | 5 | Percent | 2.2222 | uL | 0 | uL
dextran | 3 | Percent | 2.4 | uL | 0 | uL
Leg pneumo F | 0.4 | uM | 0.08 | uL | 0 | uL
Leg pneumo R | 0.4 | uM | 0.08 | uL | 0 | uL
Leg pneumo P | 0.5 | uM | 0.1 | uL | 0 | uL
IAC_F2 | 0.3 | uM | 0 | uL | 0 | uL
IAC_R | 0.3 | uM | 0.06 | uL | 0 | uL
IAC_P | 0.25 | uM | 0.05 | uL | 0 | uL
IAC Ultramer | 0 |  | 0 | uL | 0 | uL
Water |  |  | 0 | uL | 0 | uL
|  | Total: | 12.1122 | uL | 0 | uL
Component | Desired | Units | PROBE RXN | 74
mmix |  |  | 12.1122 | 896.3044
IAC_F2 |  |  | 0.6 | 44.4
IAC Ultramer ("std_8") |  |  | 0.2 | 14.8
water |  |  | 5.0878 | 376.4956
TOTAL |  |  | 18 | 1332
power went out with this plate after about 10-15 minutes.
reset with new denaturation and 40 cycles.
Component | Desired | Units | PROBE RXN | 115.2
mmix |  |  | 12.1122 | 1395.328
IAC_F2 |  |  | 0.6 | 69.12
IAC Ultramer ("std_8") |  |  | 0.2 | 23.04
water |  |  | 5.0878 | 586.112
TOTAL |  |  | 18 | 2073.6
Well | Target | Property | Dye | Ct | Ct_Mean | Ct_SD
A01 | Legionella sp. | Unknown | FAM | 34.46 | 35.5236 | 0.5221
A02 | Legionella sp. | Unknown | FAM | 35.54
A03 | Legionella sp. | Unknown | FAM | 36.41
A04 | Legionella sp. | Unknown | FAM | 35.94
A05 | Legionella sp. | Unknown | FAM | 36.27
A06 | Legionella sp. | Unknown | FAM | 34.84
A07 | Legionella sp. | Unknown | FAM | 35.26
A08 | Legionella sp. | Unknown | FAM | 35.84
A09 | Legionella sp. | Unknown | FAM | 35.04 |  |  | Row | TubeName | Conc (uM) | ng | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul) | log(copy#) | Failed | Fraction Positive
A10 | Legionella sp. | Unknown | FAM | 35.33 |  |  |  | Tube_7 | 1e-07 | 0.000305 | 6274180.9313 | 6022000 | 12044 | 4.0808 | 0% | 100%
A11 | Legionella sp. | Unknown | FAM | 35.06 |  |  |  | Tube_8 | 1e-08 | 3.05e-05 | 627418.0931 | 602200 | 1204.4 | 3.0808 | 0% | 100%
A12 | Legionella sp. | Unknown | FAM | 35.23 |  |  | A | Tube_9 | 1e-09 | 3.05e-06 | 62741.8093 | 60220 | 120.44 | 2.0808 | 0% | 100%
B01 | Legionella sp. | Unknown | FAM | 37.12 | 38.3929 | 0.74674 | B | Tube_10 | 1e-10 | 3.05e-07 | 6274.1809 | 6022 | 12.044 | 1.0808 | 45.45% | 54.55%
B02 | Legionella sp. | Unknown | FAM | 38.53 |  |  | C | Tube_11 | 1e-11 | 3.05e-08 | 627.4181 | 602.2 | 1.2044 | 0.080771 | 100% | 0%
B03 | Legionella sp. | Unknown | FAM |  |  |  | D | Tube_12 | 1e-12 | 3.05e-09 | 62.7418 | 60.22 | 0.12044 | -0.91923 | 100% | 0%
B04 | Legionella sp. | Unknown | FAM | 38.45 |  |  | E | Tube_13 | 1e-13 | 3.05e-10 | 6.2742 | 6.022 | 0.012044 | -1.9192 | 100% | 0% |  |  |  |  | y = 100 + (0 - 100)/(1 + (x/0.5807707)^3)^m
B05 | Legionella sp. | Unknown | FAM | 37.63 |  |  | F | Tube_14 | 1e-14 | 3.05e-11 | 0.62742 | 0.6022 | 0.0012044 | -2.9192 | 100% | 0%
B06 | Legionella sp. | Unknown | FAM | [17 empty] | y = 100 + (0 - 100)/(1 + (x/0.5807707)^3)^m
B07 | Legionella sp. | Unknown | FAM | 37.57 |  |  |  |  | Conc | log(conc) | Fraction Positive | Model |  | interpolating copy number/2ul
B08 | Legionella sp. | Unknown | FAM | [5 empty] | 1e-07 | -7 | 100 | #NUM! |  | 120440000000 | 4.5475e-13
B09 | Legionella sp. | Unknown | FAM | [5 empty] | 1e-08 | -8 | 100 | 100 | [11 empty] | y = 100 + (0 - 100)/(1 + (x/0.8)^3)^m
B10 | Legionella sp. | Unknown | FAM | 39.66 |  |  |  |  | 1e-09 | -9 | 100 | 100 |  | LOD Conc(uM) | percent
B11 | Legionella sp. | Unknown | FAM | 37.98 |  |  |  |  | 1e-10 | -10 | 55 | 54.9998 |  | -9.9274 | 95
B12 | Legionella sp. | Unknown | FAM | 38.93 |  |  |  |  | 1e-11 | -11 | 0 | 6.2219e-07 |  | 1.1821e-10 | uM | 142.3691
C01 | Legionella sp. | Unknown | FAM | [5 empty] | 1e-12 | -12 | 0 | -2.074e-07 |  | 14.2369 | DNA copies/ul
C02 | Legionella sp. | Unknown | FAM | [5 empty] | 1e-13 | -13 | 0 | -2.074e-07 |  | 71184.5525
C03 | Legionella sp. | Unknown | FAM | [5 empty] | 1e-14 | -14 | 0 | -2.074e-07 |  | 142.3691 | DNA copies/ul
C04 | Legionella sp. | Unknown | FAM | [8 empty] | y = 100 + (-2.073971e-7 - 100)/(1 + (x/-9.858086)^-193.1934)^13.02813 |  | Above is corrected version. 20210623 hjk
C05 | Legionella sp. | Unknown | FAM
C06 | Legionella sp. | Unknown | FAM | [5 empty] | model didn't do well with using log(copy#) prob because it was close to 0 and log(0) = undefined.
C07 | Legionella sp. | Unknown | FAM | [5 empty] | so I used the log (conc(um)), and that yielded a more accurate model.
C08 | Legionella sp. | Unknown | FAM
C09 | Legionella sp. | Unknown | FAM
C10 | Legionella sp. | Unknown | FAM
C11 | Legionella sp. | Unknown | FAM
C12 | Legionella sp. | Unknown | FAM
D01 | Legionella sp. | Unknown | FAM
D02 | Legionella sp. | Unknown | FAM
D03 | Legionella sp. | Unknown | FAM
D04 | Legionella sp. | Unknown | FAM
D05 | Legionella sp. | Unknown | FAM
D06 | Legionella sp. | Unknown | FAM
D07 | Legionella sp. | Unknown | FAM
D08 | Legionella sp. | Unknown | FAM
D09 | Legionella sp. | Unknown | FAM
D10 | Legionella sp. | Unknown | FAM
D11 | Legionella sp. | Unknown | FAM
D12 | Legionella sp. | Unknown | FAM | [10 empty] | 0.8 | -0.09691
E01 | Legionella sp. | Unknown | FAM
E02 | Legionella sp. | Unknown | FAM
E03 | Legionella sp. | Unknown | FAM
E04 | Legionella sp. | Unknown | FAM
E05 | Legionella sp. | Unknown | FAM | [17 empty] | https://bmcbioinformatics.biomedcentral.com/articles/10.1186/1471-2105-9-221
E06 | Legionella sp. | Unknown | FAM
E07 | Legionella sp. | Unknown | FAM
E08 | Legionella sp. | Unknown | FAM
E09 | Legionella sp. | Unknown | FAM
E10 | Legionella sp. | Unknown | FAM
E11 | Legionella sp. | Unknown | FAM
E12 | Legionella sp. | Unknown | FAM
F01 | Legionella sp. | Unknown | FAM
F02 | Legionella sp. | Unknown | FAM
F03 | Legionella sp. | Unknown | FAM
F04 | Legionella sp. | Unknown | FAM
F05 | Legionella sp. | Unknown | FAM
F06 | Legionella sp. | Unknown | FAM
F07 | Legionella sp. | Unknown | FAM | [5 empty] | TubeName | Conc (uM) | ng | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul)
F08 | Legionella sp. | Unknown | FAM | [5 empty] | IDT | 100 | 122000 |  | 2.4088e+15
F09 | Legionella sp. | Unknown | FAM
F10 | Legionella sp. | Unknown | FAM | [5 empty] | Tube_0 | 10 | 30500 |  | 6.022e+14
F11 | Legionella sp. | Unknown | FAM
F12 | Legionella sp. | Unknown | FAM | [5 empty] | Tube_1 | 0.1 | 305 | 6.2742e+12 | 6.022e+12 | 12044000000
G01 | Legionella sp. | Unknown | FAM | 9.21 |  |  |  |  | Tube_2 | 0.01 | 30.5 | 627418093126.6887 | 602200000000.0001 | 1204400000
G02 | Legionella sp. | Unknown | FAM | 9.9 |  |  |  |  | Tube_3 | 0.001 | 3.05 | 62741809312.6689 | 60220000000 | 120440000
G03 | Legionella sp. | Unknown | FAM | 10.1 |  |  |  |  | Tube_4 | 0.0001 | 0.305 | 6274180931.2669 | 6022000000 | 12044000
G04 | Legionella sp. | Unknown | FAM | 13.55 |  |  |  |  | Tube_5 | 1e-05 | 0.0305 | 627418093.1267 | 602200000 | 1204400
G05 | Legionella sp. | Unknown | FAM | 13.74 |  |  |  |  | Tube_6 | 1e-06 | 0.00305 | 62741809.3127 | 60220000 | 120440
G06 | Legionella sp. | Unknown | FAM | 13.81 |  |  |  |  | Tube_7 | 1e-07 | 0.000305 | 6274180.9313 | 6022000 | 12044
G07 | Legionella sp. | Unknown | FAM | 17.88 |  |  |  |  | Tube_8 | 1e-08 | 3.05e-05 | 627418.0931 | 602200 | 1204.4
G08 | Legionella sp. | Unknown | FAM | 17.8 |  |  |  |  | Tube_9 | 1e-09 | 3.05e-06 | 62741.8093 | 60220 | 120.44
G09 | Legionella sp. | Unknown | FAM | 17.74 |  |  |  |  | Tube_10 | 1e-10 | 3.05e-07 | 6274.1809 | 6022 | 12.044
G10 | Legionella sp. | Unknown | FAM | 20.94 |  |  |  |  | Tube_11 | 1e-11 | 3.05e-08 | 627.4181 | 602.2 | 1.2044
G11 | Legionella sp. | Unknown | FAM | 21.7 |  |  |  |  | Tube_12 | 1e-12 | 3.05e-09 | 62.7418 | 60.22 | 0.12044
G12 | Legionella sp. | Unknown | FAM | 22.12 |  |  |  |  | Tube_13 | 1e-13 | 3.05e-10 | 6.2742 | 6.022 | 0.012044
H01 | Legionella sp. | Unknown | FAM | 25.19 |  |  |  |  | Tube_14 | 1e-14 | 3.05e-11 | 0.62742 | 0.6022 | 0.0012044
H02 | Legionella sp. | Unknown | FAM | 25.27 |  |  |  |  | Tube_15 | 1e-15 | 3.05e-12 | 0.062742 | 0.06022 | 0.00012044
H03 | Legionella sp. | Unknown | FAM | 25.16
H04 | Legionella sp. | Unknown | FAM | 28.46
H05 | Legionella sp. | Unknown | FAM | 28.87
H06 | Legionella sp. | Unknown | FAM | 28.68
H07 | Legionella sp. | Unknown | FAM | 32.26
H08 | Legionella sp. | Unknown | FAM | 32.39
H09 | Legionella sp. | Unknown | FAM | 32.28
H10 | Legionella sp. | Unknown | FAM | [6 empty] | 20210623 HK re-evaluating
H11 | Legionella sp. | Unknown | FAM
H12 | Legionella sp. | Unknown | FAM | 38.51 | [5 empty] | TubeName | Conc (uM) | ng | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul) |  |  | copy number (tube) | Conc | log(conc) | Fraction Positive | Model
[10 empty] | IDT | 92.35 | 122000 |  | 2.2245e+15 |  |  |  |  | 1e-09 | -9 | 100 | 100
[19 empty] | 1e-10 | -10 | 55 | 55.0002
[10 empty] | Tube_0 | 0.9235 | 30500 |  | 5.5613e+14 |  |  |  |  | 1e-11 | -11 | 0 | 6.1293e-07
[19 empty] | 1e-12 | -12 | 0 | -2.0431e-07
[10 empty] | Tube_1 | 0.009235 | 305 | 6.2742e+12 | 5.5613e+12 | 11122634000 |  |  |  | 1e-13 | -13 | 0 | -2.0431e-07
[10 empty] | Tube_2 | 0.0009235 | 30.5 | 627418093126.6887 | 556131700000.0001 | 1112263400 |  |  |  | 1e-14 | -14 | 0 | -2.0431e-07
[10 empty] | Tube_3 | 9.235e-05 | 3.05 | 62741809312.6689 | 55613170000 | 111226340 |  | copies/rxn |  |  |  | interpolating
[10 empty] | Tube_4 | 9.235e-06 | 0.305 | 6274180931.2669 | 5561317000 | 11122634 |  | 142.3177 | 71158.832 | 1.1816e-10 | -9.9275 |  | 95
[10 empty] | Tube_5 | 9.235e-07 | 0.0305 | 627418093.1267 | 556131700 | 1112263.4
[10 empty] | Tube_6 | 9.235e-08 | 0.00305 | 62741809.3127 | 55613170 | 111226.34
[10 empty] | Tube_7 | 9.235e-09 | 0.000305 | 6274180.9313 | 5561317 | 11122.634
[10 empty] | Tube_8 | 9.235e-10 | 3.05e-05 | 627418.0931 | 556131.7 | 1112.2634
[10 empty] | Tube_9 | 9.235e-11 | 3.05e-06 | 62741.8093 | 55613.17 | 111.2263 | [7 empty] | y = 100 + (-2.043069e-7 - 100)/(1 + (x/-9.857024)^-193.3437)^13.32201
[10 empty] | Tube_10 | 9.235e-12 | 3.05e-07 | 6274.1809 | 5561.317 | 11.1226 |  | slope | y-int
[10 empty] | Tube_11 | 9.235e-13 | 3.05e-08 | 627.4181 | 556.1317 | 1.1123 |  | 1.2044e+12 | 0
[10 empty] | Tube_12 | 9.235e-14 | 3.05e-09 | 62.7418 | 55.6132 | 0.11123 |  | interpolation
[10 empty] | Tube_13 | 9.235e-15 | 3.05e-10 | 6.2742 | 5.5613 | 0.011123 |  | 142.3177
[10 empty] | Tube_14 | 9.235e-16 | 3.05e-11 | 0.62742 | 0.55613 | 0.0011123
[10 empty] | Tube_15 | 9.235e-17 | 3.05e-12 | 0.062742 | 0.055613 | 0.00011123

## lyoph_LOD
Well | Target | Property | Dye | Ct | Ct Mean | Ct SD | Concentration | Aver. Con. | Con. SD | Sample ID | Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct
C03 | Internal Control | Unknown | CY5 | [7 empty] | C03 | Internal Control | Unknown | CY5 |  | C03 | Legionella sp. | Unknown | FAM |  | BAC_20ul
C03 | Legionella sp. | Unknown | FAM | [7 empty] | C04 | Internal Control | Unknown | CY5 | 35.66 | C04 | Legionella sp. | Unknown | FAM |  | BAC_10ul
C04 | Internal Control | Unknown | CY5 | 35.66 | [6 empty] | C05 | Internal Control | Unknown | CY5 | 36.45 | C05 | Legionella sp. | Unknown | FAM |  | BAC_5ul |  | TubeName | Conc (uM) | ng | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul) | log(copy#) | from this assay | from this assay | from this assay | samples from Q16 (in this folder copy pasted) | Failed | Fraction Positive
C04 | Legionella sp. | Unknown | FAM | [7 empty] | C06 | Internal Control | Unknown | CY5 | 35.97 | C06 | Legionella sp. | Unknown | FAM |  | BAC_1ul |  | Tube_7 | 1e-07 | 0.000305 | 6274180.9313 | 6022000 | 12044 | 4.0808 |  | 28.95 | 29.92 | 28.04 | 0% | 100%
C05 | Internal Control | Unknown | CY5 | 36.45 | [6 empty] | C07 | Internal Control | Unknown | CY5 | 35.14 | C07 | Legionella sp. | Unknown | FAM | 38.63 | Wei_extracted sample |  | Tube_8 | 1e-08 | 3.05e-05 | 627418.0931 | 602200 | 1204.4 | 3.0808 | 31.19 | 31.22 | 30.78 | 32.09 | 0% | 100%
C05 | Legionella sp. | Unknown | FAM | [7 empty] | C08 | Internal Control | Unknown | CY5 |  | C08 | Legionella sp. | Unknown | FAM | 19.56 | gDNA |  | Tube_9 | 1e-09 | 3.05e-06 | 62741.8093 | 60220 | 120.44 | 2.0808 | 34.37 | 34.48 | 34.8 | 35.4 | 0% | 100%
C06 | Internal Control | Unknown | CY5 | 35.97 | [6 empty] | C09 | Internal Control | Unknown | CY5 | 41.95 | C09 | Legionella sp. | Unknown | FAM | 21.58 | std_5 |  | Tube_10 | 1e-10 | 3.05e-07 | 6274.1809 | 6022 | 12.044 | 1.0808 | 37.72 | 35.76 |  | 38.14 | 25% | 75%
C06 | Legionella sp. | Unknown | FAM | [7 empty] | C10 | Internal Control | Unknown | CY5 | 37.36 | C10 | Legionella sp. | Unknown | FAM |  | NTC |  | Tube_11 | 1e-11 | 3.05e-08 | 627.4181 | 602.2 | 1.2044 | 0.080771 |  |  |  |  | 100% | 0%
C07 | Internal Control | Unknown | CY5 | 35.14 | [6 empty] | D03 | Internal Control | Unknown | CY5 |  | D03 | Legionella sp. | Unknown | FAM |  | 28.95 | 29.92 | Tube_12 | 1e-12 | 3.05e-09 | 62.7418 | 60.22 | 0.12044 | -0.91923 |  |  |  |  | 100% | 0%
C07 | Legionella sp. | Unknown | FAM | 38.63 | [6 empty] | D04 | Internal Control | Unknown | CY5 |  | D04 | Legionella sp. | Unknown | FAM | 31.19 | 31.22 | 30.78 | Tube_13 | 1e-13 | 3.05e-10 | 6.2742 | 6.022 | 0.012044 | -1.9192 |  |  |  |  | 100% | 0%
C08 | Internal Control | Unknown | CY5 | [7 empty] | D05 | Internal Control | Unknown | CY5 | 37.27 | D05 | Legionella sp. | Unknown | FAM | 34.37 | 34.48 | 34.8 | Tube_14 | 1e-14 | 3.05e-11 | 0.62742 | 0.6022 | 0.0012044 | -2.9192 |  |  |  |  | 100% | 0%
C08 | Legionella sp. | Unknown | FAM | 19.56 | [6 empty] | D06 | Internal Control | Unknown | CY5 | 37.19 | D06 | Legionella sp. | Unknown | FAM | 37.72 | 35.76 | [6 empty] | slope | y-int
C09 | Internal Control | Unknown | CY5 | 41.95 | [6 empty] | D07 | Internal Control | Unknown | CY5 | 36.82 | D07 | Legionella sp. | Unknown | FAM | [8 empty] | 120440000000 | 4.5475e-13
C09 | Legionella sp. | Unknown | FAM | 21.58 | [6 empty] | D08 | Internal Control | Unknown | CY5 | 36.48 | D08 | Legionella sp. | Unknown | FAM |  |  |  |  | Conc (uM) | Log(Conc) | Fraction Positive
C10 | Internal Control | Unknown | CY5 | 37.36 | [6 empty] | D09 | Internal Control | Unknown | CY5 | 36.54 | D09 | Legionella sp. | Unknown | FAM |  |  |  |  | 1e-07 | -7 | 100 | #NUM!
C10 | Legionella sp. | Unknown | FAM | [7 empty] | D10 | Internal Control | Unknown | CY5 | 36.54 | D10 | Legionella sp. | Unknown | FAM |  |  |  |  | 1e-08 | -8 | 100 | 100
D03 | Internal Control | Unknown | CY5 | [7 empty] | E03 | Internal Control | Unknown | CY5 |  | E03 | Legionella sp. | Unknown | FAM | 28.95 |  |  |  | 1e-09 | -9 | 100 | 100 |  | LOD Conc(uM) | percent
D03 | Legionella sp. | Unknown | FAM | [7 empty] | E04 | Internal Control | Unknown | CY5 | 42.62 | E04 | Legionella sp. | Unknown | FAM | 31.22 |  |  |  | 1e-10 | -10 | 75 | 74.9998 |  | -9.9598 | 95
D04 | Internal Control | Unknown | CY5 | [7 empty] | E05 | Internal Control | Unknown | CY5 |  | E05 | Legionella sp. | Unknown | FAM | 34.48 |  |  |  | 1e-11 | -11 | 0 | 3.9959e-07 |  | 1.0971e-10
D04 | Legionella sp. | Unknown | FAM | 31.19 | [6 empty] | E06 | Internal Control | Unknown | CY5 | 38.19 | E06 | Legionella sp. | Unknown | FAM | 35.76 |  |  |  | 1e-12 | -12 | 0 | -1.332e-07 |  | 13.2136 | copies/2ul
D05 | Internal Control | Unknown | CY5 | 37.27 | [6 empty] | E07 | Internal Control | Unknown | CY5 | 36.66 | E07 | Legionella sp. | Unknown | FAM |  |  |  |  | 1e-13 | -13 | 0 | -1.332e-07
D05 | Legionella sp. | Unknown | FAM | 34.37 | [6 empty] | E08 | Internal Control | Unknown | CY5 |  | E08 | Legionella sp. | Unknown | FAM |  |  |  |  | 1e-14 | -14 | 0 | -1.332e-07
D06 | Internal Control | Unknown | CY5 | 37.19 | [6 empty] | E09 | Internal Control | Unknown | CY5 | 38.67 | E09 | Legionella sp. | Unknown | FAM
D06 | Legionella sp. | Unknown | FAM | 37.72 | [6 empty] | E10 | Internal Control | Unknown | CY5 | 36.68 | E10 | Legionella sp. | Unknown | FAM
D07 | Internal Control | Unknown | CY5 | 36.82 | [6 empty] | F03 | Internal Control | Unknown | CY5 |  | F03 | Legionella sp. | Unknown | FAM | 29.92 |  |  |  | y = 100 + (-1.331975e-7 - 100)/(1 + (x/-9.88257)^-203.7598)^16.07043
D07 | Legionella sp. | Unknown | FAM | [7 empty] | F04 | Internal Control | Unknown | CY5 | 42.65 | F04 | Legionella sp. | Unknown | FAM | 30.78
D08 | Internal Control | Unknown | CY5 | 36.48 | [6 empty] | F05 | Internal Control | Unknown | CY5 | 36.87 | F05 | Legionella sp. | Unknown | FAM | 34.8
D08 | Legionella sp. | Unknown | FAM | [7 empty] | F06 | Internal Control | Unknown | CY5 | 37.06 | F06 | Legionella sp. | Unknown | FAM
D09 | Internal Control | Unknown | CY5 | 36.54 | [6 empty] | F07 | Internal Control | Unknown | CY5 |  | F07 | Legionella sp. | Unknown | FAM
D09 | Legionella sp. | Unknown | FAM | [7 empty] | F08 | Internal Control | Unknown | CY5 | 37.23 | F08 | Legionella sp. | Unknown | FAM
D10 | Internal Control | Unknown | CY5 | 36.54 | [6 empty] | F09 | Internal Control | Unknown | CY5 |  | F09 | Legionella sp. | Unknown | FAM
D10 | Legionella sp. | Unknown | FAM | [7 empty] | F10 | Internal Control | Unknown | CY5 | 36.7 | F10 | Legionella sp. | Unknown | FAM
E03 | Internal Control | Unknown | CY5 | [7 empty] | C03 | Legionella sp. | Unknown | FAM
E03 | Legionella sp. | Unknown | FAM | 28.95 | [6 empty] | C04 | Legionella sp. | Unknown | FAM
E04 | Internal Control | Unknown | CY5 | 42.62 | [6 empty] | C05 | Legionella sp. | Unknown | FAM
E04 | Legionella sp. | Unknown | FAM | 31.22 | [6 empty] | C06 | Legionella sp. | Unknown | FAM
E05 | Internal Control | Unknown | CY5 | [7 empty] | C07 | Legionella sp. | Unknown | FAM | 38.63
E05 | Legionella sp. | Unknown | FAM | 34.48 | [6 empty] | C08 | Legionella sp. | Unknown | FAM | 19.56
E06 | Internal Control | Unknown | CY5 | 38.19 | [6 empty] | C09 | Legionella sp. | Unknown | FAM | 21.58
E06 | Legionella sp. | Unknown | FAM | 35.76 | [6 empty] | C10 | Legionella sp. | Unknown | FAM
E07 | Internal Control | Unknown | CY5 | 36.66 | [6 empty] | D03 | Legionella sp. | Unknown | FAM
E07 | Legionella sp. | Unknown | FAM | [7 empty] | D04 | Legionella sp. | Unknown | FAM | 31.19
E08 | Internal Control | Unknown | CY5 | [7 empty] | D05 | Legionella sp. | Unknown | FAM | 34.37
E08 | Legionella sp. | Unknown | FAM | [7 empty] | D06 | Legionella sp. | Unknown | FAM | 37.72
E09 | Internal Control | Unknown | CY5 | 38.67 | [6 empty] | D07 | Legionella sp. | Unknown | FAM
E09 | Legionella sp. | Unknown | FAM | [7 empty] | D08 | Legionella sp. | Unknown | FAM
E10 | Internal Control | Unknown | CY5 | 36.68 | [6 empty] | D09 | Legionella sp. | Unknown | FAM
E10 | Legionella sp. | Unknown | FAM | [7 empty] | D10 | Legionella sp. | Unknown | FAM
F03 | Internal Control | Unknown | CY5 | [7 empty] | E03 | Legionella sp. | Unknown | FAM | 28.95
F03 | Legionella sp. | Unknown | FAM | 29.92 | [6 empty] | E04 | Legionella sp. | Unknown | FAM | 31.22
F04 | Internal Control | Unknown | CY5 | 42.65 | [6 empty] | E05 | Legionella sp. | Unknown | FAM | 34.48
F04 | Legionella sp. | Unknown | FAM | 30.78 | [6 empty] | E06 | Legionella sp. | Unknown | FAM | 35.76
F05 | Internal Control | Unknown | CY5 | 36.87 | [6 empty] | E07 | Legionella sp. | Unknown | FAM
F05 | Legionella sp. | Unknown | FAM | 34.8 | [6 empty] | E08 | Legionella sp. | Unknown | FAM
F06 | Internal Control | Unknown | CY5 | 37.06 | [6 empty] | E09 | Legionella sp. | Unknown | FAM
F06 | Legionella sp. | Unknown | FAM | [7 empty] | E10 | Legionella sp. | Unknown | FAM
F07 | Internal Control | Unknown | CY5 | [7 empty] | F03 | Legionella sp. | Unknown | FAM | 29.92
F07 | Legionella sp. | Unknown | FAM | [7 empty] | F04 | Legionella sp. | Unknown | FAM | 30.78
F08 | Internal Control | Unknown | CY5 | 37.23 | [6 empty] | F05 | Legionella sp. | Unknown | FAM | 34.8
F08 | Legionella sp. | Unknown | FAM | [7 empty] | F06 | Legionella sp. | Unknown | FAM
F09 | Internal Control | Unknown | CY5 | [7 empty] | F07 | Legionella sp. | Unknown | FAM
F09 | Legionella sp. | Unknown | FAM | [7 empty] | F08 | Legionella sp. | Unknown | FAM
F10 | Internal Control | Unknown | CY5 | 36.7 | [6 empty] | F09 | Legionella sp. | Unknown | FAM
F10 | Legionella sp. | Unknown | FAM | [7 empty] | F10 | Legionella sp. | Unknown | FAM
This discussion continued in analysis LOD under fold production run. Refer to that. 13 copies is incorrect. 20210706 HJK

## run5_LOD_vs_LOD
Testing the LOD of Leg spp vs Leg pneumophila (lyophilized) with Greg extracted Legionella pneumophila gDNA.
Rceived a tube from Greg and told at 1e5 copies.
Diluted in 5 tubes
Row 1 | 100000 | original | gDNA (1e5)
Row 2 | 10000 | Removed 20ul from previous tube, diluted in 180ul water | gDNA (1e4)
Row 3 | 1000 | Removed 20ul from previous tube, diluted in 180ul water | gDNA (1e3)
Row 4 | 100 | Removed 20ul from previous tube, diluted in 180ul water | gDNA (1e2)
Row 5 | 10 | Removed 20ul from previous tube, diluted in 180ul water | gDNA (1e1)
Row 6 | 1 | Removed 20ul from previous tube, diluted in 180ul water | gDNA (1e0)
Row 7 | std_6 | This was either Leg spp ultramer or Leg pneumo ultramer depending on the tubes | pos
Row 8 | NTC |  | NTC
Lyophilized Leg spp came from batch 004. This batch has not yet been QC'ed but I was unable to wait.
OT-2 pipetted into wells. Created 6 8-well strip tubes first of Leg spp. Followed by 6, 8-well strip tubes of Leg pneumophila.
ThermalCycler | Well | SampleName | Target | Type | Dye | Ct
Q96 | A03 | gDNA (1e5) | Internal Control | Leg spp | CY5 | 38.79
Q96 | A03 | gDNA (1e5) | Legionella sp. | Leg spp | FAM | 23.45
Q96 | A04 | gDNA (1e4) | Internal Control | Leg spp | CY5 | 36.53
Q96 | A04 | gDNA (1e4) | Legionella sp. | Leg spp | FAM | 25.84
Q96 | A05 | gDNA (1e3) | Internal Control | Leg spp | CY5 | 36.7
Q96 | A05 | gDNA (1e3) | Legionella sp. | Leg spp | FAM | 29.29
Q96 | A06 | gDNA (1e2) | Internal Control | Leg spp | CY5 | 36.67
Q96 | A06 | gDNA (1e2) | Legionella sp. | Leg spp | FAM | 34.9
Q96 | A07 | gDNA (1e1) | Internal Control | Leg spp | CY5
Q96 | A07 | gDNA (1e1) | Legionella sp. | Leg spp | FAM
Q96 | A08 | gDNA (1e0) | Internal Control | Leg spp | CY5
Q96 | A08 | gDNA (1e0) | Legionella sp. | Leg spp | FAM
Q96 | A09 | pos | Internal Control | Leg spp | CY5
Q96 | A09 | pos | Legionella sp. | Leg spp | FAM | 27.49
Q96 | A10 | NTC | Internal Control | Leg spp | CY5 | 36.26
Q96 | A10 | NTC | Legionella sp. | Leg spp | FAM
Q96 | B03 | gDNA (1e5) | Internal Control | Leg pneumo | CY5 | 34.71
Q96 | B03 | gDNA (1e5) | Legionella sp. | Leg pneumo | FAM | 20.44
Q96 | B04 | gDNA (1e4) | Internal Control | Leg pneumo | CY5 | 35.79
Q96 | B04 | gDNA (1e4) | Legionella sp. | Leg pneumo | FAM | 23.74
Q96 | B05 | gDNA (1e3) | Internal Control | Leg pneumo | CY5 | 33.88
Q96 | B05 | gDNA (1e3) | Legionella sp. | Leg pneumo | FAM | 27.54
Q96 | B06 | gDNA (1e2) | Internal Control | Leg pneumo | CY5 | 31.58
Q96 | B06 | gDNA (1e2) | Legionella sp. | Leg pneumo | FAM | 34.6
Q96 | B07 | gDNA (1e1) | Internal Control | Leg pneumo | CY5 | 31.6
Q96 | B07 | gDNA (1e1) | Legionella sp. | Leg pneumo | FAM
Q96 | B08 | gDNA (1e0) | Internal Control | Leg pneumo | CY5 | 31
Q96 | B08 | gDNA (1e0) | Legionella sp. | Leg pneumo | FAM
Q96 | B09 | pos | Internal Control | Leg pneumo | CY5 | 33.66
Q96 | B09 | pos | Legionella sp. | Leg pneumo | FAM | 25.6
Q96 | B10 | NTC | Internal Control | Leg pneumo | CY5 | 31.07
Q96 | B10 | NTC | Legionella sp. | Leg pneumo | FAM
Q96 | C03 | gDNA (1e5) | Internal Control | Leg spp | CY5
Q96 | C03 | gDNA (1e5) | Legionella sp. | Leg spp | FAM | 24.16
Q96 | C04 | gDNA (1e4) | Internal Control | Leg spp | CY5 | 35.59
Q96 | C04 | gDNA (1e4) | Legionella sp. | Leg spp | FAM | 24.83
Q96 | C05 | gDNA (1e3) | Internal Control | Leg spp | CY5 | 36.16
Q96 | C05 | gDNA (1e3) | Legionella sp. | Leg spp | FAM | 29.33
Q96 | C06 | gDNA (1e2) | Internal Control | Leg spp | CY5 | 38.42
Q96 | C06 | gDNA (1e2) | Legionella sp. | Leg spp | FAM | 34.39
Q96 | C07 | gDNA (1e1) | Internal Control | Leg spp | CY5
Q96 | C07 | gDNA (1e1) | Legionella sp. | Leg spp | FAM
Q96 | C08 | gDNA (1e0) | Internal Control | Leg spp | CY5
Q96 | C08 | gDNA (1e0) | Legionella sp. | Leg spp | FAM
Q96 | C09 | pos | Internal Control | Leg spp | CY5 | 35.58
Q96 | C09 | pos | Legionella sp. | Leg spp | FAM | 25.93
Q96 | C10 | NTC | Internal Control | Leg spp | CY5
Q96 | C10 | NTC | Legionella sp. | Leg spp | FAM
Q96 | D03 | gDNA (1e5) | Internal Control | Leg pneumo | CY5 | 41.15
Q96 | D03 | gDNA (1e5) | Legionella sp. | Leg pneumo | FAM | 22.16
Q96 | D04 | gDNA (1e4) | Internal Control | Leg pneumo | CY5 | 36.12
Q96 | D04 | gDNA (1e4) | Legionella sp. | Leg pneumo | FAM | 24.15
Q96 | D05 | gDNA (1e3) | Internal Control | Leg pneumo | CY5 | 32.2
Q96 | D05 | gDNA (1e3) | Legionella sp. | Leg pneumo | FAM | 27.21
Q96 | D06 | gDNA (1e2) | Internal Control | Leg pneumo | CY5 | 31.19
Q96 | D06 | gDNA (1e2) | Legionella sp. | Leg pneumo | FAM | 34.04
Q96 | D07 | gDNA (1e1) | Internal Control | Leg pneumo | CY5 | 31.33
Q96 | D07 | gDNA (1e1) | Legionella sp. | Leg pneumo | FAM | 38.99
Q96 | D08 | gDNA (1e0) | Internal Control | Leg pneumo | CY5 | 31.09
Q96 | D08 | gDNA (1e0) | Legionella sp. | Leg pneumo | FAM
Q96 | D09 | pos | Internal Control | Leg pneumo | CY5 | 33.98
Q96 | D09 | pos | Legionella sp. | Leg pneumo | FAM | 25.27
Q96 | D10 | NTC | Internal Control | Leg pneumo | CY5 | 30.99
Q96 | D10 | NTC | Legionella sp. | Leg pneumo | FAM
Q96 | E03 | gDNA (1e5) | Internal Control | Leg spp | CY5 | 39.39
Q96 | E03 | gDNA (1e5) | Legionella sp. | Leg spp | FAM | 22.98
Q96 | E04 | gDNA (1e4) | Internal Control | Leg spp | CY5 | 37.5
Q96 | E04 | gDNA (1e4) | Legionella sp. | Leg spp | FAM | 25.59
Q96 | E05 | gDNA (1e3) | Internal Control | Leg spp | CY5 | 36.37
Q96 | E05 | gDNA (1e3) | Legionella sp. | Leg spp | FAM | 28.92
Q96 | E06 | gDNA (1e2) | Internal Control | Leg spp | CY5 | 35.21
Q96 | E06 | gDNA (1e2) | Legionella sp. | Leg spp | FAM | 34.52
Q96 | E07 | gDNA (1e1) | Internal Control | Leg spp | CY5 | 37.56
Q96 | E07 | gDNA (1e1) | Legionella sp. | Leg spp | FAM
Q96 | E08 | gDNA (1e0) | Internal Control | Leg spp | CY5 | 37.04
Q96 | E08 | gDNA (1e0) | Legionella sp. | Leg spp | FAM
Q96 | E09 | pos | Internal Control | Leg spp | CY5 | 35.91
Q96 | E09 | pos | Legionella sp. | Leg spp | FAM | 25.96
Q96 | E10 | NTC | Internal Control | Leg spp | CY5
Q96 | E10 | NTC | Legionella sp. | Leg spp | FAM
Q96 | F03 | gDNA (1e5) | Internal Control | Leg pneumo | CY5 | 34.62
Q96 | F03 | gDNA (1e5) | Legionella sp. | Leg pneumo | FAM | 21.71
Q96 | F04 | gDNA (1e4) | Internal Control | Leg pneumo | CY5 | 34.79
Q96 | F04 | gDNA (1e4) | Legionella sp. | Leg pneumo | FAM | 23.86
Q96 | F05 | gDNA (1e3) | Internal Control | Leg pneumo | CY5 | 32.99
Q96 | F05 | gDNA (1e3) | Legionella sp. | Leg pneumo | FAM | 27.45
Q96 | F06 | gDNA (1e2) | Internal Control | Leg pneumo | CY5 | 30.85
Q96 | F06 | gDNA (1e2) | Legionella sp. | Leg pneumo | FAM | 34.49
Q96 | F07 | gDNA (1e1) | Internal Control | Leg pneumo | CY5 | 31.04
Q96 | F07 | gDNA (1e1) | Legionella sp. | Leg pneumo | FAM
Q96 | F08 | gDNA (1e0) | Internal Control | Leg pneumo | CY5 | 31.33
Q96 | F08 | gDNA (1e0) | Legionella sp. | Leg pneumo | FAM | 37.75
Q96 | F09 | pos | Internal Control | Leg pneumo | CY5 | 33.62
Q96 | F09 | pos | Legionella sp. | Leg pneumo | FAM | 25.41
Q96 | F10 | NTC | Internal Control | Leg pneumo | CY5 | 31.62
Q96 | F10 | NTC | Legionella sp. | Leg pneumo | FAM
Q96 | G03 | gDNA (1e5) | Internal Control | Leg spp | CY5 | 39.82
Q96 | G03 | gDNA (1e5) | Legionella sp. | Leg spp | FAM | 23.94
Q96 | G04 | gDNA (1e4) | Internal Control | Leg spp | CY5 | 37.39
Q96 | G04 | gDNA (1e4) | Legionella sp. | Leg spp | FAM | 25.53
Q96 | G05 | gDNA (1e3) | Internal Control | Leg spp | CY5 | 36.2
Q96 | G05 | gDNA (1e3) | Legionella sp. | Leg spp | FAM | 28.03
Q96 | G06 | gDNA (1e2) | Internal Control | Leg spp | CY5 | 35.72
Q96 | G06 | gDNA (1e2) | Legionella sp. | Leg spp | FAM | 34.9
Q96 | G07 | gDNA (1e1) | Internal Control | Leg spp | CY5 | 38.23
Q96 | G07 | gDNA (1e1) | Legionella sp. | Leg spp | FAM
Q96 | G08 | gDNA (1e0) | Internal Control | Leg spp | CY5 | 36.48
Q96 | G08 | gDNA (1e0) | Legionella sp. | Leg spp | FAM
Q96 | G09 | pos | Internal Control | Leg spp | CY5
Q96 | G09 | pos | Legionella sp. | Leg spp | FAM | 27.79
Q96 | G10 | NTC | Internal Control | Leg spp | CY5 | 35.94
Q96 | G10 | NTC | Legionella sp. | Leg spp | FAM
Q96 | H03 | gDNA (1e5) | Internal Control | Leg pneumo | CY5 | 38
Q96 | H03 | gDNA (1e5) | Legionella sp. | Leg pneumo | FAM | 22.22
Q96 | H04 | gDNA (1e4) | Internal Control | Leg pneumo | CY5 | 35.26
Q96 | H04 | gDNA (1e4) | Legionella sp. | Leg pneumo | FAM | 24.84
Q96 | H05 | gDNA (1e3) | Internal Control | Leg pneumo | CY5 | 32.49
Q96 | H05 | gDNA (1e3) | Legionella sp. | Leg pneumo | FAM | 27.48
Q96 | H06 | gDNA (1e2) | Internal Control | Leg pneumo | CY5 | 30.82
Q96 | H06 | gDNA (1e2) | Legionella sp. | Leg pneumo | FAM | 33.56
Q96 | H07 | gDNA (1e1) | Internal Control | Leg pneumo | CY5 | 30.91
Q96 | H07 | gDNA (1e1) | Legionella sp. | Leg pneumo | FAM
Q96 | H08 | gDNA (1e0) | Internal Control | Leg pneumo | CY5 | 31.42
Q96 | H08 | gDNA (1e0) | Legionella sp. | Leg pneumo | FAM | 40.62
Q96 | H09 | pos | Internal Control | Leg pneumo | CY5 | 33.7
Q96 | H09 | pos | Legionella sp. | Leg pneumo | FAM | 24.09
Q96 | H10 | NTC | Internal Control | Leg pneumo | CY5 | 31.11
Q96 | H10 | NTC | Legionella sp. | Leg pneumo | FAM
miniS | A01 | gDNA (1e5) | Internal Control | Leg spp | CY5 | 35.8
miniS | A01 | gDNA (1e5) | Legionella sp | Leg spp | FAM | 23.02
miniS | A02 | gDNA (1e4) | Internal Control | Leg spp | CY5 | 37.34
miniS | A02 | gDNA (1e4) | Legionella sp | Leg spp | FAM | 26.76
miniS | A03 | gDNA (1e3) | Internal Control | Leg spp | CY5 | 36.55
miniS | A03 | gDNA (1e3) | Legionella sp | Leg spp | FAM | 29.71
miniS | A04 | gDNA (1e2) | Internal Control | Leg spp | CY5 | 35.7
miniS | A04 | gDNA (1e2) | Legionella sp | Leg spp | FAM | 35.3
miniS | A05 | gDNA (1e1) | Internal Control | Leg spp | CY5 | 36.78
miniS | A05 | gDNA (1e1) | Legionella sp | Leg spp | FAM | 40.65
miniS | A06 | gDNA (1e0) | Internal Control | Leg spp | CY5 | 37.59
miniS | A06 | gDNA (1e0) | Legionella sp | Leg spp | FAM
miniS | A07 | pos | Internal Control | Leg spp | CY5 | 36.37
miniS | A07 | pos | Legionella sp | Leg spp | FAM | 26.94
miniS | A08 | NTC | Internal Control | Leg spp | CY5 | 36.13
miniS | A08 | NTC | Legionella sp | Leg spp | FAM
miniS | B01 | gDNA (1e5) | Internal Control | Leg pneumo | CY5 | 34.21
miniS | B01 | gDNA (1e5) | Legionella sp | Leg pneumo | FAM | 20.42
miniS | B02 | gDNA (1e4) | Internal Control | Leg pneumo | CY5 | 35.41
miniS | B02 | gDNA (1e4) | Legionella sp | Leg pneumo | FAM | 24.93
miniS | B03 | gDNA (1e3) | Internal Control | Leg pneumo | CY5 | 32.21
miniS | B03 | gDNA (1e3) | Legionella sp | Leg pneumo | FAM | 27.62
miniS | B04 | gDNA (1e2) | Internal Control | Leg pneumo | CY5 | 31.44
miniS | B04 | gDNA (1e2) | Legionella sp | Leg pneumo | FAM | 34.99
miniS | B05 | gDNA (1e1) | Internal Control | Leg pneumo | CY5 | 31.16
miniS | B05 | gDNA (1e1) | Legionella sp | Leg pneumo | FAM | 36.63
miniS | B06 | gDNA (1e0) | Internal Control | Leg pneumo | CY5 | 30.9
miniS | B06 | gDNA (1e0) | Legionella sp | Leg pneumo | FAM | 38.05
miniS | B07 | pos | Internal Control | Leg pneumo | CY5 | 32.68
miniS | B07 | pos | Legionella sp | Leg pneumo | FAM | 25.49
miniS | B08 | NTC | Internal Control | Leg pneumo | CY5 | 31.24
miniS | B08 | NTC | Legionella sp | Leg pneumo | FAM | 38.39
Q16 | A1-1 | gDNA (1e5) | Legionella sp. | Leg spp | FAM | 24.45
Q16 | A1-4 | gDNA (1e5) | Internal Control | Leg spp | CY5
Q16 | A2-1 | gDNA (1e4) | Legionella sp. | Leg spp | FAM | 26.95
Q16 | A2-4 | gDNA (1e4) | Internal Control | Leg spp | CY5 | 34.91
Q16 | A3-1 | gDNA (1e3) | Legionella sp. | Leg spp | FAM | 30.57
Q16 | A3-4 | gDNA (1e3) | Internal Control | Leg spp | CY5 | 34.59
Q16 | A4-1 | gDNA (1e2) | Legionella sp. | Leg spp | FAM | 36.75
Q16 | A4-4 | gDNA (1e2) | Internal Control | Leg spp | CY5
Q16 | A5-1 | gDNA (1e1) | Legionella sp. | Leg spp | FAM
Q16 | A5-4 | gDNA (1e1) | Internal Control | Leg spp | CY5
Q16 | A6-1 | gDNA (1e0) | Legionella sp. | Leg spp | FAM
Q16 | A6-4 | gDNA (1e0) | Internal Control | Leg spp | CY5
Q16 | A7-1 | pos | Legionella sp. | Leg spp | FAM
Q16 | A7-4 | pos | Internal Control | Leg spp | CY5
Q16 | A8-1 | NTC | Legionella sp. | Leg spp | FAM
Q16 | A8-4 | NTC | Internal Control | Leg spp | CY5
Q16 | B1-1 | gDNA (1e5) | Legionella sp. | Leg pneumo | FAM | 21.63
Q16 | B1-4 | gDNA (1e5) | Internal Control | Leg pneumo | CY5 | 39.4
Q16 | B2-1 | gDNA (1e4) | Legionella sp. | Leg pneumo | FAM | 26.11
Q16 | B2-4 | gDNA (1e4) | Internal Control | Leg pneumo | CY5 | 34.57
Q16 | B3-1 | gDNA (1e3) | Legionella sp. | Leg pneumo | FAM | 29.8
Q16 | B3-4 | gDNA (1e3) | Internal Control | Leg pneumo | CY5 | 33.76
Q16 | B4-1 | gDNA (1e2) | Legionella sp. | Leg pneumo | FAM | 36.59
Q16 | B4-4 | gDNA (1e2) | Internal Control | Leg pneumo | CY5 | 32.96
Q16 | B5-1 | gDNA (1e1) | Legionella sp. | Leg pneumo | FAM
Q16 | B5-4 | gDNA (1e1) | Internal Control | Leg pneumo | CY5 | 32.86
Q16 | B6-1 | gDNA (1e0) | Legionella sp. | Leg pneumo | FAM | 39.67
Q16 | B6-4 | gDNA (1e0) | Internal Control | Leg pneumo | CY5 | 33.12
Q16 | B7-1 | pos | Legionella sp. | Leg pneumo | FAM | 28.12
Q16 | B7-4 | pos | Internal Control | Leg pneumo | CY5 | 34.44
Q16 | B8-1 | NTC | Legionella sp. | Leg pneumo | FAM
Q16 | B8-4 | NTC | Internal Control | Leg pneumo | CY5 | 34.14

## Pne_vs_Spp LOD
Dye | FAM
Average of Ct | Column Labels
Row Labels | gDNA (1e5) | gDNA (1e4) | gDNA (1e3) | gDNA (1e2) | gDNA (1e1) | gDNA (1e0) | Grand Total
Leg pneumo | 21.43 | 24.605 | 27.85 | 34.7117 | 37.81 | 39.0225 | 29.443
miniS | 20.42 | 24.93 | 27.62 | 34.99 | 36.63 | 38.05 | 30.44
Q16 | 21.63 | 26.11 | 29.8 | 36.59 |  | 39.67 | 30.76
Q96 | 21.6325 | 24.1475 | 27.42 | 34.1725 | 38.99 | 39.185 | 28.7816
Leg spp | 23.6667 | 25.9167 | 29.3083 | 35.1267 | 40.65 | #DIV/0! | 28.9904
miniS | 23.02 | 26.76 | 29.71 | 35.3 | 40.65 | #DIV/0! | 31.088
Q16 | 24.45 | 26.95 | 30.57 | 36.75 |  |  | 29.68
Q96 | 23.6325 | 25.4475 | 28.8925 | 34.6775 | #DIV/0! | #DIV/0! | 28.1625
Grand Total | 22.5483 | 25.2608 | 28.5792 | 34.9192 | 38.7567 | 39.0225 | 29.2373

## Pne_vs_Spp LOD IAC (2)
Dye | CY5
Max of Ct | Column Labels
Row Labels | gDNA (1e5) | gDNA (1e4) | gDNA (1e3) | gDNA (1e2) | gDNA (1e1) | gDNA (1e0) | Grand Total
Leg pneumo | 41.15 | 36.12 | 33.88 | 32.96 | 32.86 | 33.12 | 41.15
miniS | 34.21 | 35.41 | 32.21 | 31.44 | 31.16 | 30.9 | 35.41
Q16 | 39.4 | 34.57 | 33.76 | 32.96 | 32.86 | 33.12 | 39.4
Q96 | 41.15 | 36.12 | 33.88 | 31.58 | 31.6 | 31.42 | 41.15
Leg spp | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.59 | 39.82
miniS | 35.8 | 37.34 | 36.55 | 35.7 | 36.78 | 37.59 | 37.59
Q16 |  | 34.91 | 34.59 |  |  |  | 34.91
Q96 | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.04 | 39.82
Grand Total | 41.15 | 37.5 | 36.7 | 38.42 | 38.23 | 37.59 | 41.15

## Pne_vs_Spp LOD IAC
Dye | CY5
Max of Ct | Column Labels
Row Labels | gDNA (1e5) | gDNA (1e4) | gDNA (1e3) | gDNA (1e2) | gDNA (1e1) | gDNA (1e0) | Grand Total
Leg pneumo | 41.15 | 36.12 | 33.88 | 32.96 | 32.86 | 33.12 | 41.15
miniS | 34.21 | 35.41 | 32.21 | 31.44 | 31.16 | 30.9 | 35.41
Q16 | 39.4 | 34.57 | 33.76 | 32.96 | 32.86 | 33.12 | 39.4
Q96 | 41.15 | 36.12 | 33.88 | 31.58 | 31.6 | 31.42 | 41.15
Leg spp | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.59 | 39.82
miniS | 35.8 | 37.34 | 36.55 | 35.7 | 36.78 | 37.59 | 37.59
Q16 |  | 34.91 | 34.59 |  |  |  | 34.91
Q96 | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.04 | 39.82
Grand Total | 41.15 | 37.5 | 36.7 | 38.42 | 38.23 | 37.59 | 41.15

## LOQ
Well | Target | Property | copies/rxn | log(copies/rxn) | Set | Replicate | Dye | Ct | Well | Target | Property | copies/rxn | log(copies/rxn) | Set | Replicate | Dye | Ct
A01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A01 | Internal Control | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 1 | CY5 | 38.03
A01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | A12 | Internal Control | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 1 | CY5 | 37.81
A02 | Internal Control | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 1 | CY5 | 38.03 | B01 | Internal Control | std_3 | 120440000 | 8.0808 | stds1_7 | rep 1 | CY5 | 34.51
A02 | Legionella sp. | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 1 | FAM |  | B12 | Internal Control | std_4 | 12044000 | 7.0808 | stds1_7 | rep 1 | CY5 | 33.81
A03 | Internal Control | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 2 | CY5 |  | C01 | Internal Control | std_5 | 1204400 | 6.0808 | stds1_7 | rep 1 | CY5 | 35.33
A03 | Legionella sp. | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 2 | FAM |  | C12 | Internal Control | std_6 | 120440 | 5.0808 | stds1_7 | rep 1 | CY5 | 33.18
A04 | Internal Control | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 3 | CY5 | 36.21 | D01 | Internal Control | std_7 | 12044 | 4.0808 | stds1_7 | rep 1 | CY5 | 35.26
A04 | Legionella sp. | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 3 | FAM |  | D12 | Internal Control | NTC | 0 | #NUM! | stds1_7 | rep 1 | CY5 | 30.49
A05 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 1 | CY5 | 35.72 | E01 | Internal Control | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 2 | CY5
A05 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 1 | FAM | 11.97 | E12 | Internal Control | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 2 | CY5 | 33.85
A06 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 2 | CY5 | 36.95 | F01 | Internal Control | std_3 | 120440000 | 8.0808 | stds1_7 | rep 2 | CY5 | 35.93
A06 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 2 | FAM | 11.92 | F12 | Internal Control | std_4 | 12044000 | 7.0808 | stds1_7 | rep 2 | CY5 | 34.89
A07 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 3 | CY5 | 40.04 | G01 | Internal Control | std_5 | 1204400 | 6.0808 | stds1_7 | rep 2 | CY5 | 33.51
A07 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 3 | FAM | 12.73 | G12 | Internal Control | std_6 | 120440 | 5.0808 | stds1_7 | rep 2 | CY5 | 32.06
A08 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 4 | CY5 | 33.76 | H01 | Internal Control | std_7 | 12044 | 4.0808 | stds1_7 | rep 2 | CY5 | 31.95
A08 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 4 | FAM | 10.07 | H12 | Internal Control | NTC | 0 | #NUM! | stds1_7 | rep 2 | CY5 | 30.29
A09 | Internal Control | std_3 | 120440000 | 8.0808 | stds3_9 | rep 1 | CY5 | 37.67 | A02 | Internal Control | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 3 | CY5 | 36.21
A09 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds3_9 | rep 1 | FAM | 18.27 | B02 | Internal Control | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 3 | CY5 | 35.35
A10 | Internal Control | std_3 | 120440000 | 8.0808 | stds3_9 | rep 2 | CY5 |  | C02 | Internal Control | std_3 | 120440000 | 8.0808 | stds1_7 | rep 3 | CY5 | 36.35
A10 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds3_9 | rep 2 | FAM | 14.01 | D02 | Internal Control | std_4 | 12044000 | 7.0808 | stds1_7 | rep 3 | CY5 | 35.13
A11 | Internal Control | std_3 | 120440000 | 8.0808 | stds3_9 | rep 3 | CY5 | 38.2 | E02 | Internal Control | std_5 | 1204400 | 6.0808 | stds1_7 | rep 3 | CY5 | 35.15
A11 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds3_9 | rep 3 | FAM | 17.09 | F02 | Internal Control | std_6 | 120440 | 5.0808 | stds1_7 | rep 3 | CY5 | 32.63
A12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | G02 | Internal Control | std_7 | 12044 | 4.0808 | stds1_7 | rep 3 | CY5 | 33.49
A12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H02 | Internal Control | NTC | 0 | #NUM! | stds1_7 | rep 3 | CY5 | 30.9
B01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A03 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 1 | CY5 | 35.72
B01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | B03 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 1 | CY5 | 34.52
B02 | Internal Control | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 1 | CY5 | 37.81 | C03 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 1 | CY5 | 34.67
B02 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 1 | FAM | 6.63 | D03 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 1 | CY5 | 32.98
B03 | Internal Control | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 2 | CY5 | 33.85 | E03 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 1 | CY5 | 33.64
B03 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 2 | FAM | 11.5 | F03 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 1 | CY5 | 30.4
B04 | Internal Control | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 3 | CY5 | 35.35 | G03 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 1 | CY5 | 30.52
B04 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 3 | FAM | 9.45 | H03 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 1 | CY5 | 30.55
B05 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 1 | CY5 | 34.52 | A04 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 2 | CY5 | 36.95
B05 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 1 | FAM | 15.06 | B04 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 2 | CY5 | 39.92
B06 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 2 | CY5 | 39.92 | C04 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 2 | CY5
B06 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 2 | FAM | 14.48 | D04 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 2 | CY5 | 35.1
B07 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 3 | CY5 |  | E04 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 2 | CY5 | 32.42
B07 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 3 | FAM | 15.87 | F04 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 2 | CY5 | 30.56
B08 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 4 | CY5 | 37.88 | G04 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 2 | CY5 | 30.89
B08 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 4 | FAM | 14.32 | H04 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 2 | CY5 | 30.13
B09 | Internal Control | std_4 | 12044000 | 7.0808 | stds3_9 | rep 1 | CY5 | 32.7 | A05 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 3 | CY5 | 40.04
B09 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds3_9 | rep 1 | FAM | 18.3 | B05 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 3 | CY5
B10 | Internal Control | std_4 | 12044000 | 7.0808 | stds3_9 | rep 2 | CY5 | 36.25 | C05 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 3 | CY5 | 38.43
B10 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds3_9 | rep 2 | FAM | 17.26 | D05 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 3 | CY5 | 35.84
B11 | Internal Control | std_4 | 12044000 | 7.0808 | stds3_9 | rep 3 | CY5 |  | E05 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 3 | CY5 | 42.55
B11 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds3_9 | rep 3 | FAM | 19.43 | F05 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 3 | CY5 | 32.67
B12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | G05 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 3 | CY5 | 30.15
B12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H05 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 3 | CY5 | 30.53
C01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A06 | Internal Control | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 4 | CY5 | 33.76
C01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | B06 | Internal Control | std_3 | 120440000 | 8.0808 | stds2_8 | rep 4 | CY5 | 37.88
C02 | Internal Control | std_3 | 120440000 | 8.0808 | stds1_7 | rep 1 | CY5 | 34.51 | C06 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 4 | CY5 | 33.95
C02 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds1_7 | rep 1 | FAM | 13.22 | D06 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 4 | CY5 | 38.66
C03 | Internal Control | std_3 | 120440000 | 8.0808 | stds1_7 | rep 2 | CY5 | 35.93 | E06 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 4 | CY5 | 34.4
C03 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds1_7 | rep 2 | FAM | 13.45 | F06 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 4 | CY5 | 31.09
C04 | Internal Control | std_3 | 120440000 | 8.0808 | stds1_7 | rep 3 | CY5 | 36.35 | G06 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 4 | CY5 | 30.03
C04 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds1_7 | rep 3 | FAM | 13.49 | H06 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 4 | CY5 | 30.87
C05 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 1 | CY5 | 34.67 | A07 | Internal Control | std_3 | 120440000 | 8.0808 | stds3_9 | rep 1 | CY5 | 37.67
C05 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 1 | FAM | 19.42 | B07 | Internal Control | std_4 | 12044000 | 7.0808 | stds3_9 | rep 1 | CY5 | 32.7
C06 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 2 | CY5 |  | C07 | Internal Control | std_5 | 1204400 | 6.0808 | stds3_9 | rep 1 | CY5
C06 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 2 | FAM | 16.84 | D07 | Internal Control | std_6 | 120440 | 5.0808 | stds3_9 | rep 1 | CY5 | 31.93
C07 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 3 | CY5 | 38.43 | E07 | Internal Control | std_7 | 12044 | 4.0808 | stds3_9 | rep 1 | CY5 | 31.2
C07 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 3 | FAM | 19.23 | F07 | Internal Control | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 1 | CY5 | 30.94
C08 | Internal Control | std_4 | 12044000 | 7.0808 | stds2_8 | rep 4 | CY5 | 33.95 | G07 | Internal Control | std_9 | 120.44 | 2.0808 | stds3_9 | rep 1 | CY5 | 31.36
C08 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 4 | FAM | 19.08 | H07 | Internal Control | NTC | 0 | #NUM! | stds3_9 | rep 1 | CY5 | 30.14
C09 | Internal Control | std_5 | 1204400 | 6.0808 | stds3_9 | rep 1 | CY5 |  | A08 | Internal Control | std_3 | 120440000 | 8.0808 | stds3_9 | rep 2 | CY5
C09 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds3_9 | rep 1 | FAM | 21.87 | B08 | Internal Control | std_4 | 12044000 | 7.0808 | stds3_9 | rep 2 | CY5 | 36.25
C10 | Internal Control | std_5 | 1204400 | 6.0808 | stds3_9 | rep 2 | CY5 | 34.99 | C08 | Internal Control | std_5 | 1204400 | 6.0808 | stds3_9 | rep 2 | CY5 | 34.99
C10 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds3_9 | rep 2 | FAM | 20.7 | D08 | Internal Control | std_6 | 120440 | 5.0808 | stds3_9 | rep 2 | CY5 | 36.12
C11 | Internal Control | std_5 | 1204400 | 6.0808 | stds3_9 | rep 3 | CY5 | 36.94 | E08 | Internal Control | std_7 | 12044 | 4.0808 | stds3_9 | rep 2 | CY5 | 31.36
C11 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds3_9 | rep 3 | FAM | 21.93 | F08 | Internal Control | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 2 | CY5 | 30.36
C12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | G08 | Internal Control | std_9 | 120.44 | 2.0808 | stds3_9 | rep 2 | CY5 | 32.07
C12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H08 | Internal Control | NTC | 0 | #NUM! | stds3_9 | rep 2 | CY5 | 30.19
D01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A09 | Internal Control | std_3 | 120440000 | 8.0808 | stds3_9 | rep 3 | CY5 | 38.2
D01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | B09 | Internal Control | std_4 | 12044000 | 7.0808 | stds3_9 | rep 3 | CY5
D02 | Internal Control | std_4 | 12044000 | 7.0808 | stds1_7 | rep 1 | CY5 | 33.81 | C09 | Internal Control | std_5 | 1204400 | 6.0808 | stds3_9 | rep 3 | CY5 | 36.94
D02 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds1_7 | rep 1 | FAM | 17.21 | D09 | Internal Control | std_6 | 120440 | 5.0808 | stds3_9 | rep 3 | CY5 | 32.61
D03 | Internal Control | std_4 | 12044000 | 7.0808 | stds1_7 | rep 2 | CY5 | 34.89 | E09 | Internal Control | std_7 | 12044 | 4.0808 | stds3_9 | rep 3 | CY5 | 33.02
D03 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds1_7 | rep 2 | FAM | 16.96 | F09 | Internal Control | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 3 | CY5 | 31.28
D04 | Internal Control | std_4 | 12044000 | 7.0808 | stds1_7 | rep 3 | CY5 | 35.13 | G09 | Internal Control | std_9 | 120.44 | 2.0808 | stds3_9 | rep 3 | CY5 | 32.8
D04 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds1_7 | rep 3 | FAM | 17.44 | H09 | Internal Control | NTC | 0 | #NUM! | stds3_9 | rep 3 | CY5 | 29.94
D05 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 1 | CY5 | 32.98 | A01 | Legionella sp. | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 1 | FAM |  | Stds | Set | copy#/rxn | log(copy#) | rep 1 | rep 2 | rep 3 |  | avg | stdev | 2*sd
D05 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 1 | FAM | 22.36 | A12 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 1 | FAM | 6.63 | std_1 | stds1_7 | 12044000000 | 10.0808 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
D06 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 2 | CY5 | 35.1 | B01 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds1_7 | rep 1 | FAM | 13.22 | std_2 | stds1_7 | 1204400000 | 9.0808 | 6.63 | 11.5 | 9.45 |  | 9.1933 | 2.4451 | 4.8902
D06 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 2 | FAM | 22.58 | B12 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds1_7 | rep 1 | FAM | 17.21 | std_3 | stds1_7 | 120440000 | 8.0808 | 13.22 | 13.45 | 13.49 |  | 13.3867 | 0.14572 | 0.29143
D07 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 3 | CY5 | 35.84 | C01 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds1_7 | rep 1 | FAM | 20.8 | std_4 | stds1_7 | 12044000 | 7.0808 | 17.21 | 16.96 | 17.44 |  | 17.2033 | 0.24007 | 0.48014
D07 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 3 | FAM | 21.01 | C12 | Legionella sp. | std_6 | 120440 | 5.0808 | stds1_7 | rep 1 | FAM | 24.08 | std_5 | stds1_7 | 1204400 | 6.0808 | 20.8 | 21.38 | 20.89 |  | 21.0233 | 0.31214 | 0.62429
D08 | Internal Control | std_5 | 1204400 | 6.0808 | stds2_8 | rep 4 | CY5 | 38.66 | D01 | Legionella sp. | std_7 | 12044 | 4.0808 | stds1_7 | rep 1 | FAM | 29.08 | std_6 | stds1_7 | 120440 | 5.0808 | 24.08 | 24.46 | 24.18 |  | 24.24 | 0.19698 | 0.39395
D08 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 4 | FAM | 22.73 | D12 | Legionella sp. | NTC | 0 | #NUM! | stds1_7 | rep 1 | FAM | 37.88 | std_7 | stds1_7 | 12044 | 4.0808 | 29.08 | 28.12 | 29.38 |  | 28.86 | 0.65818 | 1.3164
D09 | Internal Control | std_6 | 120440 | 5.0808 | stds3_9 | rep 1 | CY5 | 31.93 | E01 | Legionella sp. | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 2 | FAM
D09 | Legionella sp. | std_6 | 120440 | 5.0808 | stds3_9 | rep 1 | FAM | 24.85 | E12 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 2 | FAM | 11.5
D10 | Internal Control | std_6 | 120440 | 5.0808 | stds3_9 | rep 2 | CY5 | 36.12 | F01 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds1_7 | rep 2 | FAM | 13.45
D10 | Legionella sp. | std_6 | 120440 | 5.0808 | stds3_9 | rep 2 | FAM | 24.71 | F12 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds1_7 | rep 2 | FAM | 16.96
D11 | Internal Control | std_6 | 120440 | 5.0808 | stds3_9 | rep 3 | CY5 | 32.61 | G01 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds1_7 | rep 2 | FAM | 21.38
D11 | Legionella sp. | std_6 | 120440 | 5.0808 | stds3_9 | rep 3 | FAM | 24.11 | G12 | Legionella sp. | std_6 | 120440 | 5.0808 | stds1_7 | rep 2 | FAM | 24.46
D12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | H01 | Legionella sp. | std_7 | 12044 | 4.0808 | stds1_7 | rep 2 | FAM | 28.12
D12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H12 | Legionella sp. | NTC | 0 | #NUM! | stds1_7 | rep 2 | FAM
E01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A02 | Legionella sp. | std_1 | 12044000000 | 10.0808 | stds1_7 | rep 3 | FAM | [14 empty] | Leg pneumo seq
E01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | B02 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds1_7 | rep 3 | FAM | 9.45 | [13 empty] | CCGGATCATTATTAACCATCACCTGTGCCACCGTAGATAAAGTCACATACATTTGCTGTATAGCAATCGCGCGAGACAAT
E02 | Internal Control | std_5 | 1204400 | 6.0808 | stds1_7 | rep 1 | CY5 | 35.33 | C02 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds1_7 | rep 3 | FAM | 13.49 | [13 empty] | Leg pneumo pos control quantified with spec on 20210408
E02 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds1_7 | rep 1 | FAM | 20.8 | D02 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds1_7 | rep 3 | FAM | 17.44 | [14 empty] | abs@260 | Ext Co | conc(uM) | dil factor
E03 | Internal Control | std_5 | 1204400 | 6.0808 | stds1_7 | rep 2 | CY5 | 33.51 | E02 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds1_7 | rep 3 | FAM | 20.89 | [13 empty] | undiluted | 69.8944 | 779300 | 89.6887 |  |  | 92.3549
E03 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds1_7 | rep 2 | FAM | 21.38 | F02 | Legionella sp. | std_6 | 120440 | 5.0808 | stds1_7 | rep 3 | FAM | 24.18 | [13 empty] | std_0 | 0.7405 | 779300 | 0.95021 | 95.0212
E04 | Internal Control | std_5 | 1204400 | 6.0808 | stds1_7 | rep 3 | CY5 | 35.15 | G02 | Legionella sp. | std_7 | 12044 | 4.0808 | stds1_7 | rep 3 | FAM | 29.38 |  |  |  |  | Replicate
E04 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds1_7 | rep 3 | FAM | 20.89 | H02 | Legionella sp. | NTC | 0 | #NUM! | stds1_7 | rep 3 | FAM | 37.01 |  | Set | copy#/rxn | log(copy#) | rep 1 | rep 2 | rep 3 | rep 4 | avg | stdev | 2*sd | CV | Name | TubeName | Conc (uM) | ng | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul)
E05 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 1 | CY5 | 33.64 | A03 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 1 | FAM | 11.97 |  | stds2_8 | 1204400000 | 9.0808 | 11.97 | 11.92 | 12.73 | 10.07 | 11.6725 | 1.1308 | 2.2616 | 9.688% | Tube_2 | IDT | 92.35 | 122000 |  | 2.2245e+15
E05 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 1 | FAM | 26.78 | B03 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 1 | FAM | 15.06 |  | stds2_8 | 120440000 | 8.0808 | 15.06 | 14.48 | 15.87 | 14.32 | 14.9325 | 0.70121 | 1.4024 | 4.696% | Tube_3
E06 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 2 | CY5 | 32.42 | C03 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 1 | FAM | 19.42 |  | stds2_8 | 12044000 | 7.0808 | 19.42 | 16.84 | 19.23 | 19.08 | 18.6425 | 1.2097 | 2.4194 | 6.489% | Tube_4 | Tube_0 | 0.9235 | 30500 |  | 5.5613e+14
E06 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 2 | FAM | 26.59 | D03 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 1 | FAM | 22.36 |  | stds2_8 | 1204400 | 6.0808 | 22.36 | 22.58 | 21.01 | 22.73 | 22.17 | 0.78812 | 1.5762 | 3.555% | Tube_5
E07 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 3 | CY5 | 42.55 | E03 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 1 | FAM | 26.78 |  | stds2_8 | 120440 | 5.0808 | 26.78 | 26.59 | 24.98 | 28.13 | 26.62 | 1.2905 | 2.581 | 4.848% | Tube_6 | Tube_1 | 0.009235 | 305 | 6.2742e+12 | 5.5613e+12 | 11122634000
E07 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 3 | FAM | 24.98 | F03 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 1 | FAM | 28.39 |  | stds2_8 | 12044 | 4.0808 | 28.39 | 29.6 | 29.92 | 31.06 | 29.7425 | 1.098 | 2.196 | 3.692% | Tube_7 | Tube_2 | 0.0009235 | 30.5 | 627418093126.6887 | 556131700000.0001 | 1112263400
E08 | Internal Control | std_6 | 120440 | 5.0808 | stds2_8 | rep 4 | CY5 | 34.4 | G03 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 1 | FAM | 33.37 |  | stds2_8 | 1204.4 | 3.0808 | 33.37 | 33.91 | 34.5 | 34.64 | 34.105 | 0.58324 | 1.1665 | 1.71% | Tube_8 | Tube_3 | 9.235e-05 | 3.05 | 62741809312.6689 | 55613170000 | 111226340
E08 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 4 | FAM | 28.13 | H03 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 1 | FAM |  |  | stds2_8 | 0 | #NUM! | [9 empty] | Tube_4 | 9.235e-06 | 0.305 | 6274180931.2669 | 5561317000 | 11122634
E09 | Internal Control | std_7 | 12044 | 4.0808 | stds3_9 | rep 1 | CY5 | 31.2 | A04 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 2 | FAM | 11.92 | [13 empty] | Tube_5 | 9.235e-07 | 0.0305 | 627418093.1267 | 556131700 | 1112263.4
E09 | Legionella sp. | std_7 | 12044 | 4.0808 | stds3_9 | rep 1 | FAM | 28.77 | B04 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 2 | FAM | 14.48 | [5 empty] | slope | y-int | [6 empty] | Tube_6 | 9.235e-08 | 0.00305 | 62741809.3127 | 55613170 | 111226.34
E10 | Internal Control | std_7 | 12044 | 4.0808 | stds3_9 | rep 2 | CY5 | 31.36 | C04 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 2 | FAM | 16.84 | [5 empty] | -3.7463 | 45.3351 | [6 empty] | Tube_7 | 9.235e-09 | 0.000305 | 6274180.9313 | 5561317 | 11122.634
E10 | Legionella sp. | std_7 | 12044 | 4.0808 | stds3_9 | rep 2 | FAM | 28.17 | D04 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 2 | FAM | 22.58 | [13 empty] | Tube_8 | 9.235e-10 | 3.05e-05 | 627418.0931 | 556131.7 | 1112.2634
E11 | Internal Control | std_7 | 12044 | 4.0808 | stds3_9 | rep 3 | CY5 | 33.02 | E04 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 2 | FAM | 26.59 |  |  |  |  | PCR Efficiency | 84.9% | [7 empty] | Tube_9 | 9.235e-11 | 3.05e-06 | 62741.8093 | 55613.17 | 111.2263
E11 | Legionella sp. | std_7 | 12044 | 4.0808 | stds3_9 | rep 3 | FAM | 28.99 | F04 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 2 | FAM | 29.6 | [13 empty] | Tube_10 | 9.235e-12 | 3.05e-07 | 6274.1809 | 5561.317 | 11.1226
E12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | G04 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 2 | FAM | 33.91 | [13 empty] | Tube_11 | 9.235e-13 | 3.05e-08 | 627.4181 | 556.1317 | 1.1123
E12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H04 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 2 | FAM | [14 empty] | Tube_12 | 9.235e-14 | 3.05e-09 | 62.7418 | 55.6132 | 0.11123
F01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A05 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 3 | FAM | 12.73 | [13 empty] | Tube_13 | 9.235e-15 | 3.05e-10 | 6.2742 | 5.5613 | 0.011123
F01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | B05 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 3 | FAM | 15.87 | [13 empty] | Tube_14 | 9.235e-16 | 3.05e-11 | 0.62742 | 0.55613 | 0.0011123
F02 | Internal Control | std_6 | 120440 | 5.0808 | stds1_7 | rep 1 | CY5 | 33.18 | C05 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 3 | FAM | 19.23 | [13 empty] | Tube_15 | 9.235e-17 | 3.05e-12 | 0.062742 | 0.055613 | 0.00011123
F02 | Legionella sp. | std_6 | 120440 | 5.0808 | stds1_7 | rep 1 | FAM | 24.08 | D05 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 3 | FAM | 21.01
F03 | Internal Control | std_6 | 120440 | 5.0808 | stds1_7 | rep 2 | CY5 | 32.06 | E05 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 3 | FAM | 24.98
F03 | Legionella sp. | std_6 | 120440 | 5.0808 | stds1_7 | rep 2 | FAM | 24.46 | F05 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 3 | FAM | 29.92
F04 | Internal Control | std_6 | 120440 | 5.0808 | stds1_7 | rep 3 | CY5 | 32.63 | G05 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 3 | FAM | 34.5
F04 | Legionella sp. | std_6 | 120440 | 5.0808 | stds1_7 | rep 3 | FAM | 24.18 | H05 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 3 | FAM
F05 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 1 | CY5 | 30.4 | A06 | Legionella sp. | std_2 | 1204400000 | 9.0808 | stds2_8 | rep 4 | FAM | 10.07
F05 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 1 | FAM | 28.39 | B06 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds2_8 | rep 4 | FAM | 14.32
F06 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 2 | CY5 | 30.56 | C06 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds2_8 | rep 4 | FAM | 19.08
F06 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 2 | FAM | 29.6 | D06 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds2_8 | rep 4 | FAM | 22.73
F07 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 3 | CY5 | 32.67 | E06 | Legionella sp. | std_6 | 120440 | 5.0808 | stds2_8 | rep 4 | FAM | 28.13
F07 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 3 | FAM | 29.92 | F06 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 4 | FAM | 31.06
F08 | Internal Control | std_7 | 12044 | 4.0808 | stds2_8 | rep 4 | CY5 | 31.09 | G06 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 4 | FAM | 34.64
F08 | Legionella sp. | std_7 | 12044 | 4.0808 | stds2_8 | rep 4 | FAM | 31.06 | H06 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 4 | FAM
F09 | Internal Control | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 1 | CY5 | 30.94 | A07 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds3_9 | rep 1 | FAM | 18.27 |  |  |  |  | Replicate
F09 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 1 | FAM | 31.73 | B07 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds3_9 | rep 1 | FAM | 18.3 | Stds | Set | copy#/rxn | log(copy#) | rep 1 | rep 2 | rep 3 |  | avg | stdev | 2*sd
F10 | Internal Control | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 2 | CY5 | 30.36 | C07 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds3_9 | rep 1 | FAM | 21.87 | std_3 | stds3_9 | 120440000 | 8.0808 | 18.27 | 14.01 | 17.09 |  | 16.4567 | 2.1995 | 4.399
F10 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 2 | FAM | 31.48 | D07 | Legionella sp. | std_6 | 120440 | 5.0808 | stds3_9 | rep 1 | FAM | 24.85 | std_4 | stds3_9 | 12044000 | 7.0808 | 18.3 | 17.26 | 19.43 |  | 18.33 | 1.0853 | 2.1706
F11 | Internal Control | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 3 | CY5 | 31.28 | E07 | Legionella sp. | std_7 | 12044 | 4.0808 | stds3_9 | rep 1 | FAM | 28.77 | std_5 | stds3_9 | 1204400 | 6.0808 | 21.87 | 20.7 | 21.93 |  | 21.5 | 0.69347 | 1.3869
F11 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 3 | FAM | 31.76 | F07 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 1 | FAM | 31.73 | std_6 | stds3_9 | 120440 | 5.0808 | 24.85 | 24.71 | 24.11 |  | 24.5567 | 0.39311 | 0.78621
F12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | G07 | Legionella sp. | std_9 | 120.44 | 2.0808 | stds3_9 | rep 1 | FAM | 32.69 | std_7 | stds3_9 | 12044 | 4.0808 | 28.77 | 28.17 | 28.99 |  | 28.6433 | 0.42442 | 0.84884
F12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H07 | Legionella sp. | NTC | 0 | #NUM! | stds3_9 | rep 1 | FAM | 39.78 | std_8 | stds3_9 | 1204.4 | 3.0808 | 31.73 | 31.48 | 31.76 |  | 31.6567 | 0.15373 | 0.30746
G01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | A08 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds3_9 | rep 2 | FAM | 14.01 | std_9 | stds3_9 | 120.44 | 2.0808 | 32.69 | 33.01 | 33.88 |  | 33.1933 | 0.61582 | 1.2316
G01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | B08 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds3_9 | rep 2 | FAM | 17.26
G02 | Internal Control | std_7 | 12044 | 4.0808 | stds1_7 | rep 1 | CY5 | 35.26 | C08 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds3_9 | rep 2 | FAM | 20.7
G02 | Legionella sp. | std_7 | 12044 | 4.0808 | stds1_7 | rep 1 | FAM | 29.08 | D08 | Legionella sp. | std_6 | 120440 | 5.0808 | stds3_9 | rep 2 | FAM | 24.71
G03 | Internal Control | std_7 | 12044 | 4.0808 | stds1_7 | rep 2 | CY5 | 31.95 | E08 | Legionella sp. | std_7 | 12044 | 4.0808 | stds3_9 | rep 2 | FAM | 28.17
G03 | Legionella sp. | std_7 | 12044 | 4.0808 | stds1_7 | rep 2 | FAM | 28.12 | F08 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 2 | FAM | 31.48
G04 | Internal Control | std_7 | 12044 | 4.0808 | stds1_7 | rep 3 | CY5 | 33.49 | G08 | Legionella sp. | std_9 | 120.44 | 2.0808 | stds3_9 | rep 2 | FAM | 33.01
G04 | Legionella sp. | std_7 | 12044 | 4.0808 | stds1_7 | rep 3 | FAM | 29.38 | H08 | Legionella sp. | NTC | 0 | #NUM! | stds3_9 | rep 2 | FAM | 37.37
G05 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 1 | CY5 | 30.52 | A09 | Legionella sp. | std_3 | 120440000 | 8.0808 | stds3_9 | rep 3 | FAM | 17.09
G05 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 1 | FAM | 33.37 | B09 | Legionella sp. | std_4 | 12044000 | 7.0808 | stds3_9 | rep 3 | FAM | 19.43
G06 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 2 | CY5 | 30.89 | C09 | Legionella sp. | std_5 | 1204400 | 6.0808 | stds3_9 | rep 3 | FAM | 21.93
G06 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 2 | FAM | 33.91 | D09 | Legionella sp. | std_6 | 120440 | 5.0808 | stds3_9 | rep 3 | FAM | 24.11
G07 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 3 | CY5 | 30.15 | E09 | Legionella sp. | std_7 | 12044 | 4.0808 | stds3_9 | rep 3 | FAM | 28.99
G07 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 3 | FAM | 34.5 | F09 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds3_9 | rep 3 | FAM | 31.76
G08 | Internal Control | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 4 | CY5 | 30.03 | G09 | Legionella sp. | std_9 | 120.44 | 2.0808 | stds3_9 | rep 3 | FAM | 33.88
G08 | Legionella sp. | std_8 | 1204.4 | 3.0808 | stds2_8 | rep 4 | FAM | 34.64 | H09 | Legionella sp. | NTC | 0 | #NUM! | stds3_9 | rep 3 | FAM
G09 | Internal Control | std_9 | 120.44 | 2.0808 | stds3_9 | rep 1 | CY5 | 31.36 | A10
G09 | Legionella sp. | std_9 | 120.44 | 2.0808 | stds3_9 | rep 1 | FAM | 32.69 | A10
G10 | Internal Control | std_9 | 120.44 | 2.0808 | stds3_9 | rep 2 | CY5 | 32.07 | B10
G10 | Legionella sp. | std_9 | 120.44 | 2.0808 | stds3_9 | rep 2 | FAM | 33.01 | B10
G11 | Internal Control | std_9 | 120.44 | 2.0808 | stds3_9 | rep 3 | CY5 | 32.8 | C10
G11 | Legionella sp. | std_9 | 120.44 | 2.0808 | stds3_9 | rep 3 | FAM | 33.88 | C10
G12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | D10
G12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | D10
H01 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | E10
H01 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | E10
H02 | Internal Control | NTC | 0 | #NUM! | stds1_7 | rep 1 | CY5 | 30.49 | F10
H02 | Legionella sp. | NTC | 0 | #NUM! | stds1_7 | rep 1 | FAM | 37.88 | F10
H03 | Internal Control | NTC | 0 | #NUM! | stds1_7 | rep 2 | CY5 | 30.29 | G10
H03 | Legionella sp. | NTC | 0 | #NUM! | stds1_7 | rep 2 | FAM |  | G10
H04 | Internal Control | NTC | 0 | #NUM! | stds1_7 | rep 3 | CY5 | 30.9 | H10
H04 | Legionella sp. | NTC | 0 | #NUM! | stds1_7 | rep 3 | FAM | 37.01 | H10
H05 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 1 | CY5 | 30.55 | A11
H05 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 1 | FAM |  | A11
H06 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 2 | CY5 | 30.13 | B11
H06 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 2 | FAM |  | B11
H07 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 3 | CY5 | 30.53 | C11
H07 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 3 | FAM |  | C11
H08 | Internal Control | NTC | 0 | #NUM! | stds2_8 | rep 4 | CY5 | 30.87 | D11
H08 | Legionella sp. | NTC | 0 | #NUM! | stds2_8 | rep 4 | FAM |  | D11
H09 | Internal Control | NTC | 0 | #NUM! | stds3_9 | rep 1 | CY5 | 30.14 | E11
H09 | Legionella sp. | NTC | 0 | #NUM! | stds3_9 | rep 1 | FAM | 39.78 | E11
H10 | Internal Control | NTC | 0 | #NUM! | stds3_9 | rep 2 | CY5 | 30.19 | F11
H10 | Legionella sp. | NTC | 0 | #NUM! | stds3_9 | rep 2 | FAM | 37.37 | F11
H11 | Internal Control | NTC | 0 | #NUM! | stds3_9 | rep 3 | CY5 | 29.94 | G11
H11 | Legionella sp. | NTC | 0 | #NUM! | stds3_9 | rep 3 | FAM |  | G11
H12 | Internal Control | empty | empty | #VALUE! | empty | empty | CY5 |  | H11
H12 | Legionella sp. | empty | empty | #VALUE! | empty | empty | FAM |  | H11
Note on AT102: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: These two values should agree, but they don't. Ok, thy're close enough. Take avg

## variability_Q16
Dye | CY5
Max of Ct | Column Labels
Row Labels | gDNA (1e5) | gDNA (1e4) | gDNA (1e3) | gDNA (1e2) | gDNA (1e1) | gDNA (1e0) | Grand Total
Leg spp | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.59 | 39.82
miniS | 35.8 | 37.34 | 36.55 | 35.7 | 36.78 | 37.59 | 37.59
Q16 |  | 34.91 | 34.59 |  |  |  | 34.91
Q96 | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.04 | 39.82
Grand Total | 39.82 | 37.5 | 36.7 | 38.42 | 38.23 | 37.59 | 39.82

## 1st LOD
Well | Target | Property | Dye | Ct | Ct Mean | Ct SD | Concentration | Aver. Con. | Con. SD | ThermalCycler | Well | Target | Property | replicate | Dye | Ct | Copy#/RXN | ThermalCycler | Well | Target | Property | replicate | Dye | Ct | Copy#/RXN
A03 | Internal Control | Unknown | CY5 | 35.93 | [5 empty] | Q96 | A03 | Internal Control | lyophilized | rep 1 | CY5 | 35.93 | 10^3 | Q96 | A03 | Legionella sp. | lyophilized | rep 1 | FAM | 27.23 | 10^3
A03 | Legionella sp. | Unknown | FAM | 27.23 | [5 empty] | Q96 | A03 | Legionella sp. | lyophilized | rep 1 | FAM | 27.23 | 10^3 | Q96 | A04 | Legionella sp. | lyophilized | rep 1 | FAM | 29.75 | 10^2 | TubeName | Conc (uM) | ng | pg/rxn | Copy Number (from ng) | Copy Number (from conc) | Copy Number/rxn (2ul)
A04 | Internal Control | Unknown | CY5 | 33.22 | [5 empty] | Q96 | A04 | Internal Control | lyophilized | rep 1 | CY5 | 33.22 | 10^2 | Q96 | A05 | Legionella sp. | lyophilized | rep 1 | FAM | 31.93 | 10^1 | IDT | 92.35 | 122000 |  |  | 2.2245e+15 |  |  |  | Estimating LOD Lyophilized (Q96)
A04 | Legionella sp. | Unknown | FAM | 29.75 | [5 empty] | Q96 | A04 | Legionella sp. | lyophilized | rep 1 | FAM | 29.75 | 10^2 | Q96 | A06 | Legionella sp. | lyophilized | rep 1 | FAM | 38.03 | 10^0
A05 | Internal Control | Unknown | CY5 | 31.96 | [5 empty] | Q96 | A05 | Internal Control | lyophilized | rep 1 | CY5 | 31.96 | 10^1 | Q96 | A07 | Legionella sp. | lyophilized | rep 1 | FAM |  | 10^-1 | Tube_0 | 0.9235 | 30500 |  |  | 5.5613e+14 | [5 empty] | Copy#/RXN | TOTAL | SUCCESS | FRACTION POS
A05 | Legionella sp. | Unknown | FAM | 31.93 | [5 empty] | Q96 | A05 | Legionella sp. | lyophilized | rep 1 | FAM | 31.93 | 10^1 | Q96 | A08 | Legionella sp. | lyophilized | rep 1 | FAM |  | 10^-2 | [11 empty] | 10^3 | 4 | 4 | 1 | 100
A06 | Internal Control | Unknown | CY5 | 32.9 | [5 empty] | Q96 | A06 | Internal Control | lyophilized | rep 1 | CY5 | 32.9 | 10^0 | Q96 | A09 | Legionella sp. | lyophilized | rep 1 | FAM |  | 10^-3 | Tube_1 | 0.009235 | 305 | 0.61 | 6.2742e+12 | 5.5613e+12 | 11122634000 | 5.5613e+12 | 11122634000 |  |  | 10^2 | 4 | 4 | 1 | 100
A06 | Legionella sp. | Unknown | FAM | 38.03 | [5 empty] | Q96 | A06 | Legionella sp. | lyophilized | rep 1 | FAM | 38.03 | 10^0 | Q96 | A10 | Legionella sp. | lyophilized | rep 1 | FAM |  | NTC | Tube_2 | 0.0009235 | 30.5 | 0.061 | 627418093126.6887 | 556131700000.0001 | 1112263400 | 556131700000 | 1112263400 |  |  | 10^1 | 4 | 4 | 1 | 100
A07 | Internal Control | Unknown | CY5 | 31.97 | [5 empty] | Q96 | A07 | Internal Control | lyophilized | rep 1 | CY5 | 31.97 | 10^-1 | Q96 | C03 | Legionella sp. | lyophilized | rep 2 | FAM | 26.83 | 10^3 | Tube_3 | 9.235e-05 | 3.05 | 0.0061 | 62741809312.6689 | 55613170000 | 111226340 | 55613170000 | 111226340 |  |  | 10^0 | 4 | 2 | 0.5 | 50
A07 | Legionella sp. | Unknown | FAM | 36.94 | [5 empty] | Q96 | A07 | Legionella sp. | lyophilized | rep 1 | FAM | 36.94 | 10^-1 | Q96 | C04 | Legionella sp. | lyophilized | rep 2 | FAM | 29.71 | 10^2 | Tube_4 | 9.235e-06 | 0.305 | 0.00061 | 6274180931.2669 | 5561317000 | 11122634 | 5561317000 | 11122634 |  |  | 10^-1 | 4 | 1 | 0.25 | 25
A08 | Internal Control | Unknown | CY5 | 32.54 | [5 empty] | Q96 | A08 | Internal Control | lyophilized | rep 1 | CY5 | 32.54 | 10^-2 | Q96 | C05 | Legionella sp. | lyophilized | rep 2 | FAM | 32.84 | 10^1 | Tube_5 | 9.235e-07 | 0.0305 | 6.1e-05 | 627418093.1267 | 556131700 | 1112263.4 | 556131700 | 1112263.4 |  |  | 10^-2 | 4 | 0 | 0 | 0
A08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | A08 | Legionella sp. | lyophilized | rep 1 | FAM |  | 10^-2 | Q96 | C06 | Legionella sp. | lyophilized | rep 2 | FAM | 36.86 | 10^0 | Tube_6 | 9.235e-08 | 0.00305 | 6.1e-06 | 62741809.3127 | 55613170 | 111226.34 | 55613170 | 111226.34 |  |  | 10^-3 | 4 | 0 | 0 | 0
A09 | Internal Control | Unknown | CY5 | 32.78 | [5 empty] | Q96 | A09 | Internal Control | lyophilized | rep 1 | CY5 | 32.78 | 10^-3 | Q96 | C07 | Legionella sp. | lyophilized | rep 2 | FAM | 38.81 | 10^-1 | Tube_7 | 9.235e-09 | 0.000305 | 6.1e-07 | 6274180.9313 | 5561317 | 11122.634 | 5561317 | 11122.634 |  |  | NTC | 4 | 0 | 0 | 0
A09 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | A09 | Legionella sp. | lyophilized | rep 1 | FAM |  | 10^-3 | Q96 | C08 | Legionella sp. | lyophilized | rep 2 | FAM |  | 10^-2 | Tube_8 | 9.235e-10 | 3.05e-05 | 6.1e-08 | 627418.0931 | 556131.7 | 1112.2634 | 556131.7 | 1112.2634
A10 | Internal Control | Unknown | CY5 | 32.88 | [5 empty] | Q96 | A10 | Internal Control | lyophilized | rep 1 | CY5 | 32.88 | NTC | Q96 | C09 | Legionella sp. | lyophilized | rep 2 | FAM |  | 10^-3 | Tube_9 | 9.235e-11 | 3.05e-06 | 6.1e-09 | 62741.8093 | 55613.17 | 111.2263 | 55613.17 | 111.2263
A10 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | A10 | Legionella sp. | lyophilized | rep 1 | FAM |  | NTC | Q96 | C10 | Legionella sp. | lyophilized | rep 2 | FAM | 37.27 | NTC | Tube_10 | 9.235e-12 | 3.05e-07 | 6.1e-10 | 6274.1809 | 5561.317 | 11.1226 | 5561.317 | 11.1226
B03 | Internal Control | Unknown | CY5 | 41.78 | [5 empty] | Q96 | B03 | Internal Control | beads | rep 1 | CY5 | 41.78 | 10^3 | Q96 | E03 | Legionella sp. | lyophilized | rep 3 | FAM | 27.24 | 10^3 | Tube_11 | 9.235e-13 | 3.05e-08 | 6.1e-11 | 627.4181 | 556.1317 | 1.1123 | 556.1317 | 1.1123
B03 | Legionella sp. | Unknown | FAM | 31.65 | [5 empty] | Q96 | B03 | Legionella sp. | beads | rep 1 | FAM | 31.65 | 10^3 | Q96 | E04 | Legionella sp. | lyophilized | rep 3 | FAM | 31.04 | 10^2 | Tube_12 | 9.235e-14 | 3.05e-09 | 6.1e-12 | 62.7418 | 55.6132 | 0.11123 | 55.6132 | 0.11123 |  |  |  | Conc (uM) | Log(Conc) | Fraction Pos | EQN
B04 | Internal Control | Unknown | CY5 | 34.07 | [5 empty] | Q96 | B04 | Internal Control | beads | rep 1 | CY5 | 34.07 | 10^2 | Q96 | E05 | Legionella sp. | lyophilized | rep 3 | FAM | 32.83 | 10^1 | Tube_13 | 9.235e-15 | 3.05e-10 |  | 6.2742 | 5.5613 | 0.011123 | [5 empty] | 9.235e-10 | -9.0346 | 100 | 101.2638
B04 | Legionella sp. | Unknown | FAM | 35.7 | [5 empty] | Q96 | B04 | Legionella sp. | beads | rep 1 | FAM | 35.7 | 10^2 | Q96 | E06 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^0 | Tube_14 | 9.235e-16 | 3.05e-11 |  | 0.62742 | 0.55613 | 0.0011123 | [5 empty] | 9.235e-11 | -10.0346 | 100 | 101.2635
B05 | Internal Control | Unknown | CY5 | 34.26 | [5 empty] | Q96 | B05 | Internal Control | beads | rep 1 | CY5 | 34.26 | 10^1 | Q96 | E07 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^-1 | Tube_15 | 9.235e-17 | 3.05e-12 |  | 0.062742 | 0.055613 | 0.00011123 |  |  | 1st method |  |  | 9.235e-12 | -11.0346 | 100 | 96.0249
B05 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | B05 | Legionella sp. | beads | rep 1 | FAM |  | 10^1 | Q96 | E08 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^-2 | [9 empty] | this one prob more accurate |  |  | 9.235e-13 | -12.0346 | 50 | 54.1619
B06 | Internal Control | Unknown | CY5 | 34.68 | [5 empty] | Q96 | B06 | Internal Control | beads | rep 1 | CY5 | 34.68 | 10^0 | Q96 | E09 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^-3 | [12 empty] | 9.235e-14 | -13.0346 | 25 | 18.6373
B06 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | B06 | Legionella sp. | beads | rep 1 | FAM |  | 10^0 | Q96 | E10 | Legionella sp. | lyophilized | rep 3 | FAM |  | NTC | [12 empty] | 9.235e-15 | -14.0346 | 0 | 4.3032
B07 | Internal Control | Unknown | CY5 | 33.43 | [5 empty] | Q96 | B07 | Internal Control | beads | rep 1 | CY5 | 33.43 | 10^-1 | Q96 | G03 | Legionella sp. | lyophilized | rep 4 | FAM | 27.27 | 10^3 | [12 empty] | 9.235e-16 | -15.0346 | 0 | -0.65486
B07 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | B07 | Legionella sp. | beads | rep 1 | FAM |  | 10^-1 | Q96 | G04 | Legionella sp. | lyophilized | rep 4 | FAM | 29.94 | 10^2 | [10 empty] | 100.3607 | 5018.0346 | 8.3328e-12 | -11.0792 |  | 95
B08 | Internal Control | Unknown | CY5 | 32.94 | [5 empty] | Q96 | B08 | Internal Control | beads | rep 1 | CY5 | 32.94 | 10^-2 | Q96 | G05 | Legionella sp. | lyophilized | rep 4 | FAM | 33.03 | 10^1
B08 | Legionella sp. | Unknown | FAM | 38.35 | [5 empty] | Q96 | B08 | Legionella sp. | beads | rep 1 | FAM | 38.35 | 10^-2 | Q96 | G06 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^0 | [15 empty] | y = 101.2638 + (-3.43411 - 101.2638)/(1 + (x/-3.869408)^-15.23477)^25697910
B09 | Internal Control | Unknown | CY5 | 34.31 | [5 empty] | Q96 | B09 | Internal Control | beads | rep 1 | CY5 | 34.31 | 10^-3 | Q96 | G07 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^-1 | [12 empty] | Conc (uM) | Log(Conc) | Fraction Pos | EQN
B09 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | B09 | Legionella sp. | beads | rep 1 | FAM |  | 10^-3 | Q96 | G08 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^-2 | [13 empty] | -9.0346 | 100 | #NUM!
B10 | Internal Control | Unknown | CY5 | 34.35 | [5 empty] | Q96 | B10 | Internal Control | beads | rep 1 | CY5 | 34.35 | NTC | Q96 | G09 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^-3 | [13 empty] | -10.0346 | 100 | 100
B10 | Legionella sp. | Unknown | FAM | 39.55 | [5 empty] | Q96 | B10 | Legionella sp. | beads | rep 1 | FAM | 39.55 | NTC | Q96 | G10 | Legionella sp. | lyophilized | rep 4 | FAM |  | NTC | [9 empty] | compared above to see if losing datapoint was sig. |  |  |  | -11.0346 | 100 | 100
C03 | Internal Control | Unknown | CY5 | 36.58 | [5 empty] | Q96 | C03 | Internal Control | lyophilized | rep 2 | CY5 | 36.58 | 10^3 | Q16 | A1-1 | Legionella sp. | lyophilized | rep 5 | FAM | 27.34 | 10^3 | [9 empty] | 20210623 |  |  |  | -12.0346 | 50 | 49.9995
C03 | Legionella sp. | Unknown | FAM | 26.83 | [5 empty] | Q96 | C03 | Legionella sp. | lyophilized | rep 2 | FAM | 26.83 | 10^3 | Q16 | A2-1 | Legionella sp. | lyophilized | rep 5 | FAM | 31.98 | 10^2 | [13 empty] | -13.0346 | 0 | 6.9275e-07
C04 | Internal Control | Unknown | CY5 | 32.11 | [5 empty] | Q96 | C04 | Internal Control | lyophilized | rep 2 | CY5 | 32.11 | 10^2 | Q16 | A3-1 | Legionella sp. | lyophilized | rep 5 | FAM | 32.81 | 10^1 | [13 empty] | -14.0346 | 0 | -3.4638e-07
C04 | Legionella sp. | Unknown | FAM | 29.71 | [5 empty] | Q96 | C04 | Legionella sp. | lyophilized | rep 2 | FAM | 29.71 | 10^2 | Q16 | A4-1 | Legionella sp. | lyophilized | rep 5 | FAM | 37.65 | 10^0 | [13 empty] | -15.0346 | 0
C05 | Internal Control | Unknown | CY5 | 32.26 | [5 empty] | Q96 | C05 | Internal Control | lyophilized | rep 2 | CY5 | 32.26 | 10^1 | Q16 | A5-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | 10^-1 | [10 empty] | 13.4363 | 671.8162 | 1.1156e-12 | -11.9525 |  | 95
C05 | Legionella sp. | Unknown | FAM | 32.84 | [5 empty] | Q96 | C05 | Legionella sp. | lyophilized | rep 2 | FAM | 32.84 | 10^1 | Q16 | A6-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | 10^-2
C06 | Internal Control | Unknown | CY5 | 32.05 | [5 empty] | Q96 | C06 | Internal Control | lyophilized | rep 2 | CY5 | 32.05 | 10^0 | Q16 | A7-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | 10^-3 | [15 empty] | y = 100 + (-3.463824e-7 - 100)/(1 + (x/-11.87615)^-226.0071)^14.19253
C06 | Legionella sp. | Unknown | FAM | 36.86 | [5 empty] | Q96 | C06 | Legionella sp. | lyophilized | rep 2 | FAM | 36.86 | 10^0 | Q16 | A8-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | NTC
C07 | Internal Control | Unknown | CY5 | 31.99 | [5 empty] | Q96 | C07 | Internal Control | lyophilized | rep 2 | CY5 | 31.99 | 10^-1 | Q96 | B03 | Legionella sp. | beads | rep 1 | FAM | 31.65 | 10^3
C07 | Legionella sp. | Unknown | FAM | 38.81 | [5 empty] | Q96 | C07 | Legionella sp. | lyophilized | rep 2 | FAM | 38.81 | 10^-1 | Q96 | B04 | Legionella sp. | beads | rep 1 | FAM | 35.7 | 10^2
C08 | Internal Control | Unknown | CY5 | 32.16 | [5 empty] | Q96 | C08 | Internal Control | lyophilized | rep 2 | CY5 | 32.16 | 10^-2 | Q96 | B05 | Legionella sp. | beads | rep 1 | FAM |  | 10^1 | [9 empty] | Estimating LOD Beads
C08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | C08 | Legionella sp. | lyophilized | rep 2 | FAM |  | 10^-2 | Q96 | B06 | Legionella sp. | beads | rep 1 | FAM |  | 10^0 | [11 empty] | Copy#/RXN | TOTAL | SUCCESS | FRACTION POS
C09 | Internal Control | Unknown | CY5 | 32.13 | [5 empty] | Q96 | C09 | Internal Control | lyophilized | rep 2 | CY5 | 32.13 | 10^-3 | Q96 | B07 | Legionella sp. | beads | rep 1 | FAM |  | 10^-1 | [11 empty] | 10^3 | 4 | 4 | 1 | 100
C09 | Legionella sp. | Unknown | FAM | 39.46 | [5 empty] | Q96 | C09 | Legionella sp. | lyophilized | rep 2 | FAM | 39.46 | 10^-3 | Q96 | B08 | Legionella sp. | beads | rep 1 | FAM |  | 10^-2 | [11 empty] | 10^2 | 4 | 4 | 1 | 100
C10 | Internal Control | Unknown | CY5 | 33.25 | [5 empty] | Q96 | C10 | Internal Control | lyophilized | rep 2 | CY5 | 33.25 | NTC | Q96 | B09 | Legionella sp. | beads | rep 1 | FAM |  | 10^-3 | [11 empty] | 10^1 | 4 | 2 | 0.5 | 50
C10 | Legionella sp. | Unknown | FAM | 37.27 | [5 empty] | Q96 | C10 | Legionella sp. | lyophilized | rep 2 | FAM | 37.27 | NTC | Q96 | B10 | Legionella sp. | beads | rep 1 | FAM | 39.55 | NTC | [11 empty] | 10^0 | 4 | 1 | 0.25 | 25
D03 | Internal Control | Unknown | CY5 | 42.7 | [5 empty] | Q96 | D03 | Internal Control | beads | rep 2 | CY5 | 42.7 | 10^3 | Q96 | D03 | Legionella sp. | beads | rep 2 | FAM | 29.44 | 10^3 | [11 empty] | 10^-1 | 4 | 0 | 0 | 0
D03 | Legionella sp. | Unknown | FAM | 29.44 | [5 empty] | Q96 | D03 | Legionella sp. | beads | rep 2 | FAM | 29.44 | 10^3 | Q96 | D04 | Legionella sp. | beads | rep 2 | FAM | 33.08 | 10^2 | [11 empty] | 10^-2 | 4 | 0 | 0 | 0
D04 | Internal Control | Unknown | CY5 | 34 | [5 empty] | Q96 | D04 | Internal Control | beads | rep 2 | CY5 | 34 | 10^2 | Q96 | D05 | Legionella sp. | beads | rep 2 | FAM | 35.46 | 10^1 | [11 empty] | 10^-3 | 4 | 0 | 0 | 0
D04 | Legionella sp. | Unknown | FAM | 33.08 | [5 empty] | Q96 | D04 | Legionella sp. | beads | rep 2 | FAM | 33.08 | 10^2 | Q96 | D06 | Legionella sp. | beads | rep 2 | FAM | 44.72 | 10^0 | [11 empty] | NTC | 4 | 0 | 0 | 0
D05 | Internal Control | Unknown | CY5 | 33.83 | [5 empty] | Q96 | D05 | Internal Control | beads | rep 2 | CY5 | 33.83 | 10^1 | Q96 | D07 | Legionella sp. | beads | rep 2 | FAM |  | 10^-1
D05 | Legionella sp. | Unknown | FAM | 35.46 | [5 empty] | Q96 | D05 | Legionella sp. | beads | rep 2 | FAM | 35.46 | 10^1 | Q96 | D08 | Legionella sp. | beads | rep 2 | FAM |  | 10^-2
D06 | Internal Control | Unknown | CY5 | 33.57 | [5 empty] | Q96 | D06 | Internal Control | beads | rep 2 | CY5 | 33.57 | 10^0 | Q96 | D09 | Legionella sp. | beads | rep 2 | FAM |  | 10^-3 | [13 empty] | Log(Conc) | Positive | EQN
D06 | Legionella sp. | Unknown | FAM | 44.72 | [5 empty] | Q96 | D06 | Legionella sp. | beads | rep 2 | FAM | 44.72 | 10^0 | Q96 | D10 | Legionella sp. | beads | rep 2 | FAM |  | NTC | [13 empty] | -9.0346 | 100 | 102.053
D07 | Internal Control | Unknown | CY5 | 33.53 | [5 empty] | Q96 | D07 | Internal Control | beads | rep 2 | CY5 | 33.53 | 10^-1 | Q96 | F03 | Legionella sp. | beads | rep 3 | FAM | 31.81 | 10^3 | [13 empty] | -10.0346 | 100 | 96.6497
D07 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | D07 | Legionella sp. | beads | rep 2 | FAM |  | 10^-1 | Q96 | F04 | Legionella sp. | beads | rep 3 | FAM | 32.44 | 10^2 | [13 empty] | -11.0346 | 50 | 53.4655
D08 | Internal Control | Unknown | CY5 | 33.7 | [5 empty] | Q96 | D08 | Internal Control | beads | rep 2 | CY5 | 33.7 | 10^-2 | Q96 | F05 | Legionella sp. | beads | rep 3 | FAM | 34.85 | 10^1 | [13 empty] | -12.0346 | 25 | 18.4863
D08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | D08 | Legionella sp. | beads | rep 2 | FAM |  | 10^-2 | Q96 | F06 | Legionella sp. | beads | rep 3 | FAM |  | 10^0 | [13 empty] | -13.0346 | 0 | 4.8902
D09 | Internal Control | Unknown | CY5 | 34.06 | [5 empty] | Q96 | D09 | Internal Control | beads | rep 2 | CY5 | 34.06 | 10^-3 | Q96 | F07 | Legionella sp. | beads | rep 3 | FAM |  | 10^-1 | [13 empty] | -14.0346 | 0 | 0.28526
D09 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | D09 | Legionella sp. | beads | rep 2 | FAM |  | 10^-3 | Q96 | F08 | Legionella sp. | beads | rep 3 | FAM |  | 10^-2 | [13 empty] | -15.0346 | 0 | -1.3035
D10 | Internal Control | Unknown | CY5 | 33.24 | [5 empty] | Q96 | D10 | Internal Control | beads | rep 2 | CY5 | 33.24 | NTC | Q96 | F09 | Legionella sp. | beads | rep 3 | FAM |  | 10^-3 | [10 empty] | 954.0593 | 47702.9629 | 7.9214e-11 | -10.1012 |  | 94.999
D10 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | D10 | Legionella sp. | beads | rep 2 | FAM |  | NTC | Q96 | F10 | Legionella sp. | beads | rep 3 | FAM |  | NTC
E03 | Internal Control | Unknown | CY5 | 33.77 | [5 empty] | Q96 | E03 | Internal Control | lyophilized | rep 3 | CY5 | 33.77 | 10^3 | Q96 | H03 | Legionella sp. | beads | rep 4 | FAM | 27.53 | 10^3 | [15 empty] | y = 102.0532 + (-2.267969 - 102.0532)/(1 + (x/-3.26389)^-14.25711)^26634450
E03 | Legionella sp. | Unknown | FAM | 27.24 | [5 empty] | Q96 | E03 | Legionella sp. | lyophilized | rep 3 | FAM | 27.24 | 10^3 | Q96 | H04 | Legionella sp. | beads | rep 4 | FAM | 35.93 | 10^2
E04 | Internal Control | Unknown | CY5 | 32.87 | [5 empty] | Q96 | E04 | Internal Control | lyophilized | rep 3 | CY5 | 32.87 | 10^2 | Q96 | H05 | Legionella sp. | beads | rep 4 | FAM |  | 10^1 | [9 empty] | 2nd Estimating LOD Lyophilized (Q96)
E04 | Legionella sp. | Unknown | FAM | 31.04 | [5 empty] | Q96 | E04 | Legionella sp. | lyophilized | rep 3 | FAM | 31.04 | 10^2 | Q96 | H06 | Legionella sp. | beads | rep 4 | FAM |  | 10^0 | [9 empty] | estimating using GraphPad.
E05 | Internal Control | Unknown | CY5 | 32.85 | [5 empty] | Q96 | E05 | Internal Control | lyophilized | rep 3 | CY5 | 32.85 | 10^1 | Q96 | H07 | Legionella sp. | beads | rep 4 | FAM |  | 10^-1 | [9 empty] | 5PL, logarithmic
E05 | Legionella sp. | Unknown | FAM | 32.83 | [5 empty] | Q96 | E05 | Legionella sp. | lyophilized | rep 3 | FAM | 32.83 | 10^1 | Q96 | H08 | Legionella sp. | beads | rep 4 | FAM |  | 10^-2
E06 | Internal Control | Unknown | CY5 | 32.69 | [5 empty] | Q96 | E06 | Internal Control | lyophilized | rep 3 | CY5 | 32.69 | 10^0 | Q96 | H09 | Legionella sp. | beads | rep 4 | FAM |  | 10^-3
E06 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | E06 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^0 | Q96 | H10 | Legionella sp. | beads | rep 4 | FAM |  | NTC
E07 | Internal Control | Unknown | CY5 | 32.23 | [5 empty] | Q96 | E07 | Internal Control | lyophilized | rep 3 | CY5 | 32.23 | 10^-1 | Q16 | B1-1 | Legionella sp. | beads | rep 5 | FAM | 28.15 | 10^3
E07 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | E07 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^-1 | Q16 | B2-1 | Legionella sp. | beads | rep 5 | FAM | 30.74 | 10^2
E08 | Internal Control | Unknown | CY5 | 32.66 | [5 empty] | Q96 | E08 | Internal Control | lyophilized | rep 3 | CY5 | 32.66 | 10^-2 | Q16 | B3-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^1
E08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | E08 | Legionella sp. | lyophilized | rep 3 | FAM |  | 10^-2 | Q16 | B4-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^0
E09 | Internal Control | Unknown | CY5 | 32.39 | [5 empty] | Q96 | E09 | Internal Control | lyophilized | rep 3 | CY5 | 32.39 | 10^-3 | Q16 | B5-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^-1
E09 | Legionella sp. | Unknown | FAM | 36.78 | [5 empty] | Q96 | E09 | Legionella sp. | lyophilized | rep 3 | FAM | 36.78 | 10^-3 | Q16 | B6-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^-2
E10 | Internal Control | Unknown | CY5 | 31.83 | [5 empty] | Q96 | E10 | Internal Control | lyophilized | rep 3 | CY5 | 31.83 | NTC | Q16 | B7-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^-3
E10 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | E10 | Legionella sp. | lyophilized | rep 3 | FAM |  | NTC | Q16 | B8-1 | Legionella sp. | beads | rep 5 | FAM |  | NTC
F03 | Internal Control | Unknown | CY5 | 44.31 | [5 empty] | Q96 | F03 | Internal Control | beads | rep 3 | CY5 | 44.31 | 10^3 | Q96 | A03 | Internal Control | lyophilized | rep 1 | CY5 | 35.93 | 10^3
F03 | Legionella sp. | Unknown | FAM | 31.81 | [5 empty] | Q96 | F03 | Legionella sp. | beads | rep 3 | FAM | 31.81 | 10^3 | Q96 | A04 | Internal Control | lyophilized | rep 1 | CY5 | 33.22 | 10^2
F04 | Internal Control | Unknown | CY5 | 34.96 | [5 empty] | Q96 | F04 | Internal Control | beads | rep 3 | CY5 | 34.96 | 10^2 | Q96 | A05 | Internal Control | lyophilized | rep 1 | CY5 | 31.96 | 10^1
F04 | Legionella sp. | Unknown | FAM | 32.44 | [5 empty] | Q96 | F04 | Legionella sp. | beads | rep 3 | FAM | 32.44 | 10^2 | Q96 | A06 | Internal Control | lyophilized | rep 1 | CY5 | 32.9 | 10^0
F05 | Internal Control | Unknown | CY5 | 33.32 | [5 empty] | Q96 | F05 | Internal Control | beads | rep 3 | CY5 | 33.32 | 10^1 | Q96 | A07 | Internal Control | lyophilized | rep 1 | CY5 | 31.97 | 10^-1
F05 | Legionella sp. | Unknown | FAM | 34.85 | [5 empty] | Q96 | F05 | Legionella sp. | beads | rep 3 | FAM | 34.85 | 10^1 | Q96 | A08 | Internal Control | lyophilized | rep 1 | CY5 | 32.54 | 10^-2
F06 | Internal Control | Unknown | CY5 | 33.86 | [5 empty] | Q96 | F06 | Internal Control | beads | rep 3 | CY5 | 33.86 | 10^0 | Q96 | A09 | Internal Control | lyophilized | rep 1 | CY5 | 32.78 | 10^-3
F06 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | F06 | Legionella sp. | beads | rep 3 | FAM |  | 10^0 | Q96 | A10 | Internal Control | lyophilized | rep 1 | CY5 | 32.88 | NTC
F07 | Internal Control | Unknown | CY5 | 34.19 | [5 empty] | Q96 | F07 | Internal Control | beads | rep 3 | CY5 | 34.19 | 10^-1 | Q96 | C03 | Internal Control | lyophilized | rep 2 | CY5 | 36.58 | 10^3 | [11 empty] | -11.1862 | 6.5139e-12 | 3922.696 | 78.4539
F07 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | F07 | Legionella sp. | beads | rep 3 | FAM |  | 10^-1 | Q96 | C04 | Internal Control | lyophilized | rep 2 | CY5 | 32.11 | 10^2
F08 | Internal Control | Unknown | CY5 | 33.33 | [5 empty] | Q96 | F08 | Internal Control | beads | rep 3 | CY5 | 33.33 | 10^-2 | Q96 | C05 | Internal Control | lyophilized | rep 2 | CY5 | 32.26 | 10^1
F08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | F08 | Legionella sp. | beads | rep 3 | FAM |  | 10^-2 | Q96 | C06 | Internal Control | lyophilized | rep 2 | CY5 | 32.05 | 10^0
F09 | Internal Control | Unknown | CY5 | 34.73 | [5 empty] | Q96 | F09 | Internal Control | beads | rep 3 | CY5 | 34.73 | 10^-3 | Q96 | C07 | Internal Control | lyophilized | rep 2 | CY5 | 31.99 | 10^-1
F09 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | F09 | Legionella sp. | beads | rep 3 | FAM |  | 10^-3 | Q96 | C08 | Internal Control | lyophilized | rep 2 | CY5 | 32.16 | 10^-2
F10 | Internal Control | Unknown | CY5 | 32.77 | [5 empty] | Q96 | F10 | Internal Control | beads | rep 3 | CY5 | 32.77 | NTC | Q96 | C09 | Internal Control | lyophilized | rep 2 | CY5 | 32.13 | 10^-3 | [9 empty] | 3rd Estimating LOD Lyophilized (Q96)
F10 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | F10 | Legionella sp. | beads | rep 3 | FAM |  | NTC | Q96 | C10 | Internal Control | lyophilized | rep 2 | CY5 | 33.25 | NTC | [9 empty] | estimating using GraphPad.
G03 | Internal Control | Unknown | CY5 | 35.29 | [5 empty] | Q96 | G03 | Internal Control | lyophilized | rep 4 | CY5 | 35.29 | 10^3 | Q96 | E03 | Internal Control | lyophilized | rep 3 | CY5 | 33.77 | 10^3 | [9 empty] | 4PL, logarithmic
G03 | Legionella sp. | Unknown | FAM | 27.27 | [5 empty] | Q96 | G03 | Legionella sp. | lyophilized | rep 4 | FAM | 27.27 | 10^3 | Q96 | E04 | Internal Control | lyophilized | rep 3 | CY5 | 32.87 | 10^2
G04 | Internal Control | Unknown | CY5 | 31.99 | [5 empty] | Q96 | G04 | Internal Control | lyophilized | rep 4 | CY5 | 31.99 | 10^2 | Q96 | E05 | Internal Control | lyophilized | rep 3 | CY5 | 32.85 | 10^1 | [11 empty] | -10.8078 | 1.5568e-11 | 9374.9887 | 187.4998
G04 | Legionella sp. | Unknown | FAM | 29.94 | [5 empty] | Q96 | G04 | Legionella sp. | lyophilized | rep 4 | FAM | 29.94 | 10^2 | Q96 | E06 | Internal Control | lyophilized | rep 3 | CY5 | 32.69 | 10^0
G05 | Internal Control | Unknown | CY5 | 32.96 | [5 empty] | Q96 | G05 | Internal Control | lyophilized | rep 4 | CY5 | 32.96 | 10^1 | Q96 | E07 | Internal Control | lyophilized | rep 3 | CY5 | 32.23 | 10^-1 | [11 empty] | +error | 17.328
G05 | Legionella sp. | Unknown | FAM | 33.03 | [5 empty] | Q96 | G05 | Legionella sp. | lyophilized | rep 4 | FAM | 33.03 | 10^1 | Q96 | E08 | Internal Control | lyophilized | rep 3 | CY5 | 32.66 | 10^-2 | [11 empty] | -error | 17.328
G06 | Internal Control | Unknown | CY5 | 31.8 | [5 empty] | Q96 | G06 | Internal Control | lyophilized | rep 4 | CY5 | 31.8 | 10^0 | Q96 | E09 | Internal Control | lyophilized | rep 3 | CY5 | 32.39 | 10^-3
G06 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | G06 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^0 | Q96 | E10 | Internal Control | lyophilized | rep 3 | CY5 | 31.83 | NTC | [11 empty] | -11.5345 | 2.9208e-12 | 1758.8984 | 35.178
G07 | Internal Control | Unknown | CY5 | 31.85 | [5 empty] | Q96 | G07 | Internal Control | lyophilized | rep 4 | CY5 | 31.85 | 10^-1 | Q96 | G03 | Internal Control | lyophilized | rep 4 | CY5 | 35.29 | 10^3 | [11 empty] | -10.3093 | 4.9059e-11 | 29543.4188 | 590.8684
G07 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | G07 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^-1 | Q96 | G04 | Internal Control | lyophilized | rep 4 | CY5 | 31.99 | 10^2
G08 | Internal Control | Unknown | CY5 | 32.36 | [5 empty] | Q96 | G08 | Internal Control | lyophilized | rep 4 | CY5 | 32.36 | 10^-2 | Q96 | G05 | Internal Control | lyophilized | rep 4 | CY5 | 32.96 | 10^1
G08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | G08 | Legionella sp. | lyophilized | rep 4 | FAM |  | 10^-2 | Q96 | G06 | Internal Control | lyophilized | rep 4 | CY5 | 31.8 | 10^0
G09 | Internal Control | Unknown | CY5 | 31.74 | [5 empty] | Q96 | G09 | Internal Control | lyophilized | rep 4 | CY5 | 31.74 | 10^-3 | Q96 | G07 | Internal Control | lyophilized | rep 4 | CY5 | 31.85 | 10^-1
G09 | Legionella sp. | Unknown | FAM | 30.87 | [5 empty] | Q96 | G09 | Legionella sp. | lyophilized | rep 4 | FAM | 30.87 | 10^-3 | Q96 | G08 | Internal Control | lyophilized | rep 4 | CY5 | 32.36 | 10^-2
G10 | Internal Control | Unknown | CY5 | 33.11 | [5 empty] | Q96 | G10 | Internal Control | lyophilized | rep 4 | CY5 | 33.11 | NTC | Q96 | G09 | Internal Control | lyophilized | rep 4 | CY5 | 31.74 | 10^-3
G10 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | G10 | Legionella sp. | lyophilized | rep 4 | FAM |  | NTC | Q96 | G10 | Internal Control | lyophilized | rep 4 | CY5 | 33.11 | NTC
H03 | Internal Control | Unknown | CY5 | 38.9 | [5 empty] | Q96 | H03 | Internal Control | beads | rep 4 | CY5 | 38.9 | 10^3 | Q16 | A1-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.42 | 10^3
H03 | Legionella sp. | Unknown | FAM | 27.53 | [5 empty] | Q96 | H03 | Legionella sp. | beads | rep 4 | FAM | 27.53 | 10^3 | Q16 | A2-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.68 | 10^2
H04 | Internal Control | Unknown | CY5 | 36 | [5 empty] | Q96 | H04 | Internal Control | beads | rep 4 | CY5 | 36 | 10^2 | Q16 | A3-4 | Internal Control | lyophilized | rep 5 | CY5 | 32.75 | 10^1
H04 | Legionella sp. | Unknown | FAM | 35.93 | [5 empty] | Q96 | H04 | Legionella sp. | beads | rep 4 | FAM | 35.93 | 10^2 | Q16 | A4-4 | Internal Control | lyophilized | rep 5 | CY5 | 34.16 | 10^0
H05 | Internal Control | Unknown | CY5 | 33.34 | [5 empty] | Q96 | H05 | Internal Control | beads | rep 4 | CY5 | 33.34 | 10^1 | Q16 | A5-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.67 | 10^-1
H05 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | H05 | Legionella sp. | beads | rep 4 | FAM |  | 10^1 | Q16 | A6-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.75 | 10^-2
H06 | Internal Control | Unknown | CY5 | 33.83 | [5 empty] | Q96 | H06 | Internal Control | beads | rep 4 | CY5 | 33.83 | 10^0 | Q16 | A7-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.06 | 10^-3
H06 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | H06 | Legionella sp. | beads | rep 4 | FAM |  | 10^0 | Q16 | A8-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.79 | NTC
H07 | Internal Control | Unknown | CY5 | 32.52 | [5 empty] | Q96 | H07 | Internal Control | beads | rep 4 | CY5 | 32.52 | 10^-1 | Q96 | B03 | Internal Control | beads | rep 1 | CY5 | 41.78 | 10^3
H07 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | H07 | Legionella sp. | beads | rep 4 | FAM |  | 10^-1 | Q96 | B04 | Internal Control | beads | rep 1 | CY5 | 34.07 | 10^2 | [10 empty] | the error associated with 95% could actually mean that it's more like
H08 | Internal Control | Unknown | CY5 | 32.47 | [5 empty] | Q96 | H08 | Internal Control | beads | rep 4 | CY5 | 32.47 | 10^-2 | Q96 | B05 | Internal Control | beads | rep 1 | CY5 | 34.26 | 10^1 | [9 empty] | 95+17.328 | 112.328
H08 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | H08 | Legionella sp. | beads | rep 4 | FAM |  | 10^-2 | Q96 | B06 | Internal Control | beads | rep 1 | CY5 | 34.68 | 10^0 | [9 empty] | 95-17.328 | 77.672
H09 | Internal Control | Unknown | CY5 | 32.64 | [5 empty] | Q96 | H09 | Internal Control | beads | rep 4 | CY5 | 32.64 | 10^-3 | Q96 | B07 | Internal Control | beads | rep 1 | CY5 | 33.43 | 10^-1
H09 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | H09 | Legionella sp. | beads | rep 4 | FAM |  | 10^-3 | Q96 | B08 | Internal Control | beads | rep 1 | CY5 | 32.94 | 10^-2 | [9 empty] | What is the Log(uM) associated with these points?
H10 | Internal Control | Unknown | CY5 | 33.93 | [5 empty] | Q96 | H10 | Internal Control | beads | rep 4 | CY5 | 33.93 | NTC | Q96 | B09 | Internal Control | beads | rep 1 | CY5 | 34.31 | 10^-3
H10 | Legionella sp. | Unknown | FAM | [6 empty] | Q96 | H10 | Legionella sp. | beads | rep 4 | FAM |  | NTC | Q96 | B10 | Internal Control | beads | rep 1 | CY5 | 34.35 | NTC | [9 empty] | conc at 112 doesn't exist, asymptote at 103.
[10 empty] | Q16 | A1-1 | Legionella sp. | lyophilized | rep 5 | FAM | 27.34 | 10^3 | Q96 | D03 | Internal Control | beads | rep 2 | CY5 | 42.7 | 10^3
[10 empty] | Q16 | A1-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.42 | 10^3 | Q96 | D04 | Internal Control | beads | rep 2 | CY5 | 34 | 10^2
[10 empty] | Q16 | A2-1 | Legionella sp. | lyophilized | rep 5 | FAM | 31.98 | 10^2 | Q96 | D05 | Internal Control | beads | rep 2 | CY5 | 33.83 | 10^1
[10 empty] | Q16 | A2-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.68 | 10^2 | Q96 | D06 | Internal Control | beads | rep 2 | CY5 | 33.57 | 10^0
[10 empty] | Q16 | A3-1 | Legionella sp. | lyophilized | rep 5 | FAM | 32.81 | 10^1 | Q96 | D07 | Internal Control | beads | rep 2 | CY5 | 33.53 | 10^-1
[10 empty] | Q16 | A3-4 | Internal Control | lyophilized | rep 5 | CY5 | 32.75 | 10^1 | Q96 | D08 | Internal Control | beads | rep 2 | CY5 | 33.7 | 10^-2
[10 empty] | Q16 | A4-1 | Legionella sp. | lyophilized | rep 5 | FAM | 37.65 | 10^0 | Q96 | D09 | Internal Control | beads | rep 2 | CY5 | 34.06 | 10^-3
[10 empty] | Q16 | A4-4 | Internal Control | lyophilized | rep 5 | CY5 | 34.16 | 10^0 | Q96 | D10 | Internal Control | beads | rep 2 | CY5 | 33.24 | NTC
[10 empty] | Q16 | A5-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | 10^-1 | Q96 | F03 | Internal Control | beads | rep 3 | CY5 | 44.31 | 10^3
[10 empty] | Q16 | A5-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.67 | 10^-1 | Q96 | F04 | Internal Control | beads | rep 3 | CY5 | 34.96 | 10^2
[10 empty] | Q16 | A6-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | 10^-2 | Q96 | F05 | Internal Control | beads | rep 3 | CY5 | 33.32 | 10^1
[10 empty] | Q16 | A6-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.75 | 10^-2 | Q96 | F06 | Internal Control | beads | rep 3 | CY5 | 33.86 | 10^0
[10 empty] | Q16 | A7-1 | Legionella sp. | lyophilized | rep 5 | FAM | 37.83 | 10^-3 | Q96 | F07 | Internal Control | beads | rep 3 | CY5 | 34.19 | 10^-1
[10 empty] | Q16 | A7-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.06 | 10^-3 | Q96 | F08 | Internal Control | beads | rep 3 | CY5 | 33.33 | 10^-2
[10 empty] | Q16 | A8-1 | Legionella sp. | lyophilized | rep 5 | FAM |  | NTC | Q96 | F09 | Internal Control | beads | rep 3 | CY5 | 34.73 | 10^-3
[10 empty] | Q16 | A8-4 | Internal Control | lyophilized | rep 5 | CY5 | 33.79 | NTC | Q96 | F10 | Internal Control | beads | rep 3 | CY5 | 32.77 | NTC
[10 empty] | Q16 | B1-1 | Legionella sp. | beads | rep 5 | FAM | 28.15 | 10^3 | Q96 | H03 | Internal Control | beads | rep 4 | CY5 | 38.9 | 10^3
[10 empty] | Q16 | B1-4 | Internal Control | beads | rep 5 | CY5 |  | 10^3 | Q96 | H04 | Internal Control | beads | rep 4 | CY5 | 36 | 10^2 | [9 empty] | dangerous to take the highest value because this is asymptote. Should I just use 100?
[10 empty] | Q16 | B2-1 | Legionella sp. | beads | rep 5 | FAM | 30.74 | 10^2 | Q96 | H05 | Internal Control | beads | rep 4 | CY5 | 33.34 | 10^1
[10 empty] | Q16 | B2-4 | Internal Control | beads | rep 5 | CY5 | 30.73 | 10^2 | Q96 | H06 | Internal Control | beads | rep 4 | CY5 | 33.83 | 10^0
[10 empty] | Q16 | B3-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^1 | Q96 | H07 | Internal Control | beads | rep 4 | CY5 | 32.52 | 10^-1
[10 empty] | Q16 | B3-4 | Internal Control | beads | rep 5 | CY5 | 33.3 | 10^1 | Q96 | H08 | Internal Control | beads | rep 4 | CY5 | 32.47 | 10^-2
[10 empty] | Q16 | B4-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^0 | Q96 | H09 | Internal Control | beads | rep 4 | CY5 | 32.64 | 10^-3
[10 empty] | Q16 | B4-4 | Internal Control | beads | rep 5 | CY5 | 33.27 | 10^0 | Q96 | H10 | Internal Control | beads | rep 4 | CY5 | 33.93 | NTC
[10 empty] | Q16 | B5-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^-1 | Q16 | B1-4 | Internal Control | beads | rep 5 | CY5 |  | 10^3
[10 empty] | Q16 | B5-4 | Internal Control | beads | rep 5 | CY5 | 33.1 | 10^-1 | Q16 | B2-4 | Internal Control | beads | rep 5 | CY5 | 30.73 | 10^2
[10 empty] | Q16 | B6-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^-2 | Q16 | B3-4 | Internal Control | beads | rep 5 | CY5 | 33.3 | 10^1
[10 empty] | Q16 | B6-4 | Internal Control | beads | rep 5 | CY5 | 34.64 | 10^-2 | Q16 | B4-4 | Internal Control | beads | rep 5 | CY5 | 33.27 | 10^0
[10 empty] | Q16 | B7-1 | Legionella sp. | beads | rep 5 | FAM |  | 10^-3 | Q16 | B5-4 | Internal Control | beads | rep 5 | CY5 | 33.1 | 10^-1
[10 empty] | Q16 | B7-4 | Internal Control | beads | rep 5 | CY5 | 33.74 | 10^-3 | Q16 | B6-4 | Internal Control | beads | rep 5 | CY5 | 34.64 | 10^-2
[10 empty] | Q16 | B8-1 | Legionella sp. | beads | rep 5 | FAM |  | NTC | Q16 | B7-4 | Internal Control | beads | rep 5 | CY5 | 33.74 | 10^-3
[10 empty] | Q16 | B8-4 | Internal Control | beads | rep 5 | CY5 | 34.17 | NTC | Q16 | B8-4 | Internal Control | beads | rep 5 | CY5 | 34.17 | NTC

## 2nd LOD
Well | Target | Property | Dye | Ct | Thermal Cycler | Well | Target | Property | Dye | Ct | Rep | Sample
C03 | Internal Control | Unknown | CY5 | 33.02 | Q96 | C03 | Internal Control | Unknown | CY5 | 33.02 | rep 1 | #11
C03 | Legionella sp. | Unknown | FAM | 40.67 | Q96 | C04 | Internal Control | Unknown | CY5 | 32.98 | rep 1 | #12 | FROM pivot table
C04 | Internal Control | Unknown | CY5 | 32.98 | Q96 | C05 | Internal Control | Unknown | CY5 | 32.99 | rep 1 | #13 | Target | Leg pneumo
C04 | Legionella sp. | Unknown | FAM | 39.73 | Q96 | C06 | Internal Control | Unknown | CY5 | 33.03 | rep 1 | #14
C05 | Internal Control | Unknown | CY5 | 32.99 | Q96 | C07 | Internal Control | Unknown | CY5 | 32.39 | rep 1 | #11_2 | Max of Ct |  | Sample
C05 | Legionella sp. | Unknown | FAM | 41.96 | Q96 | C08 | Internal Control | Unknown | CY5 | 33 | rep 1 | #12_2 | Thermal Cycler | Rep | #11 | #11_2 | #12 | #12_2 | #13 | #13_2 | #14 | Grand Total
C06 | Internal Control | Unknown | CY5 | 33.03 | Q96 | C09 | Internal Control | Unknown | CY5 | 33.55 | rep 1 | #13_2 | Q16 | rep 5 | 0 | 39.27 | 0 | 0 | 0 | 39.23 | 0 | 39.27
C06 | Legionella sp. | Unknown | FAM |  | Q96 | C10 | Internal Control | Unknown | CY5 | 33.04 | rep 1 | NTC |  | rep 6 | 34.99 | 0 | 0 | 0 | 0 | 36.08 | 0 | 36.08
C07 | Internal Control | Unknown | CY5 | 32.39 | Q96 | D03 | Internal Control | Unknown | CY5 | 34.05 | rep 2 | #11 | Q16 Total |  | 34.99 | 39.27 | 0 | 0 | 0 | 39.23 | 0 | 39.27
C07 | Legionella sp. | Unknown | FAM | 34.91 | Q96 | D04 | Internal Control | Unknown | CY5 | 32.69 | rep 2 | #12 | Q96 | rep 1 | 40.67 | 34.91 | 39.73 | 0 | 41.96 | 0 | 0 | 41.96
C08 | Internal Control | Unknown | CY5 | 33 | Q96 | D05 | Internal Control | Unknown | CY5 | 32.59 | rep 2 | #13 |  | rep 2 | 38.5 | 37.25 | 0 | 35.92 | 0 | 0 | 0 | 38.5
C08 | Legionella sp. | Unknown | FAM |  | Q96 | D06 | Internal Control | Unknown | CY5 | 33.03 | rep 2 | #14 |  | rep 3 | 36.67 | 37.08 | 36.17 | 0 | 0 | 0 | 38.49 | 38.49 |  | #14 sample is likely contamination as earlier, more concentrated samples were negative.
C09 | Internal Control | Unknown | CY5 | 33.55 | Q96 | D07 | Internal Control | Unknown | CY5 | 33.45 | rep 2 | #11_2 |  | rep 4 | 38.74 | 38.72 | 36.36 | 0 | 0 | 38.46 | 0 | 38.74
C09 | Legionella sp. | Unknown | FAM |  | Q96 | D08 | Internal Control | Unknown | CY5 | 32.23 | rep 2 | #12_2 | Q96 Total |  | 40.67 | 38.72 | 39.73 | 35.92 | 41.96 | 38.46 | 38.49 | 41.96
C10 | Internal Control | Unknown | CY5 | 33.04 | Q96 | D09 | Internal Control | Unknown | CY5 | 32.6 | rep 2 | #13_2 | Grand Total |  | 40.67 | 39.27 | 39.73 | 35.92 | 41.96 | 39.23 | 38.49 | 41.96
C10 | Legionella sp. | Unknown | FAM | 37.15 | Q96 | D10 | Internal Control | Unknown | CY5 | 33.84 | rep 2 | NTC
D03 | Internal Control | Unknown | CY5 | 34.05 | Q96 | E03 | Internal Control | Unknown | CY5 | 32.75 | rep 3 | #11 |  |  | Q96
D03 | Legionella sp. | Unknown | FAM | 38.5 | Q96 | E04 | Internal Control | Unknown | CY5 | 32.85 | rep 3 | #12 |  |  | #11 | #12 | #13 | #14
D04 | Internal Control | Unknown | CY5 | 32.69 | Q96 | E05 | Internal Control | Unknown | CY5 | 32.74 | rep 3 | #13 | Total Samples |  | 8 | 8 | 8 | 8
D04 | Legionella sp. | Unknown | FAM |  | Q96 | E06 | Internal Control | Unknown | CY5 | 33.44 | rep 3 | #14 | Total Positive |  | 8 | 4 | 2 | 0
D05 | Internal Control | Unknown | CY5 | 32.59 | Q96 | E07 | Internal Control | Unknown | CY5 | 32.31 | rep 3 | #11_2 | Fraction Positive |  | 100% | 50% | 25% | 0%
D05 | Legionella sp. | Unknown | FAM |  | Q96 | E08 | Internal Control | Unknown | CY5 | 32.63 | rep 3 | #12_2
D06 | Internal Control | Unknown | CY5 | 33.03 | Q96 | E09 | Internal Control | Unknown | CY5 | 33.12 | rep 3 | #13_2
D06 | Legionella sp. | Unknown | FAM |  | Q96 | E10 | Internal Control | Unknown | CY5 | 33.87 | rep 3 | NTC | TubeName | Conc (uM) | Copy Number/Tube | Copy Number/rxn (20ul)
D07 | Internal Control | Unknown | CY5 | 33.45 | Q96 | F03 | Internal Control | Unknown | CY5 | 34.15 | rep 4 | #11 | IDT | 95.759 | 2306643141280637 | 4.6133e+13
D07 | Legionella sp. | Unknown | FAM | 37.25 | Q96 | F04 | Internal Control | Unknown | CY5 | 33.21 | rep 4 | #12
D08 | Internal Control | Unknown | CY5 | 32.23 | Q96 | F05 | Internal Control | Unknown | CY5 | 33.28 | rep 4 | #13 | Tube_0 | 0.95759 | 5.7666e+14 | 1.1533e+13 |  |  |  | 82.3 | -11.3822 | 4.1476e-12 | 2497.7027 | 49.9541
D08 | Legionella sp. | Unknown | FAM | 35.92 | Q96 | F06 | Internal Control | Unknown | CY5 | 32.87 | rep 4 | #14 | Tube_1 | 0.095759 | 5.7666e+13 | 1.1533e+12 |  |  |  | 95 | -10.7656 | 1.7155e-11 | 10330.9617 | 206.6192
D09 | Internal Control | Unknown | CY5 | 32.6 | Q96 | F07 | Internal Control | Unknown | CY5 | 33.05 | rep 4 | #11_2 | Tube_2 | 0.0095759 | 5.7666e+12 | 115332157064.0318 |  |  |  | 100 | -10.1419 | 7.2127e-11 | 43435.0925 | 868.7019
D09 | Legionella sp. | Unknown | FAM |  | Q96 | F08 | Internal Control | Unknown | CY5 | 33.55 | rep 4 | #12_2 | Tube_3 | 0.00095759 | 576660785320.1592 | 11533215706.4032
D10 | Internal Control | Unknown | CY5 | 33.84 | Q96 | F09 | Internal Control | Unknown | CY5 | 33.71 | rep 4 | #13_2 | Tube_4 | 9.5759e-05 | 57666078532.0159 | 1153321570.6403
D10 | Legionella sp. | Unknown | FAM |  | Q96 | F10 | Internal Control | Unknown | CY5 | 33.35 | rep 4 | NTC | Tube_5 | 9.5759e-06 | 5766607853.2016 | 115332157.064
E03 | Internal Control | Unknown | CY5 | 32.75 | Q96 | C03 | Leg pneumo | Unknown | FAM | 40.67 | rep 1 | #11 | Tube_6 | 9.5759e-07 | 576660785.3202 | 11533215.7064
E03 | Legionella sp. | Unknown | FAM | 36.67 | Q96 | C04 | Leg pneumo | Unknown | FAM | 39.73 | rep 1 | #12 | Tube_7 | 9.5759e-08 | 57666078.532 | 1153321.5706 | Log (uM) | Fraction Pos
E04 | Internal Control | Unknown | CY5 | 32.85 | Q96 | C05 | Leg pneumo | Unknown | FAM | 41.96 | rep 1 | #13 | Tube_8 | 9.5759e-09 | 5766607.8532 | 115332.1571 | -8.0188 | 1
E04 | Legionella sp. | Unknown | FAM | 36.17 | Q96 | C06 | Leg pneumo | Unknown | FAM |  | rep 1 | #14 | Tube_9 | 9.5759e-10 | 576660.7853 | 11533.2157 | -9.0188 | 1
E05 | Internal Control | Unknown | CY5 | 32.74 | Q96 | C07 | Leg pneumo | Unknown | FAM | 34.91 | rep 1 | #11_2 | Tube_10 | 9.5759e-11 | 57666.0785 | 1153.3216 | -10.0188 | 1
E05 | Legionella sp. | Unknown | FAM |  | Q96 | C08 | Leg pneumo | Unknown | FAM |  | rep 1 | #12_2 | Tube_11 | 9.5759e-12 | 5766.6079 | 115.3322 | -11.0188 | 1
E06 | Internal Control | Unknown | CY5 | 33.44 | Q96 | C09 | Leg pneumo | Unknown | FAM |  | rep 1 | #13_2 | Tube_12 | 9.5759e-13 | 576.6608 | 11.5332 | -12.0188 | 0.5
E06 | Legionella sp. | Unknown | FAM | 38.49 | Q96 | C10 | Leg pneumo | Unknown | FAM | 37.15 | rep 1 | NTC | Tube_13 | 9.5759e-14 | 57.6661 | 1.1533 | -13.0188 | 0.25
E07 | Internal Control | Unknown | CY5 | 32.31 | Q96 | D03 | Leg pneumo | Unknown | FAM | 38.5 | rep 2 | #11 | Tube_14 | 9.5759e-15 | 5.7666 | 0.11533 | -14.0188 | 0
E07 | Legionella sp. | Unknown | FAM | 37.08 | Q96 | D04 | Leg pneumo | Unknown | FAM |  | rep 2 | #12 | Tube_15 | 9.5759e-16 | 0.57666 | 0.011533 | -15.0188 | 0
E08 | Internal Control | Unknown | CY5 | 32.63 | Q96 | D05 | Leg pneumo | Unknown | FAM |  | rep 2 | #13
E08 | Legionella sp. | Unknown | FAM |  | Q96 | D06 | Leg pneumo | Unknown | FAM |  | rep 2 | #14
E09 | Internal Control | Unknown | CY5 | 33.12 | Q96 | D07 | Leg pneumo | Unknown | FAM | 37.25 | rep 2 | #11_2
E09 | Legionella sp. | Unknown | FAM |  | Q96 | D08 | Leg pneumo | Unknown | FAM | 35.92 | rep 2 | #12_2
E10 | Internal Control | Unknown | CY5 | 33.87 | Q96 | D09 | Leg pneumo | Unknown | FAM |  | rep 2 | #13_2
E10 | Legionella sp. | Unknown | FAM |  | Q96 | D10 | Leg pneumo | Unknown | FAM |  | rep 2 | NTC
F03 | Internal Control | Unknown | CY5 | 34.15 | Q96 | E03 | Leg pneumo | Unknown | FAM | 36.67 | rep 3 | #11
F03 | Legionella sp. | Unknown | FAM | 38.74 | Q96 | E04 | Leg pneumo | Unknown | FAM | 36.17 | rep 3 | #12
F04 | Internal Control | Unknown | CY5 | 33.21 | Q96 | E05 | Leg pneumo | Unknown | FAM |  | rep 3 | #13
F04 | Legionella sp. | Unknown | FAM | 36.36 | Q96 | E06 | Leg pneumo | Unknown | FAM | 38.49 | rep 3 | #14
F05 | Internal Control | Unknown | CY5 | 33.28 | Q96 | E07 | Leg pneumo | Unknown | FAM | 37.08 | rep 3 | #11_2
F05 | Legionella sp. | Unknown | FAM |  | Q96 | E08 | Leg pneumo | Unknown | FAM |  | rep 3 | #12_2
F06 | Internal Control | Unknown | CY5 | 32.87 | Q96 | E09 | Leg pneumo | Unknown | FAM |  | rep 3 | #13_2
F06 | Legionella sp. | Unknown | FAM |  | Q96 | E10 | Leg pneumo | Unknown | FAM |  | rep 3 | NTC
F07 | Internal Control | Unknown | CY5 | 33.05 | Q96 | F03 | Leg pneumo | Unknown | FAM | 38.74 | rep 4 | #11
F07 | Legionella sp. | Unknown | FAM | 38.72 | Q96 | F04 | Leg pneumo | Unknown | FAM | 36.36 | rep 4 | #12
F08 | Internal Control | Unknown | CY5 | 33.55 | Q96 | F05 | Leg pneumo | Unknown | FAM |  | rep 4 | #13
F08 | Legionella sp. | Unknown | FAM |  | Q96 | F06 | Leg pneumo | Unknown | FAM |  | rep 4 | #14
F09 | Internal Control | Unknown | CY5 | 33.71 | Q96 | F07 | Leg pneumo | Unknown | FAM | 38.72 | rep 4 | #11_2
F09 | Legionella sp. | Unknown | FAM | 38.46 | Q96 | F08 | Leg pneumo | Unknown | FAM |  | rep 4 | #12_2
F10 | Internal Control | Unknown | CY5 | 33.35 | Q96 | F09 | Leg pneumo | Unknown | FAM | 38.46 | rep 4 | #13_2
F10 | Legionella sp. | Unknown | FAM | 42.03 | Q96 | F10 | Leg pneumo | Unknown | FAM | 42.03 | rep 4 | NTC | Sample
[5 empty] | Q16 | A1-4 | Internal Control | Unknown | CY5 | 33.82 | rep 5 | #11 | a
[5 empty] | Q16 | A2-4 | Internal Control | Unknown | CY5 | 33.98 | rep 5 | #12 | a
[5 empty] | Q16 | A3-4 | Internal Control | Unknown | CY5 | 33.02 | rep 5 | #13 | a
[5 empty] | Q16 | A4-4 | Internal Control | Unknown | CY5 | 33.28 | rep 5 | #14 | a
[5 empty] | Q16 | A5-4 | Internal Control | Unknown | CY5 | 33.95 | rep 5 | #11_2 | a
[5 empty] | Q16 | A6-4 | Internal Control | Unknown | CY5 | 33.62 | rep 5 | #12_2 | a
[5 empty] | Q16 | A7-4 | Internal Control | Unknown | CY5 | 33.56 | rep 5 | #13_2 | a
[5 empty] | Q16 | A8-4 | Internal Control | Unknown | CY5 | 34.26 | rep 5 | NTC | a
[5 empty] | Q16 | B1-4 | Internal Control | Unknown | CY5 | 33.11 | rep 6 | #11 | b
[5 empty] | Q16 | B2-4 | Internal Control | Unknown | CY5 | 0 | rep 6 | #12 | b
[5 empty] | Q16 | B3-4 | Internal Control | Unknown | CY5 | 32.74 | rep 6 | #13 | b
[5 empty] | Q16 | B4-4 | Internal Control | Unknown | CY5 | 33.59 | rep 6 | #14 | b
[5 empty] | Q16 | B5-4 | Internal Control | Unknown | CY5 | 33.51 | rep 6 | #11_2 | b
[5 empty] | Q16 | B6-4 | Internal Control | Unknown | CY5 | 33.73 | rep 6 | #12_2 | b
[5 empty] | Q16 | B7-4 | Internal Control | Unknown | CY5 | 34.52 | rep 6 | #13_2 | b
[5 empty] | Q16 | B8-4 | Internal Control | Unknown | CY5 | 33.96 | rep 6 | NTC | b
[5 empty] | Q16 | A1-1 | Leg pneumo | Unknown | FAM | 0 | rep 5 | #11 | a
[5 empty] | Q16 | A2-1 | Leg pneumo | Unknown | FAM | 0 | rep 5 | #12 | a
[5 empty] | Q16 | A3-1 | Leg pneumo | Unknown | FAM | 0 | rep 5 | #13 | a
[5 empty] | Q16 | A4-1 | Leg pneumo | Unknown | FAM | 0 | rep 5 | #14 | a
[5 empty] | Q16 | A5-1 | Leg pneumo | Unknown | FAM | 39.27 | rep 5 | #11_2 | a
[5 empty] | Q16 | A6-1 | Leg pneumo | Unknown | FAM | 0 | rep 5 | #12_2 | a
[5 empty] | Q16 | A7-1 | Leg pneumo | Unknown | FAM | 39.23 | rep 5 | #13_2 | a
[5 empty] | Q16 | A8-1 | Leg pneumo | Unknown | FAM | 0 | rep 5 | NTC | a
[5 empty] | Q16 | B1-1 | Leg pneumo | Unknown | FAM | 34.99 | rep 6 | #11 | b
[5 empty] | Q16 | B2-1 | Leg pneumo | Unknown | FAM | 0 | rep 6 | #12 | b
[5 empty] | Q16 | B3-1 | Leg pneumo | Unknown | FAM | 0 | rep 6 | #13 | b
[5 empty] | Q16 | B4-1 | Leg pneumo | Unknown | FAM | 0 | rep 6 | #14 | b
[5 empty] | Q16 | B5-1 | Leg pneumo | Unknown | FAM | 0 | rep 6 | #11_2 | b
[5 empty] | Q16 | B6-1 | Leg pneumo | Unknown | FAM | 0 | rep 6 | #12_2 | b
[5 empty] | Q16 | B7-1 | Leg pneumo | Unknown | FAM | 36.08 | rep 6 | #13_2 | b
[5 empty] | Q16 | B8-1 | Leg pneumo | Unknown | FAM | 0 | rep 6 | NTC | b

## 3rd LOD
Well | Target | Property | Dye | Ct | Well | Target | Property | Dye | Ct | Std | Rep
A01 | Internal Control | Unknown | CY5 | 32.56 | A01 | Legionella sp. | Unknown | FAM | 34.3 | 1e2 | rep 1 | TubeName | Conc (uM) | Copy Number/Tube | Copy Number/rxn (20ul) | Total | Positive | Log(uM) | Fraction positive
A01 | Legionella sp. | Unknown | FAM | 34.3 | A02 | Legionella sp. | Unknown | FAM | 33.42 | 1e2 | rep 2 | Tube_9 | 9.5759e-10 | 576660.7853 | 11533.2157 |  |  | -9.0188 | 100%
A02 | Internal Control | Unknown | CY5 | 32.94 | A03 | Legionella sp. | Unknown | FAM | 34.25 | 1e2 | rep 3 | Tube_10 | 9.5759e-11 | 57666.0785 | 1153.3216 |  |  | -10.0188 | 100%
A02 | Legionella sp. | Unknown | FAM | 33.42 | A04 | Legionella sp. | Unknown | FAM | 35.4 | 1e2 | rep 4 | Tube_11 | 9.5759e-12 | 5766.6079 | 115.3322 | 24 | 24 | -11.0188 | 100%
A03 | Internal Control | Unknown | CY5 | 32.93 | A05 | Legionella sp. | Unknown | FAM | 34.01 | 1e2 | rep 5 | Tube_12 | 9.5759e-13 | 576.6608 | 11.5332 | 24 | 14 | -12.0188 | 58.33%
A03 | Legionella sp. | Unknown | FAM | 34.25 | A06 | Legionella sp. | Unknown | FAM | 34.49 | 1e2 | rep 6 | Tube_13 | 9.5759e-14 | 57.6661 | 1.1533 | 24 | 7 | -13.0188 | 29.17%
A04 | Internal Control | Unknown | CY5 | 32.8 | A07 | Legionella sp. | Unknown | FAM | 33.43 | 1e2 | rep 7 | Tube_14 | 9.5759e-15 | 5.7666 | 0.11533 | 12 | 3 | -14.0188 | 25%
A04 | Legionella sp. | Unknown | FAM | 35.4 | A08 | Legionella sp. | Unknown | FAM | 34.08 | 1e2 | rep 8 | Tube_15 | 9.5759e-16 | 0.57666 | 0.011533 |  |  | -15.0188 | 0%
A05 | Internal Control | Unknown | CY5 | 33.51 | A09 | Legionella sp. | Unknown | FAM | 34.03 | 1e2 | rep 9
A05 | Legionella sp. | Unknown | FAM | 34.01 | A10 | Legionella sp. | Unknown | FAM | 34.63 | 1e2 | rep 10 |  |  |  |  | Interpolated | uM | copies/tube | LOD
A06 | Internal Control | Unknown | CY5 | 33.89 | A11 | Legionella sp. | Unknown | FAM | 34.65 | 1e2 | rep 11 |  |  |  |  | -10.8386 | 1.4501e-11 | 8732.5434 | 174.6509 | copies/rxn
A06 | Legionella sp. | Unknown | FAM | 34.49 | A12 | Legionella sp. | Unknown | FAM | 35.66 | 1e2 | rep 12 |  |  | Hi | 100 | -10.2561 | 5.5456e-11 | 33395.476 | 667.9095 | copies/rxn
A07 | Internal Control | Unknown | CY5 | 32.58 | B01 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 1 |  |  | Lo | 81.3 | -11.5293 | 2.9558e-12 | 1779.9943 | 35.5999 | copies/rxn
A07 | Legionella sp. | Unknown | FAM | 33.43 | B02 | Legionella sp. | Unknown | FAM | 38.43 | 1e1 | rep 2
A08 | Internal Control | Unknown | CY5 | 33.07 | B03 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 3
A08 | Legionella sp. | Unknown | FAM | 34.08 | B04 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 4 | testing LOD with 12x8 strip tubes and Leg pneumo after concentration on spec.
A09 | Internal Control | Unknown | CY5 | 33.05 | B05 | Legionella sp. | Unknown | FAM | 36.33 | 1e1 | rep 5
A09 | Legionella sp. | Unknown | FAM | 34.03 | B06 | Legionella sp. | Unknown | FAM | 36.65 | 1e1 | rep 6 | TubeName | Conc (uM) | Copy Number/Tube | Copy Number/rxn (20ul)
A10 | Internal Control | Unknown | CY5 | 33.11 | B07 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 7 | IDT | 95.759 | 2306643141280637 | 4.6133e+13
A10 | Legionella sp. | Unknown | FAM | 34.63 | B08 | Legionella sp. | Unknown | FAM | 38.85 | 1e1 | rep 8
A11 | Internal Control | Unknown | CY5 | 33.96 | B09 | Legionella sp. | Unknown | FAM | 36.54 | 1e1 | rep 9 | Tube_0 | 0.95759 | 5.7666e+14 | 1.1533e+13
A11 | Legionella sp. | Unknown | FAM | 34.65 | B10 | Legionella sp. | Unknown | FAM | 35.99 | 1e1 | rep 10 | Tube_1 | 0.095759 | 5.7666e+13 | 1.1533e+12
A12 | Internal Control | Unknown | CY5 | 33.37 | B11 | Legionella sp. | Unknown | FAM | 38.8 | 1e1 | rep 11 | Tube_2 | 0.0095759 | 5.7666e+12 | 115332157064.0318
A12 | Legionella sp. | Unknown | FAM | 35.66 | B12 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 12 | Tube_3 | 0.00095759 | 576660785320.1592 | 11533215706.4032
B01 | Internal Control | Unknown | CY5 | 33.69 | C01 | Legionella sp. | Unknown | FAM | 36.83 | 1e0 | rep 1 | Tube_4 | 9.5759e-05 | 57666078532.0159 | 1153321570.6403
B01 | Legionella sp. | Unknown | FAM |  | C02 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 2 | Tube_5 | 9.5759e-06 | 5766607853.2016 | 115332157.064
B02 | Internal Control | Unknown | CY5 | 32.66 | C03 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 3 | Tube_6 | 9.5759e-07 | 576660785.3202 | 11533215.7064
B02 | Legionella sp. | Unknown | FAM | 38.43 | C04 | Legionella sp. | Unknown | FAM | 35.62 | 1e0 | rep 4 | Tube_7 | 9.5759e-08 | 57666078.532 | 1153321.5706
B03 | Internal Control | Unknown | CY5 | 32.22 | C05 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 5 | Tube_8 | 9.5759e-09 | 5766607.8532 | 115332.1571
B03 | Legionella sp. | Unknown | FAM |  | C06 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 6 | Tube_9 | 9.5759e-10 | 576660.7853 | 11533.2157
B04 | Internal Control | Unknown | CY5 | 33.2 | C07 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 7 | Tube_10 | 9.5759e-11 | 57666.0785 | 1153.3216
B04 | Legionella sp. | Unknown | FAM |  | C08 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 8 | Tube_11 | 9.5759e-12 | 5766.6079 | 115.3322
B05 | Internal Control | Unknown | CY5 | 32.79 | C09 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 9 | Tube_12 | 9.5759e-13 | 576.6608 | 11.5332
B05 | Legionella sp. | Unknown | FAM | 36.33 | C10 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 10 | Tube_13 | 9.5759e-14 | 57.6661 | 1.1533
B06 | Internal Control | Unknown | CY5 | 32.98 | C11 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 11 | Tube_14 | 9.5759e-15 | 5.7666 | 0.11533
B06 | Legionella sp. | Unknown | FAM | 36.65 | C12 | Legionella sp. | Unknown | FAM | 38.52 | 1e0 | rep 12 | Tube_15 | 9.5759e-16 | 0.57666 | 0.011533
B07 | Internal Control | Unknown | CY5 | 33.31 | D01 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 1
B07 | Legionella sp. | Unknown | FAM |  | D02 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 2
B08 | Internal Control | Unknown | CY5 | 32.82 | D03 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 3
B08 | Legionella sp. | Unknown | FAM | 38.85 | D04 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 4
B09 | Internal Control | Unknown | CY5 | 32.86 | D05 | Legionella sp. | Unknown | FAM | 37.63 | 1e-1 | rep 5
B09 | Legionella sp. | Unknown | FAM | 36.54 | D06 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 6
B10 | Internal Control | Unknown | CY5 | 32.95 | D07 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 7
B10 | Legionella sp. | Unknown | FAM | 35.99 | D08 | Legionella sp. | Unknown | FAM | 38.58 | 1e-1 | rep 8
B11 | Internal Control | Unknown | CY5 | 33.44 | D09 | Legionella sp. | Unknown | FAM | 38.17 | 1e-1 | rep 9
B11 | Legionella sp. | Unknown | FAM | 38.8 | D10 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 10
B12 | Internal Control | Unknown | CY5 | 33.31 | D11 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 11
B12 | Legionella sp. | Unknown | FAM |  | D12 | Legionella sp. | Unknown | FAM |  | 1e-1 | rep 12
C01 | Internal Control | Unknown | CY5 | 32.73 | E01 | Legionella sp. | Unknown | FAM | 34.03 | 1e2 | rep 13
C01 | Legionella sp. | Unknown | FAM | 36.83 | E02 | Legionella sp. | Unknown | FAM | 34.76 | 1e2 | rep 14
C02 | Internal Control | Unknown | CY5 | 32.95 | E03 | Legionella sp. | Unknown | FAM | 34.38 | 1e2 | rep 15
C02 | Legionella sp. | Unknown | FAM |  | E04 | Legionella sp. | Unknown | FAM | 33.83 | 1e2 | rep 16
C03 | Internal Control | Unknown | CY5 | 33.23 | E05 | Legionella sp. | Unknown | FAM | 34.45 | 1e2 | rep 17
C03 | Legionella sp. | Unknown | FAM |  | E06 | Legionella sp. | Unknown | FAM | 33.83 | 1e2 | rep 18
C04 | Internal Control | Unknown | CY5 | 33.06 | E07 | Legionella sp. | Unknown | FAM | 33.45 | 1e2 | rep 19
C04 | Legionella sp. | Unknown | FAM | 35.62 | E08 | Legionella sp. | Unknown | FAM | 34.76 | 1e2 | rep 20
C05 | Internal Control | Unknown | CY5 | 34.03 | E09 | Legionella sp. | Unknown | FAM | 34.66 | 1e2 | rep 21
C05 | Legionella sp. | Unknown | FAM |  | E10 | Legionella sp. | Unknown | FAM | 35.22 | 1e2 | rep 22
C06 | Internal Control | Unknown | CY5 | 33.29 | E11 | Legionella sp. | Unknown | FAM | 34.47 | 1e2 | rep 23
C06 | Legionella sp. | Unknown | FAM |  | E12 | Legionella sp. | Unknown | FAM | 34.31 | 1e2 | rep 24
C07 | Internal Control | Unknown | CY5 | 32.59 | F01 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 13
C07 | Legionella sp. | Unknown | FAM |  | F02 | Legionella sp. | Unknown | FAM | 38.25 | 1e1 | rep 14
C08 | Internal Control | Unknown | CY5 | 33.11 | F03 | Legionella sp. | Unknown | FAM | 37.05 | 1e1 | rep 15
C08 | Legionella sp. | Unknown | FAM |  | F04 | Legionella sp. | Unknown | FAM | 36.64 | 1e1 | rep 16
C09 | Internal Control | Unknown | CY5 | 33.3 | F05 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 17
C09 | Legionella sp. | Unknown | FAM |  | F06 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 18
C10 | Internal Control | Unknown | CY5 | 32.52 | F07 | Legionella sp. | Unknown | FAM | 39.38 | 1e1 | rep 19
C10 | Legionella sp. | Unknown | FAM |  | F08 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 20
C11 | Internal Control | Unknown | CY5 | 32.85 | F09 | Legionella sp. | Unknown | FAM | 37.4 | 1e1 | rep 21
C11 | Legionella sp. | Unknown | FAM |  | F10 | Legionella sp. | Unknown | FAM |  | 1e1 | rep 22
C12 | Internal Control | Unknown | CY5 | 32.67 | F11 | Legionella sp. | Unknown | FAM | 37.85 | 1e1 | rep 23
C12 | Legionella sp. | Unknown | FAM | 38.52 | F12 | Legionella sp. | Unknown | FAM | 36.23 | 1e1 | rep 24
D01 | Internal Control | Unknown | CY5 | 33.19 | G01 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 13
D01 | Legionella sp. | Unknown | FAM |  | G02 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 14
D02 | Internal Control | Unknown | CY5 | 32.95 | G03 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 15
D02 | Legionella sp. | Unknown | FAM |  | G04 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 16
D03 | Internal Control | Unknown | CY5 | 32.35 | G05 | Legionella sp. | Unknown | FAM | 37.13 | 1e0 | rep 17
D03 | Legionella sp. | Unknown | FAM |  | G06 | Legionella sp. | Unknown | FAM | 36.95 | 1e0 | rep 18
D04 | Internal Control | Unknown | CY5 | 33.18 | G07 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 19
D04 | Legionella sp. | Unknown | FAM |  | G08 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 20
D05 | Internal Control | Unknown | CY5 | 33.27 | G09 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 21
D05 | Legionella sp. | Unknown | FAM | 37.63 | G10 | Legionella sp. | Unknown | FAM | 37.35 | 1e0 | rep 22
D06 | Internal Control | Unknown | CY5 | 33.04 | G11 | Legionella sp. | Unknown | FAM |  | 1e0 | rep 23
D06 | Legionella sp. | Unknown | FAM |  | G12 | Legionella sp. | Unknown | FAM | 38.96 | 1e0 | rep 24
D07 | Internal Control | Unknown | CY5 | 34.12 | H01 | Legionella sp. | Unknown | FAM | 37.82 | NTC | rep 1
D07 | Legionella sp. | Unknown | FAM |  | H02 | Legionella sp. | Unknown | FAM |  | NTC | rep 2
D08 | Internal Control | Unknown | CY5 | 33.22 | H03 | Legionella sp. | Unknown | FAM |  | NTC | rep 3
D08 | Legionella sp. | Unknown | FAM | 38.58 | H04 | Legionella sp. | Unknown | FAM | 37.21 | NTC | rep 4
D09 | Internal Control | Unknown | CY5 | 33.05 | H05 | Legionella sp. | Unknown | FAM |  | NTC | rep 5
D09 | Legionella sp. | Unknown | FAM | 38.17 | H06 | Legionella sp. | Unknown | FAM |  | NTC | rep 6
D10 | Internal Control | Unknown | CY5 | 33.21 | H07 | Legionella sp. | Unknown | FAM |  | NTC | rep 7
D10 | Legionella sp. | Unknown | FAM |  | H08 | Legionella sp. | Unknown | FAM |  | NTC | rep 8
D11 | Internal Control | Unknown | CY5 | 33.73 | H09 | Legionella sp. | Unknown | FAM |  | NTC | rep 9
D11 | Legionella sp. | Unknown | FAM |  | H10 | Legionella sp. | Unknown | FAM | 39.39 | NTC | rep 10
D12 | Internal Control | Unknown | CY5 | 33.77 | H11 | Legionella sp. | Unknown | FAM |  | NTC | rep 11
D12 | Legionella sp. | Unknown | FAM |  | H12 | Legionella sp. | Unknown | FAM | 35.94 | NTC | rep 12
E01 | Internal Control | Unknown | CY5 | 32.57 | A01 | Internal Control | Unknown | CY5 | 32.56 | 1e2 | rep 1
E01 | Legionella sp. | Unknown | FAM | 34.03 | A02 | Internal Control | Unknown | CY5 | 32.94 | 1e2 | rep 2
E02 | Internal Control | Unknown | CY5 | 33.68 | A03 | Internal Control | Unknown | CY5 | 32.93 | 1e2 | rep 3
E02 | Legionella sp. | Unknown | FAM | 34.76 | A04 | Internal Control | Unknown | CY5 | 32.8 | 1e2 | rep 4
E03 | Internal Control | Unknown | CY5 | 33.14 | A05 | Internal Control | Unknown | CY5 | 33.51 | 1e2 | rep 5
E03 | Legionella sp. | Unknown | FAM | 34.38 | A06 | Internal Control | Unknown | CY5 | 33.89 | 1e2 | rep 6
E04 | Internal Control | Unknown | CY5 | 32.58 | A07 | Internal Control | Unknown | CY5 | 32.58 | 1e2 | rep 7
E04 | Legionella sp. | Unknown | FAM | 33.83 | A08 | Internal Control | Unknown | CY5 | 33.07 | 1e2 | rep 8
E05 | Internal Control | Unknown | CY5 | 32.86 | A09 | Internal Control | Unknown | CY5 | 33.05 | 1e2 | rep 9
E05 | Legionella sp. | Unknown | FAM | 34.45 | A10 | Internal Control | Unknown | CY5 | 33.11 | 1e2 | rep 10
E06 | Internal Control | Unknown | CY5 | 32.93 | A11 | Internal Control | Unknown | CY5 | 33.96 | 1e2 | rep 11
E06 | Legionella sp. | Unknown | FAM | 33.83 | A12 | Internal Control | Unknown | CY5 | 33.37 | 1e2 | rep 12
E07 | Internal Control | Unknown | CY5 | 32.54 | B01 | Internal Control | Unknown | CY5 | 33.69 | 1e1 | rep 1
E07 | Legionella sp. | Unknown | FAM | 33.45 | B02 | Internal Control | Unknown | CY5 | 32.66 | 1e1 | rep 2
E08 | Internal Control | Unknown | CY5 | 33.03 | B03 | Internal Control | Unknown | CY5 | 32.22 | 1e1 | rep 3
E08 | Legionella sp. | Unknown | FAM | 34.76 | B04 | Internal Control | Unknown | CY5 | 33.2 | 1e1 | rep 4
E09 | Internal Control | Unknown | CY5 | 33.01 | B05 | Internal Control | Unknown | CY5 | 32.79 | 1e1 | rep 5
E09 | Legionella sp. | Unknown | FAM | 34.66 | B06 | Internal Control | Unknown | CY5 | 32.98 | 1e1 | rep 6
E10 | Internal Control | Unknown | CY5 | 32.63 | B07 | Internal Control | Unknown | CY5 | 33.31 | 1e1 | rep 7
E10 | Legionella sp. | Unknown | FAM | 35.22 | B08 | Internal Control | Unknown | CY5 | 32.82 | 1e1 | rep 8
E11 | Internal Control | Unknown | CY5 | 33.48 | B09 | Internal Control | Unknown | CY5 | 32.86 | 1e1 | rep 9
E11 | Legionella sp. | Unknown | FAM | 34.47 | B10 | Internal Control | Unknown | CY5 | 32.95 | 1e1 | rep 10
E12 | Internal Control | Unknown | CY5 | 32.7 | B11 | Internal Control | Unknown | CY5 | 33.44 | 1e1 | rep 11
E12 | Legionella sp. | Unknown | FAM | 34.31 | B12 | Internal Control | Unknown | CY5 | 33.31 | 1e1 | rep 12
F01 | Internal Control | Unknown | CY5 | 32.35 | C01 | Internal Control | Unknown | CY5 | 32.73 | 1e0 | rep 1
F01 | Legionella sp. | Unknown | FAM |  | C02 | Internal Control | Unknown | CY5 | 32.95 | 1e0 | rep 2
F02 | Internal Control | Unknown | CY5 | 33.03 | C03 | Internal Control | Unknown | CY5 | 33.23 | 1e0 | rep 3
F02 | Legionella sp. | Unknown | FAM | 38.25 | C04 | Internal Control | Unknown | CY5 | 33.06 | 1e0 | rep 4
F03 | Internal Control | Unknown | CY5 | 33.65 | C05 | Internal Control | Unknown | CY5 | 34.03 | 1e0 | rep 5
F03 | Legionella sp. | Unknown | FAM | 37.05 | C06 | Internal Control | Unknown | CY5 | 33.29 | 1e0 | rep 6
F04 | Internal Control | Unknown | CY5 | 32.2 | C07 | Internal Control | Unknown | CY5 | 32.59 | 1e0 | rep 7
F04 | Legionella sp. | Unknown | FAM | 36.64 | C08 | Internal Control | Unknown | CY5 | 33.11 | 1e0 | rep 8
F05 | Internal Control | Unknown | CY5 | 32.67 | C09 | Internal Control | Unknown | CY5 | 33.3 | 1e0 | rep 9
F05 | Legionella sp. | Unknown | FAM |  | C10 | Internal Control | Unknown | CY5 | 32.52 | 1e0 | rep 10
F06 | Internal Control | Unknown | CY5 | 33.13 | C11 | Internal Control | Unknown | CY5 | 32.85 | 1e0 | rep 11
F06 | Legionella sp. | Unknown | FAM |  | C12 | Internal Control | Unknown | CY5 | 32.67 | 1e0 | rep 12
F07 | Internal Control | Unknown | CY5 | 32.52 | D01 | Internal Control | Unknown | CY5 | 33.19 | 1e-1 | rep 1
F07 | Legionella sp. | Unknown | FAM | 39.38 | D02 | Internal Control | Unknown | CY5 | 32.95 | 1e-1 | rep 2
F08 | Internal Control | Unknown | CY5 | 32.84 | D03 | Internal Control | Unknown | CY5 | 32.35 | 1e-1 | rep 3
F08 | Legionella sp. | Unknown | FAM |  | D04 | Internal Control | Unknown | CY5 | 33.18 | 1e-1 | rep 4
F09 | Internal Control | Unknown | CY5 | 32.51 | D05 | Internal Control | Unknown | CY5 | 33.27 | 1e-1 | rep 5
F09 | Legionella sp. | Unknown | FAM | 37.4 | D06 | Internal Control | Unknown | CY5 | 33.04 | 1e-1 | rep 6
F10 | Internal Control | Unknown | CY5 | 33.36 | D07 | Internal Control | Unknown | CY5 | 34.12 | 1e-1 | rep 7
F10 | Legionella sp. | Unknown | FAM |  | D08 | Internal Control | Unknown | CY5 | 33.22 | 1e-1 | rep 8
F11 | Internal Control | Unknown | CY5 | 33.96 | D09 | Internal Control | Unknown | CY5 | 33.05 | 1e-1 | rep 9
F11 | Legionella sp. | Unknown | FAM | 37.85 | D10 | Internal Control | Unknown | CY5 | 33.21 | 1e-1 | rep 10
F12 | Internal Control | Unknown | CY5 | 32.71 | D11 | Internal Control | Unknown | CY5 | 33.73 | 1e-1 | rep 11
F12 | Legionella sp. | Unknown | FAM | 36.23 | D12 | Internal Control | Unknown | CY5 | 33.77 | 1e-1 | rep 12
G01 | Internal Control | Unknown | CY5 | 32.92 | E01 | Internal Control | Unknown | CY5 | 32.57 | 1e2 | rep 13
G01 | Legionella sp. | Unknown | FAM |  | E02 | Internal Control | Unknown | CY5 | 33.68 | 1e2 | rep 14
G02 | Internal Control | Unknown | CY5 | 32.07 | E03 | Internal Control | Unknown | CY5 | 33.14 | 1e2 | rep 15
G02 | Legionella sp. | Unknown | FAM |  | E04 | Internal Control | Unknown | CY5 | 32.58 | 1e2 | rep 16
G03 | Internal Control | Unknown | CY5 | 32.8 | E05 | Internal Control | Unknown | CY5 | 32.86 | 1e2 | rep 17
G03 | Legionella sp. | Unknown | FAM |  | E06 | Internal Control | Unknown | CY5 | 32.93 | 1e2 | rep 18
G04 | Internal Control | Unknown | CY5 | 34.3 | E07 | Internal Control | Unknown | CY5 | 32.54 | 1e2 | rep 19
G04 | Legionella sp. | Unknown | FAM |  | E08 | Internal Control | Unknown | CY5 | 33.03 | 1e2 | rep 20
G05 | Internal Control | Unknown | CY5 | 33.59 | E09 | Internal Control | Unknown | CY5 | 33.01 | 1e2 | rep 21
G05 | Legionella sp. | Unknown | FAM | 37.13 | E10 | Internal Control | Unknown | CY5 | 32.63 | 1e2 | rep 22
G06 | Internal Control | Unknown | CY5 | 32.61 | E11 | Internal Control | Unknown | CY5 | 33.48 | 1e2 | rep 23
G06 | Legionella sp. | Unknown | FAM | 36.95 | E12 | Internal Control | Unknown | CY5 | 32.7 | 1e2 | rep 24
G07 | Internal Control | Unknown | CY5 | 33.28 | F01 | Internal Control | Unknown | CY5 | 32.35 | 1e1 | rep 13
G07 | Legionella sp. | Unknown | FAM |  | F02 | Internal Control | Unknown | CY5 | 33.03 | 1e1 | rep 14
G08 | Internal Control | Unknown | CY5 | 32.88 | F03 | Internal Control | Unknown | CY5 | 33.65 | 1e1 | rep 15
G08 | Legionella sp. | Unknown | FAM |  | F04 | Internal Control | Unknown | CY5 | 32.2 | 1e1 | rep 16
G09 | Internal Control | Unknown | CY5 | 32.73 | F05 | Internal Control | Unknown | CY5 | 32.67 | 1e1 | rep 17
G09 | Legionella sp. | Unknown | FAM |  | F06 | Internal Control | Unknown | CY5 | 33.13 | 1e1 | rep 18
G10 | Internal Control | Unknown | CY5 | 33.21 | F07 | Internal Control | Unknown | CY5 | 32.52 | 1e1 | rep 19
G10 | Legionella sp. | Unknown | FAM | 37.35 | F08 | Internal Control | Unknown | CY5 | 32.84 | 1e1 | rep 20
G11 | Internal Control | Unknown | CY5 | 32.94 | F09 | Internal Control | Unknown | CY5 | 32.51 | 1e1 | rep 21
G11 | Legionella sp. | Unknown | FAM |  | F10 | Internal Control | Unknown | CY5 | 33.36 | 1e1 | rep 22
G12 | Internal Control | Unknown | CY5 | 33.2 | F11 | Internal Control | Unknown | CY5 | 33.96 | 1e1 | rep 23
G12 | Legionella sp. | Unknown | FAM | 38.96 | F12 | Internal Control | Unknown | CY5 | 32.71 | 1e1 | rep 24
H01 | Internal Control | Unknown | CY5 | 33.07 | G01 | Internal Control | Unknown | CY5 | 32.92 | 1e0 | rep 13
H01 | Legionella sp. | Unknown | FAM | 37.82 | G02 | Internal Control | Unknown | CY5 | 32.07 | 1e0 | rep 14
H02 | Internal Control | Unknown | CY5 | 32.44 | G03 | Internal Control | Unknown | CY5 | 32.8 | 1e0 | rep 15
H02 | Legionella sp. | Unknown | FAM |  | G04 | Internal Control | Unknown | CY5 | 34.3 | 1e0 | rep 16
H03 | Internal Control | Unknown | CY5 | 32.77 | G05 | Internal Control | Unknown | CY5 | 33.59 | 1e0 | rep 17
H03 | Legionella sp. | Unknown | FAM |  | G06 | Internal Control | Unknown | CY5 | 32.61 | 1e0 | rep 18
H04 | Internal Control | Unknown | CY5 | 33.17 | G07 | Internal Control | Unknown | CY5 | 33.28 | 1e0 | rep 19
H04 | Legionella sp. | Unknown | FAM | 37.21 | G08 | Internal Control | Unknown | CY5 | 32.88 | 1e0 | rep 20
H05 | Internal Control | Unknown | CY5 | 33.53 | G09 | Internal Control | Unknown | CY5 | 32.73 | 1e0 | rep 21
H05 | Legionella sp. | Unknown | FAM |  | G10 | Internal Control | Unknown | CY5 | 33.21 | 1e0 | rep 22
H06 | Internal Control | Unknown | CY5 | 32.88 | G11 | Internal Control | Unknown | CY5 | 32.94 | 1e0 | rep 23
H06 | Legionella sp. | Unknown | FAM |  | G12 | Internal Control | Unknown | CY5 | 33.2 | 1e0 | rep 24
H07 | Internal Control | Unknown | CY5 | 32.29 | H01 | Internal Control | Unknown | CY5 | 33.07 | NTC | rep 1
H07 | Legionella sp. | Unknown | FAM |  | H02 | Internal Control | Unknown | CY5 | 32.44 | NTC | rep 2
H08 | Internal Control | Unknown | CY5 | 32.36 | H03 | Internal Control | Unknown | CY5 | 32.77 | NTC | rep 3
H08 | Legionella sp. | Unknown | FAM |  | H04 | Internal Control | Unknown | CY5 | 33.17 | NTC | rep 4
H09 | Internal Control | Unknown | CY5 | 32.99 | H05 | Internal Control | Unknown | CY5 | 33.53 | NTC | rep 5
H09 | Legionella sp. | Unknown | FAM |  | H06 | Internal Control | Unknown | CY5 | 32.88 | NTC | rep 6
H10 | Internal Control | Unknown | CY5 | 33.06 | H07 | Internal Control | Unknown | CY5 | 32.29 | NTC | rep 7
H10 | Legionella sp. | Unknown | FAM | 39.39 | H08 | Internal Control | Unknown | CY5 | 32.36 | NTC | rep 8
H11 | Internal Control | Unknown | CY5 | 33.19 | H09 | Internal Control | Unknown | CY5 | 32.99 | NTC | rep 9
H11 | Legionella sp. | Unknown | FAM |  | H10 | Internal Control | Unknown | CY5 | 33.06 | NTC | rep 10
H12 | Internal Control | Unknown | CY5 | 33.54 | H11 | Internal Control | Unknown | CY5 | 33.19 | NTC | rep 11
H12 | Legionella sp. | Unknown | FAM | 35.94 | H12 | Internal Control | Unknown | CY5 | 33.54 | NTC | rep 12

## LOD_sort
Reproducibility of Legionella LOD Calculations in 3 Separate
| 1st LOD | 2nd LOD | 3rd LOD | AVERAGE
LOD_High | 590.8684 | 868.7019 | 667.9095 | 709.1599
LOD | 187.4998 | 206.6192 | 174.6509 | 189.59
LOD_low | 35.178 | 49.9541 | 35.5999 | 40.244
#Samples | 48 | 48 | 96